## Slide 1: Active Stabilization of the SRF Resonance Frequency FERMILAB Experience
Warren Schappert Yuriy Pischalnikov
FERMILAB

## Slide 2: Microphonics and df/dP
Project X plans to use cavities with bandwidths as narrow as 20 Hz
Narrow bandwidth cavities can be detuned by changes in He pressure (df/dP) and by mechanical vibrations (microphonics)
  Power required to maintaining the gradient in a detuned cavity proportional to the square of the detuning
  Results of detuning can be either or both of
    Increased capital and operating costs
    Frequent cavity trips
  Goal to keep PEAK detuning <= cavity bandwidth
Mitigation strategies
  Passive
    Reduce pressure variations
    Reduce df/dP
    Reduce vibration levels
    Provide more RF power overhead
  Active
    Use fast or slow tuner to stabilize dynamically
    resonant frequency of cavity
Tremendous progress with passive mitigation
over the past year
Active compensation may still be
required for reliable operation
Project X meeting, ORNL/SNS, April 12-14, 2011,
Section | Number of cavities/ c.modules | Max gain per cavity (MeV) | Minimum bandwidth (Hz) | Beam Power per cavity (kW) | RF Power (kW)
SSR0 | 18/1 | 1 | 35 | 1 | 1
SSR1 | 20/2 | 2.2 | 36 | 2.2 | 5
SSR2 | 44/4 | 3.9 | 24 | 3.9 | 6
LE650 | 42/7 | 11.6 | 21 | 11.6 | 30
HE650 | 152/19 | 17.4 | 24 | 17.3 | 30

## Slide 3: Active Compensation
Piezo actuators used to reduce microphonics in CKM cavities by 20 dB (Carcagno et al. 2003 --FNAL)
Bessy able to limit pressure related detuning in1.3 GHz CW cavities to less than 1 Hz (Axel Neumann Thesis )
Other work by groups at MSU, Cornell, Argonne, JLab, …
Project X meeting, ORNL/SNS, April 12-14, 2011,

## Slide 4: CCII- 1,3GHz Tesla style cavity (our “sand-box” to gain experience in SRF cavity control field… Now is part of photo-injector at NML)
Test hardware and algorithms for NML
Development of LS LFD Algorithms
Development of Microphonics Compensation Algorithms*
*Detail study of Microphonics was never main objectives of CC2 study… microphonics up to 40-50Hz was “good enough for ILC project”
Project X meeting, ORNL/SNS, April 12-14, 2011,

## Slide 5: 325MHz Single Spoke SRF Resonator
SSR1 –originally designed for HINS.
RF- pulsed operation at 4.5K.
Slow/Fast tuner requirements:
Compact size (in beam-line axis);
250KHz slow tuning range;
Fast Tuner: 4-Piezos
to compensate LFD &
and dF/dP –
frequency fluctuation as a results of large dF/dP =140Hz/torr and 4.5K He bath pressure fluctuation…
Project X meeting, ORNL/SNS, April 12-14, 2011,
Encapsulated
Piezo assembly
Stepper Motor & Harmonics Drive
“Bullet” with SGs

## Slide 6: dF/dP compensation SSR1 with Narrow Bandwidth Coupler
Initial cold tests of SSR1 using CW provided an opportunity to gain experience with a very narrow bandwidth cavity
  Test Conditions
    4.5K
    Cavity bandwidth of about 1.5 Hz
    df/dP ~= 140 Hz/torr
    dPPTP~=5 torr
    LLRF tracking resonant frequency of cavity
  Detuning control system
    100 kHz digitizer measured RF frequency offset from an arbitrary set point
    FPGA fed f back to fast tuner
Reduced pressure related variations in cavity frequency from several hundreds Hz to
  f <= 1.3 Hz RMS

## Slide 7: dF/dP compensation SSR1 with Wide Bandwidth Coupler
Second round of CW tests provided an opportunity to gain experience with a wider bandwidth cavity and fixed RF drive frequency
  Test Conditions
    4.5K
    Cavity bandwidth ~100 Hz
    RF drive frequency fixed ~325 MHz
  Detuning control system
    104 MHz digitizer measured RF phase difference between forward and probe signals
    CPU fed  back to fast tuner
Able to limit pressure related variations
Gradient <= 0.2% RMS
Phase <= 1.2o RMS
Project X meeting, ORNL/SNS, April 12-14, 2011,

## Slide 8: Microphonics CCII experience
Example of successful passive reduction of vibration level at CCII.
Project X meeting, ORNL/SNS, April 12-14, 2011,

## Slide 9: Microphonics Compensation CCII, wide (150-165 Hz) mechanical vibration
The Piezo tuner actively damped vibrations induced by external sources. An IIR filter bank was used to isolate and reverse the phase of the particular spectral line. The phase reversed signal then fed back to the piezo tuner.
Active damping reduces the vibration at this frequency by 15db
Time, sec
Frequency, Hz
Piezo OFF
Piezo ON
Project X meeting, ORNL/SNS, April 12-14, 2011,

## Slide 10
Microphonics Compensation CCII, narrow (92 Hz) mechanical vibration
Active damping reduces the vibration at this frequency
by 15 dB.

## Slide 11: Next Step (short term)
Continue to work with SSR1 and wide bandwidth coupler for immediate future
Repeat measurements with improved controller
  Phase Measurement
    Currently ~1o/Sample @ 20S/s
    0.006o @ 1kHz possible
  Use more sophisticated filter
    Replace simple integrator with Kalman filter
  Improve controller bandwidth
    Replace CPU with FPGA
  FPGA firmware modifications required
Goal is to understand limits to active control
  Reassess plans for SSR1 when performance of improved controller has been evaluated
Project X meeting, ORNL/SNS, April 12-14, 2011,

## Slide 12: Application of Adaptive LS LFD algorithm for long RF Pulse (“proof-of concept” test related to 3-8GeV pulsed linac-Project X) (test has been performed at HTS)
FlatTop=4ms
Fill=4ms
Detuning of nine-cell elliptical cavity driven by an 8ms pulse with gradient Eacc=22MV/m.
Cavity detuning without piezo compensation aprox. several kHz.
This was sufficient to drive cavity completely off resonance.
8ms-RFpulse
10ms
Start of RF pulse
Piezo Stimulus Pulse

## Slide 13: Piezo Actuator Reliability
Repeated questions about reliability of piezo actuators based on experience at SNS
  Piezos susceptible to damage by
    Shear force
    Humidity
    High slew rates
  Highly reliable if used properly
    Diesel fuel injectors
    Laptop transformers
Problems at SNS are due to failures of mechanical capsule
Actively seeking information about long term reliability
Project X meeting, ORNL/SNS, April 12-14, 2011,

## Slide 14: SNS – piezo tuner
Project X meeting, ORNL/SNS, April 12-14, 2011,

## Slide 15: Summary
Active detuning control may be required for some or all of the Project X cavities
Possible to stabilize cavity resonant frequency to several Hz or better even at 4.5K using relatively simple closed loop
  Substantial performance improvements may be possible using a more sophisticated controller
  Plan to use SSR1 to gain experience and more fully understand factors that limit active control
Active compensation of microphonics (detuning of the cavity by external mechanical vibrations) by 15db has been demonstrated (for short period of time). This is just “illustration” rather than prove of stable and reliable operation.
Piezo actuators can be extremely reliable if treated properly
  Tuner problems at SNS result of design deficiencies
  Work still required to fully assess long term reliability

## Slide 16
Additional Slides