## Slide 1: Development of Optical Metrology ENgine (OMEN)
Ben Sheff

## Slide 2: Introduction
APS is a 7 GeV, storage-ring-based hard x-ray synchrotron radiation source at Argonne National Laboratory
  Funded by DOE Office of Science, Basic Energy Sciences
Produces collimated, intense beams of hard x-ray radiation
  Useful for a variety of applications
  Requires advanced optics to manipulate
Image courtesy of https://www1.aps.anl.gov/About/Welcome

## Slide 3: X-ray Mirrors
As beamlines get better, need better mirrors
  Residual slope error as low as 50 nrad
Autocollimator-based Long trace profiler (AC-LTP) is a high precision metrology tool1,2
  Measures slope of mirror at each point to within 50 nrad
  Extremely accurate for small range of angles, but has problems farther out
  New analysis tools needed to handle this

## Slide 4: Analysis Software
Developed Optical Metrology ENgine (OMEN) in Python
Handles standard analysis
  Easy, dynamic filtering
  Fitting options, with live-updated residue plot
  Region-of-interest selection
Tools to deal with high angles
  Stitching together overlapped, adjacent scans of mirror segments
  Adding calibration run

## Slide 5: Object Oriented Design
OMEN class handles GUI and manages companion classes
Three companion classes
  Raven handles data collection and most processing
  Lightning handles stitching and mirror profile comparison
  Medium handles input/output standard formats
Easy to modify, and to re-use parts for other purposes
Raven
Lightning
Medium1
OMEN
Image courtesy of Blazon.com/heraldry/
Image courtesy of Lyoha123, uploaded to Wikimedia commons
Image courtesy of Sid Meier’s Civilization V

## Slide 6: OMEN in Action

## Slide 7: Dealing with Highly Curved Mirrors
Preliminary Work: Take overlapping partial scans and stitch together
  Each scan over a small region over which the mirror is roughly linear
  Tilt mirror so the scan center is horizontal
  Use Lightning to stitch together partial scans
  Compare the result with full scan

## Slide 8: Calibration
Significant deviation between stitched and non-stitched
  Seems to correlate with angle measured
New high precision small angle generator against which to calibrate3
  Accurate on the order of 10 nrad
Tested small calibration run
  Inverted and interpolated data to make a look-up table
  Further investigation warranted, with more accurate measurements

## Slide 9: Conclusions and Future Direction
Code is robust, and will be useful for future analysis
  Easy to use
  Rapid analysis
Stitching code allows for new measurement capability
Further investigation will be conducted into calibration in the future
  More accurate measurements needed
Some known systematic errors introduced by analysis have yet to be corrected
  eg., All interpolation is simple linear as of yet
http://img2.goodfon.su/original/2048x1360/7/39/doroga-vdal-povorot-izgib.jpg

## Slide 10: Acknowledgements
Lee Teng Undergraduate Fellowship for their generous sponsorship and for granting me the opportunity to take part in this research.
My mentor, Dr. Assoufid for welcoming me into his group, and being willing to meet me almost every day to discuss my project.
Jun Qian and Joe Sullivan for their work building, programming, and operating the equipment, and for helping me understand what the data was and how to work with it.
The Optics Group in the X-ray Science Division of the Advanced Photon Source at Argonne National Laboratory, and Argonne National Laboratory itself for hosting me for the summer.

## Slide 11: References
1 Lahsen Assoud et al., Nucl. Instrum. Methods A 710, 31-36 (2013).
2 J. Qian, J. Sullivan, M. Erdmann, A. Khounsary, L. Assoud, Nucl. Instrum. Methods A 710, 48-51 (2013).
3 D. Shu, J. Qian, W. Liu, S. Kearney, J. Anton, J. Sullivan, and L. Assoud, Proc. SPIE 9206, Advances in Metrology for X-Ray and EUV Optics V, 920601 (October 7, 2014); doi: 10.1117/12.2084726Proc. of SPIE Vol. 9206

## Slide 12: Questions