## Slide 1: LQS Shipping Analysis & Future Plans
Carlo Santini

## Slide 2: Accelerometers
ax, ay, az
Every period (60 s) the accelerator records max, min and mean values (Data Records)
If a value excedes a threshold (2g), the accelerometer saves data from 0.25 seconds before to 0.5 seconds after the EVENT (Event Records)
|a|, a/|a|
MEASURED QUANTITIES SETTING & STORAGE
T [K]
Sensr GP1 Programmable Accelerometer
100 samples per second
t [date]
[g]

## Slide 3: Example: Data Records
Data Records:
The max, min and mean values every period (60 s)

## Slide 4: Example: Event Records
Event Records:
ax, ay, az,|a| trends during each Event.

## Slide 5: Trips Analyzed
LQS01: LBNL to FNAL, IB3 to IB1
LQS01b: LBNL to FNAL, IB3 to IB1, IB1 to IB3, FNAL to LBNL
LQS02: LBNL to FNAL
LQS03: FNAL to LBNL

## Slide 6: LQS01
From LBNL to FNAL
09/10/09 – 09/15/09
From IB3 to IB1
10/29/09 – 10/30/09

## Slide 7: LBNL-FNAL 09/10/09 – 09/15/09
EVENT
09/11/09 13:36:59
Δax = -1.779g
Δay = 1.447g
Δaz = -4.648g
Sensor 3 (Truck)
Sensor 1 (Magnet)
Sensor 2 (Magnet)
NO EVENT
9/11/09 13:36:00-13:37:00
NO EVENT
9/11/09 13:36:45-13:37:45
|Δa|= 5.183g
Max |Δa|= 0.53g
Max |Δa|= 0.90g
Very high acceleration magnitude
The shock absorbers worked very well
but

## Slide 8: Event Record

## Slide 9: IB3-IB1 10/29/09 – 10/30/09
EVENT
10/30/09 13:38:50
Δax = 1.177g
Δay = 0.112g
Δaz = 2.193g
Sensor 1 (Magnet)
|Δa|= 2.491g
EVENT
10/30/09 13:38:25
Δax = 1.381g
Δay = -1.696g
Δaz = 2.153g
Sensor 2 (Magnet)
|Δa|= 3.069g
Slightly above the threshold
There were only one EVENT above the threshold

## Slide 10: Event Record
Sensor
1
Sensor
2

## Slide 11: Conclusions
Shock Absorbers worked well.
There are data from 8 different trips of 4 different magnets (LQS01, LQS01b, LQS02, LQS03).
The only significant EVENT above the threshold did not cause adverse effects.

## Slide 12: Future (Now)
Two accelerometers will be installed in the QXFS#6’s shipping crate:
one anchored to the coil support (shocks absorbed).
one anchored to the external crate.

## Slide 13: Coil Support

## Slide 14: Coil Ready for Shipping

## Slide 15: Trucks Shipping
LTL Shipping
(Less than TruckLoad)
FTL Shipping
(Full TruckLoad)
Shipment stays in the same truck from point A to point B
Expensive (>1000$)
Faster (3-4 days to Berkley)
Packages are unloaded and re-loaded multiple times during transit.
Cheap (100-200$).
Slower (5-6 days to Berkley)
QXFS#6
MQXFS#1