## Slide 1: Beam Loss Monitors
Randy Thurman-Keup – AD/Instrumentation Physicist
John Seraphin – AD/Instrumentation Technician
Andrea Saewert – AD/Instrumentation Technician
External Beamline & Instrumentation Independent Design Review
6 October 2015

## Slide 2: Outline
M4 Line Beam Loss Monitors (BLM)
M4 Line BLM Readout Electronics
M4 Line BLM Installation Efforts
Delivery Ring BLMs

## Slide 3: M4 Line BLMs – Physical Detectors
Physical detectors will be ionization chambers
Sealed, Ar filled, concentric cylindrical electrodes
Sensitivity of 70 nC per Rad
Require 30 loss monitors
  Will be reused from the Tevatron
Outer Cathode ~40 mm DIA
Inner Anode
~6 mm DIA
Electrodes ~100 mm length

## Slide 4: M4 Line BLMs – Readout Electronics
VME based
  8 Digitizer Cards (4 ch / card)
  Timing Card
  Abort Card
  Control Card
  High Voltage Card
  Crate Controller
Filter Box
  Ferrite core for common mode noise rejection
  Additional capacitance to equalize time constants for all cables
Electronics will be reused from the Tevatron
One crate located in MC-1 will readout all 30 loss monitors

## Slide 5: M4 Line BLMs – Readout Electronics
Digitizer Card
  Operates with a pair of alternating integrators
  No deadtime
  Single integration period is ~20 ms
  Circular buffer contains 64k integration samples
  Three additional variable length sums of the single integration samples
    The sum lengths are specified by the user
  The maximum integration rate is 720 Rad/s with a single ADC count corresponding to 11 mRad/s or 0.22 mRad in ~20 ms
  Performs abort threshold comparison on single integration samples and all three sum length values

## Slide 6: M4 Line BLMs – Readout Electronics
Abort Card
  Gets over-threshold information from Digitizer
  Can mask off channels
  Can require multiplicity
  Can generate a digital abort signal for each summing length
Timing Card
  Takes TCLK, MDAT, BSCLK
    TCLK and MDAT sent to Control Card
    BSCLK used for AA marker to generate integration clock
      Also has internal timing generator for transfer lines

## Slide 7: M4 Line BLMs – Readout Electronics
Control Card
  Initializes other cards in crate based on information obtained from the VME crate processor
  Reads loss data from Digitizer and stores in circular buffers for access by crate processor
  Starts and stops integration based on TCLK events
  Changes abort settings based on MDAT machine state

## Slide 8: M4 Line BLMs - Effort
Crate must be moved from Tevatron location and installed at MC-1
Loss monitors must be installed in tunnel
Cables must be terminated
Filter box must be modified to reflect cable lengths
ACNet devices must be created
Possible minor modifications to readout software
Application software must be adapted

## Slide 9: Delivery Ring BLMs
The delivery ring BLMs will be the exact same system as the M4 line
60 loss monitors are required
3 readout electronics crates are required
These will be reused from the Tevatron
Effort is similar to M4 line, but 2-3 times as much due to more systems

## Slide 10: Summary
System design is finished
Low risk due to reuse of existing Tevatron loss monitor systems
Needed modifications are minor