## Slide 1: MICE Spectrometer Solenoid Update
Mark Palmer
FNAL
September 4, 2015

## Slide 2: Current Status: SSU
SSU
  A center coil problem was discovered during initial checkout in July
    Isolated
  Magnet warm-up was initiated
  As of July 28th, the issue had been identified as a cross-wiring of the Center Coil leads
    NOTE: 1st SS delivered to RAL and further feed-through and lead issues identified with second magnet
    Cross-wire occurred during retrofit of feed-throughs, HTS leads, and thermal connections on this magnet, which were carried out at RAL
    Magnet warm checks carried out successfully
  Magnet pumpout and cool-down initiated
    Magnet was cold on Sunday, August 23rd
    Magnet cold checks carried out successfully
  Training resumed on Wednesday, August 26th
  6 training quenches completed
    Maximum current attained: 235 Amps

## Slide 4: Current Status: SSD
SSD
  Magnet training carried out in July
  5 training quenches + 1 ramp-down quench completed
    Maximum current attained: 212 Amps
  Current issues:
    Missing voltage tap
      Need dI/dt signal to properly protect SC leads
    Ground fault circuit has a fault
      Need to diagnose

## Slide 5: SSD Training Quench

## Slide 6: Quench History
Currents shown are
for the Center Coil
of each magnet

## Slide 7: SSD Training Quench

## Slide 8: Power Supply System
Both SS magnets currently have training challenges due to the power supply system
  Cable resistance significantly increased when decision made to move power supplies to Rack Room 2
  Need larger cross-section cable to maintain PSU design specifications
  Remedies now being evaluated
    Making every effort to get the cooling channel operational before the end of the present ISIS run