## Slide 1: Cryo-Module test experience at DESY for the European XFEL
Denis Kostin, MHF-SL, DESY. 29.10.2015
LCLS-2 Cavity and Cryomodule Test Workshop 2015

## Slide 2: Outline
E-XFEL CryoModule Layout
CM Test Stand AMTF
CM Test Program and Procedures
CM Test Results and Data
Summary
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 3: 1.1 E-XFEL Module Layout
70K shield
Cavity type TESLA
Number of cavities 8
Cavity length 1.035 m
Operating frequency 1.3 GHz
R/Q 1030 Ω
Accelerating Gradient 20..31 MV/m
Quality factor 1010
Qext (input coupler) 4.6×106
Operating temperature 2 K
LHe cooling 2K / 31mbar
HERA cryoplant is used
12m
4K shield
input coupler
2 phase LHe pipe
He gas return pipe
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 4: 1.2 E-XFEL Module RF Power Coupler
warm part
waveguide part
RF power coupler
cold 70K window
cold part
warm 300K window
to cavity
cavity
actuator
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 5: 2.1 AMTF Overview
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 6: 2.2 AMTF Layout
2 cavity test stands (x4), 3 module test stands (5 MW klystron)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 7: 2.3 CM RF Test Diagram
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 8: 2.4 CM Test Waveguide RF Power Distribution
V. Katalev
adjustable RF power distribution
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 9: 2.5 СM RF Power Measurement
A.Goessel / DESY
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 10: 2.6 CM RF Conditioning / Data / Controls
A.Goessel / DESY
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 11: 3.1 CM Test Steps (1)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 12: 3.1 CM Test Steps (2)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 13: 3.1 CM Test Steps (3)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 14: 3.1 CM Test Steps (4)
CM Test Documents (EDMS):
CM Incoming Inspection
CM Final Test Report
CM Outgoing Inspection
Non-Conformity Reports
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 15: 3.2 CM Test Data
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 16: 3.3 CM Test Flow (remarks)
Optimized test procedure takes 13..15 days.
Delays are caused by non-conformities:
(cryogenic + vacuum) leaks (mainly at temporary connections to test stand)
warm coupler part
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 17: 3.4 CM Test Procedures
Procedure name | 
1 | Enabling of the Personal Interlock
2 | Disabling of the personal interlock
3 | Switching ON/OFF modulator and klystron. Locking modulator
4 | Check the hardware interlock thresholds, set technical interlock to warm mode
5 | Set technical interlock to cold mode
6 | Vacuum level check before start of coupler conditioning
7 | Warm input RF Coupler conditioning
8 | Coupler conditioning during cooldown
9 | Cavities fine tuning to 1.3 GHz using LLRF system
10 | Set the cavity input power limitation
11 | Qload, Kt calibration, Qload.HOM at 1.3Ghz with low RF Power
12 | Flat - top cavity measurement at 10 Hz repetition rate
13 | Cryomodule High RF Power and Heat Loads measurements
14 | Detune cavities
15 | LLRF check
16 | LLRF test
Defined and documented CM Test Procedures:
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 18: 3.5 CM Test Flow (improvements)
M. Wiencek / SRF2015
The most time consuming steps of the original CM test procedure: HOM spectra, reaching thermal equilibrium of “cold cables” for calibration and a long heat loads measurement on shields. Improved CM test procedure shortens the test time at 2K almost by a factor of two.
change
CM test time at 2K (cold test) [hr.]
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 19: 3.6 CM LLRF Test
LLRF relevant parameters
  Lorentz force detuning constant
  Microphonics behavior (resonances)
  Loaded Q variations with gradients
  Model: piezo diag -> detuning parameter (microphonics cancelation)
  8/9  & 7/9  mode of each cavity (-> controller!)
  Changes of QL with gradient (thermal coefficients)
  Calibration for piezo DC/AC (actuator & diag, switching)
  Calibration stepper motor & hysteresis
  QL range and opportunistic effects on cavity
Accept / reject accelerator module
  Lorentz force detuning exceeds normal value
  Extreme sensitivity to microphonics
  Abnormal behavior during fill/decay
  Sudden QL variation (multipacting in couplers)
  piezo / tuners not operable
J. Branlard
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 20: 3.7 RF Power Coupler Warm Conditioning Steps
step N | pulse length [µs] | RF power [kW] | time [hr] | comments
1 | 20 | 750 | 5..50 | activity onset 50kW
2 | 50 | 750 | 2..20 | 
3 | 100 | 750 | 2..20 | 
4 | 200 | 750 | 2..20 | 
5 | 400 | 750 | 2..20 | 
6 | 800 | 700 | 2..20 | 
7 | 1300 | 700 | 10..100 | 
8 | 1300 | 100..700 | 10 | RF power sweep
total |  |  | 50..200 | normally 2..5 days
Warm coupler conditioning at AMTF is done in 3 steps (with 5MW total RF power):
all 8 couplers are conditioned up to 450 kW;
couplers 1,2,3,4 are conditioned up to 750 kW;
couplers 5,6,7,8 are conditioned up to 750 kW.
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 21: 4.1 Warm Coupler Conditioning History
Module XM14 warm couplers conditioning at AMTF (XATB2):
0 2 7 12 17 22 27 32 37 42 47 52 57 hours
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 22: 4.2 Warm Coupler Conditioning Time
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 23: 4.3 CM Cold Test / Software Tools
cavity frequency tuning
flat-top pulse test at 2K
piezo tuning / operation
M. Wiencek / SRF2015
custom software tools for the optimized test flow
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 24: 4.4 CM Cold Test / Calibration
Module RF calibration measurement is done with LabVIEW test software at 1.3 GHz
 | Cav 1 | Cav 2 | Cav 3 | Cav 4 | Cav 5 | Cav 6 | Cav 7 | Cav 8
Qext [106] | 4.61 | 4.60 | 4.58 | 4.62 | 4.68 | 4.62 | 4.62 | 4.59
Kt [107] | 1.55 | 1.72 | 1.87 | 1.96 | 1.81 | 1.93 | 1.67 | 1.83
Qtrans [1011] | 2.51 | 3.07 | 3.64 | 3.98 | 3.42 | 3.88 | 2.92 | 3.49
QHOM1 | 6.3E+12 | 4.5E+12 | 2.8E+13 | 6.0E+12 | 4.3E+12 | 2.9E+12 | 2.0E+12 | 1.4E+12
QHOM2 | 6.2E+12 | 6.5E+13 | 4.3E+14 | 1.2E+12 | 2.5E+12 | 1.6E+13 | 6.8E+12 | 3.6E+14
Typical CM calibration data (module XM64)
Acceptance criteria:
QHOM1, QHOM2 > 1×1011
QHOM1, QHOM2 > Qtrans
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 25: 4.5 CM Cold Test Results
Typical module cold test data in final report (module XM61)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 26: 4.6 CM Performance / Gradient
M. Wiencek / SRF2015
31 MV/m power limit
23.6 MV/m
Notes: Before you start editing the slides of your talk change to the Master Slide view:
Menu button “View” > Slide Master:
Edit the following items:
1. On the Title slide (second master slide)
a) Click to add title of your talk
b) Click to add subtitle (conference, location, name of the speaker, date)
2. On the Slide Master (first master slide)
a) The 1st row in the violet header: Delete the existent text and write the title of your talk into this text field b) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area.
Close Master View

## Slide 27: 4.7 CM Performance / Quality Factor
M. Wiencek / SRF2015
1×1010
Notes: Before you start editing the slides of your talk change to the Master Slide view:
Menu button “View” > Slide Master:
Edit the following items:
1. On the Title slide (second master slide)
a) Click to add title of your talk
b) Click to add subtitle (conference, location, name of the speaker, date)
2. On the Slide Master (first master slide)
a) The 1st row in the violet header: Delete the existent text and write the title of your talk into this text field b) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area.
Close Master View

## Slide 28: 4.8 CM – Vertical Test comparison: max. Gradient
Ideal case VT:CM = 1:1
individual cavity comparison
(upper limit due to 31 MV/m limit in module test)
we lose between vertical and cryomodule test
average VT: (33.8) 30.5 MV/m (clipped at 31 MV/m) average CT: 28.4 MV/m (includes limit at 31MV/m)
D. Reschke / SRF2015

## Slide 29: 4.9 CM Test Statistics
M. Wiencek / SRF2015
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 30: 4.10 CM Cryogenic Measurements
area | static | dynamic | total | acceptance
 | static | RF 23.6 MV/m | static + total dynamic | static + total dynamic
40/80K | 97.96 W | not measured(1) | not measured(1) | 125 W
5/8 K | 10.93 W | 0.55 W(2,3) | 11.48 W | 16 W
2 K | 5.33 W(4) | 3.17 W(3) | 8.5 W(4) | 14 W
(1) Due to short measurement period, it is not possible to reach steady-state equilibrium value, and therefore to determine the dynamic heat load
(2) As RF operation during the dynamic heat load on 2K system was below 4 hours, it is not possible to reach steady-state equilibrium value, and therefore to determine the dynamic heat load
(3) RF power set to 21.7MV/m instead of 23.6MV/m
(4) Static heat loads for Feed-cap and End-cap (1.55 W) are already subtracted
Typical cryogenic losses data (module XM64)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 31: 4.11 CM Test Data Base
Oracle DB is set-up
to present and analyze
the CM test data
V. Gubarev
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 32: 5. Summary
E-XFEL CM Test at AMTF@DESY is an important step in CM production. It provides a critical data for later E-XFEL Linac operation as well as multiple CM quality checks.
Among many important CM Test steps the critical ones are:
  Warm input RF power coupler conditioning;
  Vacuum / Cryo-System leak tests at cold (2K);
  Single cavities performance tests at cold (2K);
  Auxiliaries (tuners/actuators, magnets etc.) tests at cold (2K);
Testing the cavities provides the final CM performance data and allows for individual adapted cavity RF power distribution.
Customized / adapted CM Test Field equipment is critical.
Customized / adapted CM Test software tools play an important role.
Taking CM test-stand into operation takes rather long time.
Optimized and well documented CM Test procedures are critical. CM Test flow and steps are to be adjusted / changed in operation.
CM Test data storage and analysis are of great importance.
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View