## Slide 1: Microphonics Experience at Cornell - A collection of results -
Matthias Liepe
Cornell University
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 1

## Slide 2: Outline
Microphonics results from
  CESR
  Cornell ERL Injector Cryomodule
  Cornell Main Linac Cavity HTC Testing
  HZB, Jlab FEL…
Conclusions
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 2

## Slide 3: CESR
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 3

## Slide 4: 500 MHz CESR SRF
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 4
0
0.2
0.4
0.6
0.8
1
-100
-50
0
50
100
time [sec]
detuning [Hz]
f  13 to 15 Hz
500 MHz cavities, operated at 4.2 K
Strong microphonics, broad spectrum

## Slide 5: CESR: Microphonics Spectrum
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 5
0
20
40
60
80
100
120
140
160
180
200
0
0.5
1
1.5
2
2.5
3
3.5
4
x 10
5
frequency [Hz]
fft amplitude [arb. units]
Microphonics spectrum
0
10
20
30
40
50
60
70
80
90
100
0
500
1000
1500
2000
2500
amplitude [arb. units]
frequency [Hz]
Mechanical resonances
Low frequency mechanical resonances
Resonance frequencies change with time and vary from cavity to cavity

## Slide 6: Cornell ERL Injector Cryomodule
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 6

## Slide 7: 1.3 GHz ERL Injector Cryomodule
Refrigeration
transfer lines
Waveguide
from klystron
Beam exit
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 7
Number of 2-cell cavities 5
Acceleration per cavity 1 – 3 MeV
Accelerating gradient 4.3 – 13.0 MV/m
R/Q (linac definition) 222 Ohm
Qext 4.6104 – 4.1105
Total 2K / 5K / 80K loads: 30W / 60W / 700W
Number of HOM loads 6
HOM power per cavity 40 W
Couplers per cavity 2
RF power per cavity 120 kW
Amplitude/phase stability 10-4 / 0.1 (rms)
ICM length 5 m

## Slide 8: ERL Injector: Microphonics Bursts
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 8
Significant changes over time
Step impulses at ~second rep-rate!?
After improvements

## Slide 9: ERL Injector: Source of Microphonics Bursts
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 9
Source of “bursts”:
Thermo-acoustic oscillations in
Warm Up – Cool Down plumbing
Closed end (valve) creates dead end with trapped gas

## Slide 10: Microphonics Bursts Studies
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 10
Warm up – cool down valve closed
Warm up – cool down valve slightly open

## Slide 11: 1.3 GHz ERL Injector Cryomodule: Microphonics
Vibration frequency [Hz]
Significant differences between cavities
f = 3 to 8 Hz
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 7

## Slide 12: ERL Injector: Microphonics Histogram
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 12
Graph shows histogram of 30 million samples over a period of 1 hour
fmax < 5*f
Measured during “step-impulse free time”

## Slide 13: Mechanical Coupling Characterization Measurements with a Modal Shaker
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 13

## Slide 14: Measurements with a Modal Shaker
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 14
Shaker on module support
Shaker on module top (HGRP support)
Vibration frequ. [Hz]
Vibration frequency [Hz]
Response ampl. [arb. units]
Response ampl. [arb. units]

## Slide 15: Mechanical Coupling Characterization
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 15
Excitation Point | Excitation Force | Detectable With Cavity Accelerometer | Detectable On Cavity RF Frequency (>0.1Hz modul.)
Coupler Waveguide | 110 N (25 lbs) | No | No
Coupler | 110 N (25 lbs) | No | No
Cryomodule Saw-Horse Support | 110 N (25 lbs) | Yes | No
Helium Gas Return Pipe Support | 110 N (25 lbs) | Yes | Yes
Beam Line | 10 N (2 lbs) | No | No
Helium Supply/Return | 110 N (25 lbs) | No | No
Ground vibrations and other mechanical vibrations do not strongly couple to the SRF cavities!
Main contribution to cavity microphonics comes from fast fluctuations in the He-pressure and vibrations in the cryogenic system

## Slide 16: Active Microphonics Control Studies with Fast Piezoelectric Frequency Tuner
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 16
Vibration frequ. [Hz]
phase
amplitude
Cornell ERL injector cavities
Microphonics feedback challenging because of complex transfer function piezo  f with mechanical resonances
Linearity and hysteresis issues

## Slide 17: ERL Injector Frequency Tuner
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 17
Modification of the INFN blade tuner
Tuning range:
  500 kHz stepper
  500 Hz piezo
Added piezos for microphonics compensation (R&D)

## Slide 18: Frequency Tuner Studies
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 18
Stepping motor driven tuner: not suitable for fine frequency control!
Piezo driven tuner: suitable for fine frequency control!

## Slide 19: Frequency: Piezo-Tuner Hysteresis
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 19
Piezo Hysteresis Curve: DC frequency shift vs. applied voltage to all 4 piezos:

## Slide 20: Active Microphonics Compensation
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 20
Piezo Feedback on Tuning Angle/Cavity Frequency:
 Reduces rms microphonics by up to 70%!

## Slide 21: Active Microphonics Compensation
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 21
Feedback off
Feedback on
Detuning Histogram: ~30% reduction in peak detuning

## Slide 22: Active Microphonics Compensation
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 22

## Slide 23: Cornell Main Linac Cavity HTC
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 23

## Slide 24: Cornell HTC Test of the Cornell ERL Main Linac Cavity
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 24
cavity
HOM load
HOM load
HGRP
80K shield
Gate valve
Test cryomodule installed at Wilson Lab
Main linac system test in one-cavity test module
  1st test: cavity and tuner only
  2nd test: add high power RF input
  3rd test: add HOM beamline loads

## Slide 25: Main Linac Cavity Mechanical Design
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 25
Main linac prototype module has
3 cavities without stiffening ring
3 cavities with stiffening ring at ~70% iris-equator distance

## Slide 26: Main Linac Cavity: Measured Detuning Pressure Sensitivity
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 26
df/dp as a function of tuner motor steps

## Slide 27: Main Linac Cavity: Detuning Pressure Sensitivity
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 27
df/dp as a function of tuner motor steps
Original model for production cavity, deformed by helium pressure
Model in which the welds are modelled more realistically

## Slide 28: Main Linac Cavity: Tuner Studies
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 28

## Slide 29: Main Linac Cavity: Vibration and Microphonics Studies
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 29
FFT of cavity detuning signal shows 60 Hz
harmonics, but nothing indicative of pumps

## Slide 30: Main Linac Cavity: Mechanical Resonances in HTC-1
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 30
Resonances agree fairly with ANSYS predictions for cavity in tank
Several additional resonances in spectrum

## Slide 31: Main Linac Cavity: Microphonics in HTC-1
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 31
peak detuning < 20 Hz
RMS detuning ~5 Hz

## Slide 32: Main Linac Cavity: Microphonics in HTC-3
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 32
1.6K
Std: 7.1 HZ
1.8K
Std: 9.6 Hz
2.0K
Std: 8.6 Hz
FFT shows line @
pump frequency

## Slide 33: HZB, Jlab FEL…
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 33

## Slide 34: 9-cell Test in HZB Hobicat With Cornell LLRF System and Saclay Tuner
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 34
Piezo to Detuning Transfer Function
Microphonics Spectrum
Strong mechanical resonances
RMS detuning of 4 Hz with 15 Hz peak detuning

## Slide 35: JLAB Upgrade Cryomodule Test With Cornell LLRF System
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 35
-6
-4
-2
0
2
4
0
1000
2000
3000
4000
5000
cavity detuning [Hz]
# of measurements
JLab FEL

## Slide 36: JLAB Upgrade Cryomodule Test With Cornell LLRF System
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 36
Automated resonance hunting
time [minutes]
cavity detuned by 30 kHz (400 bandwidths)
cavity on resonance
acc. field [MV/m]
0
4

## Slide 37: JLAB Upgrade Cryomodule Test With Cornell LLRF System
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 37
cavity field
constant forward (klystron) power ( 210 W)
[MV/m , kW]
piezo feedback/feedforward off
piezo feedback/feedforward on
time [seconds]

## Slide 38: JLAB Upgrade Cryomodule Test With Cornell LLRF System
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 38
0
0.2
0.4
0.6
0.8
1
0
5
10
accelerating field [MV/m]
time [sec]
 150 Hz Lorentz-force detuning (compensated by piezo feedforward ), cavity half bandwidth = 6 Hz !
15
Start-up: Field Ramp at QL = 1.2·108

## Slide 39: JLAB Upgrade Cryomodule Test With Cornell LLRF System
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 39
0
0.2
0.4
0.6
0.8
1
-5000
-4000
-3000
-2000
-1000
0
time [sec]
piezo drive signal
[arb. units]
Piezo drive signal to compensate
Lorentz-force detuning
cavity filling
0
0.2
0.4
0.6
0.8
1
-30
-20
-10
0
10
20
detuning [Hz]
time [sec]
Lorentz-force detuning without compensation: 150 Hz
remaining microphonics
cavity half bandwidth: 6 Hz

## Slide 40: Cornell HTC with LCLS-II Prototype Cavity (measurements by FNAL team)
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 40

## Slide 41: Conclusions
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 41

## Slide 42: Conclusions
  Microphonics levels vary significantly from cavity to cavity
  Microphonics spectra are often complex with many lines, making active compensation more challenging
  Ground vibrations might not be the main driving source, but the He-system might be
  Passive compensation/mechanical design for low microphonics can be powerful, but not always possible
  Active compensation feasible within limits, but challenging. Success strongly depends on actual microphonics spectrum
      Easier to reduce RMS then peak detuning value!
  Best Cornell results are ~20 Hz peak detuning for long periods of time (hours)
  Much more work needed to understand dominant sources of microphonics, how to reduce them, and to improve active compensation
M. Liepe Microphonics Workshop, FNAL October 8-9 Slide 42

## Slide 43: Thank you for your attention!