## Slide 1: Cryomodule Safety Framework
Jay Theilacker
LCLS-II 3.9 GHz Cryomodule Preliminary Design Review
November 20, 2015

## Slide 2: Introduction
Fermilab Safety Program
Multi-Laboratory Safety Program Comparison
Pressure Safety
Electrical Safety
Seismic Safety
FMEA / What-If Analysis
Summary

## Slide 3: Design Basis
Derived from the 3.9 GHz cryomodule design supplied by Fermilab to DESY FLASH, most recently the 3.9 GHz INFN/European-XFEL and 1.3 GHz LCLS-II cryomodule design
Design modified at Fermilab to satisfy LCLS-II project requirements using Fermilab Engineering, QA and ES&H Manuals
It is also necessary to ensure that Fermilab policies satisfy the requirements of SLAC

## Slide 4: Fermilab Safety Program
Fermilab ES&H Manual (FESHM) applied to the cryomodule
Chapter | Title | LCLS-II Application
5031.1 | Piping Systems | Cryomodules and cryogenic distribution system piping
5031.6 | Dressed Niobium SRF Cavity Pressure Safety | SRF cavities
5034 | Pressure Testing | Cryomodules and cryogenic distribution system.
5033 | Vacuum Vessel Safety | Cryomodules and cryogenic distribution system.
4240 | Oxygen Deficiency Hazard | Formerly 5064. Input for SLAC ODH analysis.
5032 | Cryogenic System Review | Lists documents we generate to receive an operational readiness clearance. Applies to cryomodules and cryogenic distribution system. First application will be CMTS.
9100 | Fermilab Electrical Safety Program | Cryomodules in general
9140 | Protection Against Exposed Electrical Bus | Cryomodule magnet power leads
9150 | High Voltage Coaxial Connectors | Cryomodule coupler vacuum ion pump
9160 | Low Voltage, High Current Power Distribution Systems | Cryomodule magnet and power leads
10110 | Below the Hook Lifting Devices | In case SLAC chooses to use a Fermilab lift fixture engineering design

## Slide 5: Multi-Laboratory Safety Program Comparison
SLAC reviewed the safety and QA programs at the partner laboratories
  Common areas determined to be consistent with the SLAC program and therefore the LCLS-II project requirements
  Unique areas were addressed
    Seismic (FNAL)
    ODH (SLAC)

## Slide 6: Pressure Safety
Partner labs (FNAL and JLAB) have pressure vessel, pressure piping and vacuum vessel policies in place
  Reviewed to be equivalent to SLAC policies and LCLS-II Project requirements [LCLSII-1.2-EN-0020]
Special SRF pressure vessel considerations
  Although the 3.9 GHz dressed cavity is outside the scope of ASME B&PV Code (< 6”), we will still treat it as a pressure vessel using the same process as the 1.3 GHz dressed cavity
  Unable to comply directly with ASME B&PV Code
  Governing regulation becomes 10 CFR 851
  “When national consensus codes are not applicable”...“contractors must implement measures to provide equivalent protection and ensure a level of safety greater than or equal to the level of protection afforded by the ASME”

## Slide 7: SRF Pressure Safety
Fermilab equivalent level of safety method
    Chose to address the pressure safety at the source (the cavity) as opposed to containment at the vacuum vessel boundary
    Fermilab Dressed Niobium SRF Cavity Pressure Safety
      ES&H Manual Chapter 5031.6
      References “Guidelines for the Design, Fabrication, Testing and Installation of SRF Niobium Cavities”
    Using the Fermilab method made the most sense since the cavity was already designed using this method

## Slide 8: Pressure Safety Review Differences
FNAL and SLAC review pressure systems differently
Both systems have advantages and disadvantages
Project is working toward meeting requirements of both systems
 | Fermilab | SLAC
Scope of Review | Individual components Overall system | Overall system
Time of Review | After implementation (ORC) | After design, before implementation
Code Stamped Vessels | Considers stamp to be proof of design | Requires vendor calculations, material certifications and weld/welder information

## Slide 9: Pressure System Review
Document Document #
1. Cryomodule Piping Engineering Note 3.9 TBD (1.3 EN01803)
2. Cryomodule Cavity Engineering Note 3.9 TBD (1.3 EN01774)
3. Cryomodule Vacuum Vessel 3.9 TBD (1.3 ED0002339)
4. LCLS-II CDS Relief System Requirements LCLSII-4.9-EN-0300
5. LCLS-II CDS Gas Return Pipe Pressure Drop Analysis LCLSII-4.9-EN-0292
6. LCLS-II CDS HT Shield and LT Intercept ∆P Analysis LCLSII-4.9-EN-0295

## Slide 10: LCLS-II Linac Pressure Systems
Pressure Systems Scope Design Documents
2K Return SRF Cavities, Lines B, G and H 1, 2, 4, 5
2K Supply Line A 1, 2
Low Temperature Intercept Lines C and D 1, 2, 6
High Temperature Shield Lines E and F 1, 2, 6
Insulating Vacuum 3
Status
Cavity design and implementation
Design phase
LCLS-II Linac pressure systems
Documents 1. – 3. will be reviewed internally at FNAL and then submitted to the SLAC Pressure System Program Manager for review and approval
Documents 4. – 6. have been approved by the SLAC Pressure System Program Manager
Cryomodule Operational Readiness Clearance
Will be reviewed internally at FNAL prior to cryomodule testing

## Slide 11: Electrical Safety
The cryomodule electrical systems have been design in accordance to the Fermilab ES&H Manual.
Electrical systems include:
  System FESHM Chapter
  Thermometry 9100
  Coupler vacuum ion pump 9100, 9150
  Coarse tuner motor 9100
  Fine tuner piezo actuators 9100
SLAC has a LCLS-II project technical standard for electrical safety requirements [LCLSII-1.1-TS-0435] which is consistent with the Fermilab policies
9100 | Fermilab Electrical Safety Program
9150 | High Voltage Coaxial Connectors

## Slide 12: Seismic Safety
Requirements
  LCLSII-4.5-EN-0226
    Cryomodule Seismic Design Criteria
Analysis
  A separate 3.9 GHz cryomodule seismic analysis will be very similar to the 1.3 GHz cryomodule and the cavity mass is considerably less.
  LCLSII-4.5-EN-0430
    Stress Analysis of LCLS-II 1.3 GHz Cryomodule for Seismic Load
Status
  A revision of the 1.3 GHz analysis is being finalized. It will be reviewed by SLAC to determined that it meets SLAC seismic requirements.

## Slide 13: FMEA / What-If Analysis
The direct interconnection of the cryomodule and cryogenic distribution system made it prudent to generate combined FMEA and What-If documents.
  Requirements for Operational Readiness Clearance
  FESHM 5032 Cryogenic System Review [for FNAL]
  LCLSII-1.1-TS-0382 Engineering Failure Analysis Procedure
FMEA
  LCLSII-4.5-EN-0255
    CDS and Cryomodule Failure Mode and Effects Analysis
What-If
  LCLSII-4.5-EN-0253
    CDS and Cryomodule What-If Analysis

## Slide 14: ODH / What-If Analysis
  What-If analysis is used to identify maximum credible incidents. Two different types of maximum credible incidents are determined and mitigated.
  Internal cryogenic rupture
  Mitigated by vacuum relief sizing [ED0002396]
  Input to tunnel and gallery ODH analysis [LCLSII-1.2-PM-0379]
  Loss of vacuum (beam or cryostat) to air
  Mitigated by circuit relieving [LCLSII-4.9-EN-0300]
  The CDS/Cryomodule What-If Analysis is an input to the SLAC tunnel and gallery oxygen deficiency hazard analysis. The CDS/Cryomodule groups have been supplying the SLAC personnel working on the ODH analysis with the required inputs. A preliminary ODH analysis has been prepared by SLAC.

## Slide 15: Summary
  Partner Lab Engineering, Pressure Safety and QA processes have been reviewed and approved for use on the LCLS-II project
  1.3 GHz Cryomodule and CDS pressure systems have been reviewed and approved by the SLAC Pressure System Program Manager. Additional documents specific to 3.9 GHz cryomodule will be reviewed at the FDR
  LCLS-II dressed cavities will be reviewed and approved at FNAL prior to testing
  Electrical safety being addressed through existing policies
  Seismic analysis will be under final review by the SLAC Seismic Program Manager and Building Inspection Office shortly for the 1.3 GHz cryomodule. By design, the seismic stresses in the 3.9 GHz cryomodule will be less
  Cryomodule/CDS FMEA and What-If analyses are developed, identifying maximum credible incidents and tied into the SLAC tunnel/gallery ODH analysis