## Slide 1: 35 Ton Run 2 Operation and Experiences
Fritz Schwartz
October 26, 2016

## Slide 2: Outline
Background information
  Overview of system and key components
Operating modes of 35T
  Piston Purge
  Cooldown & Filling
Issues seen during operation
  Liquid Pump Seizure
  Power outage
  Loss of LN2 cooling to condensers
  Vapor pump failure leading to gross contamination of liquid argon
Future plans for 35T system

## Slide 3: Background Information
Prototype Cryostat at Fermilab to show the suitability of the membrane technology for LAr detectors
Constructed in 2012 in a decommissioned proton beam line, adjacent to the existing Liquid Argon Purity Demonstrator (LAPD)
Reuse the existing purification and instrumentation infrastructure of LAPD
Run 1 of 35T was a successful demonstration of techniques of piston purge, cooldown and purification in the membrane cryostat of rectangular shape and warm top plate without TPC (detector)
Run 2 was initiated in 2015 to test operations of the newly installed TPC
Filtration stopped on March 19, 2016 after warm pump failed resulting in intake of atmospheric air and contaminating 35T LAr inventory

## Slide 4: Cryogenic System Overview
System as installed in PC4 facility

## Slide 5: Key System Components
Cryostats
Liquid Argon Pumps
Condensers
Filtration
Purity Monitoring

## Slide 6: Cryostats

## Slide 7: Liquid Argon Pumps
ACD submersible AC32 pumps

## Slide 8: Condenser Design
Coiled tubing circulating liquid nitrogen
Seamless tubing with connections outside vessel to decrease possibility of leak into cryostat argon
Control valves regulate the flow to each coil depending upon heat load

## Slide 9: Filtration
Water filter
  ~80 liters 4A porous sphere molecular sieve
  Regenerated using heated argon gas, followed by evacuation
  Upstream of oxygen filter because oxygen filter also removes water
Oxygen filter
  ~80 liters of copper catalyst
  A thin layer of copper on a high surface area alumina substrate
  Regenerated using a heated mixture of hydrogen/argon gas which removes oxygen from the copper to form water, followed by evacuation
  Capacity at LAr temperature is approximately 0.5g O2 per kg of filter material
Particulates
  Sintered metal filters, large surface areas
  Used to protect other equipment from debris

## Slide 10: Water & Oxygen Filter Design

## Slide 11: Purity Monitoring
Commercial/Laboratory Gas Analyzers
  Nitrogen
  Oxygen (3 with multiple ranges)
  Water (Rough dew point and higher precision)
Purity Monitor
  Creates an electron drift using xenon flash lamp and photo cathode
  In-line and cryostat installations

## Slide 12: Outline
Background information
  Overview of system and key components
Operating modes of 35T
  Piston Purge
  Cooldown & Filling
Issues seen during operation
  Liquid Pump Seizure
  Power outage
  Loss of LN2 cooling to condensers
  Vapor pump failure leading to gross contamination of liquid argon
Future plans for 35T system

## Slide 13: 35T Piston Purge & Recirculation
Using higher density of argon to push out other gases from bottom to top
Vertical velocity of 1.2 meters per hour determined best to reduce mixing

## Slide 14: 35T Piston Purge & Recirculation
H2O = 25 ppm
N2 = 18 ppm
O2 = 0.2 ppm
35T Piston Purge
35T Recirculation and Filtering
O2%, Instrument sampling other source

## Slide 15: Cooldown Sprayers
Mix liquid and gas argon
Additional nozzles provide necessary momentum to circulate flow around volume of cryostat
Design rate of 10K/hr supplied by cryostat manufacturer

## Slide 16: Cooldown and Fill of 35T
LAPD Pump trip
Transfer from
LAPD
GAr transfer from LAPD (over weekend) ~1.5 tons
Fill from
LAr Trailer
Membrane
Temperature
sensors
Gas Supply Issue

## Slide 17: Results of Run 2
Several issues occurred during this run – did provide valuable lessons for future systems
PrM2 & 3 lifetimes
Sitewide Power Outage
LN2 Cooling Loss
Tubing Break
Pump Start
3 ms lifetime
Failures

## Slide 18: Outline
Background information
  Overview of system and key components
Operating modes of 35T
  Piston Purge
  Cooldown & Filling
Issues seen during operation
  Liquid Pump Seizure
  Power outage
  Loss of LN2 cooling to condensers
  Vapor pump failure leading to gross contamination of liquid argon
Future plans for 35T system

## Slide 19: Liquid Pump Seizure
With 35T cryostat completely filled, the submersible pumps were seized up and blowing fuses
Pumps were rebuilt after Run 1, suspected that they were not tested prior to trying to operate them now
Controls group bypassed fast acting fuses with disconnect fuses and installed simple ON/OFF switch to bump motor several times (momentarily ON, 1 minute OFF)
The pump finally broke free after 12 repetitions
Restored original wiring once pumps were freed

## Slide 20: Power Outage
The power outage provided an opportunity to identify purity gradients within the liquid
Pump Off

## Slide 21: Power Outage
Small thermal gradients were also noticed

## Slide 22: CFD Analysis (Erik Voirin)
Suspected that the liquid return was the primary cause
CFD analysis performed to confirm
Liquid Return
Pump Suction

## Slide 23: Loss of LN2 Cooling
Loose valve plug on LN2 supply line
DF-560 O2 Analyzer
RTD Temp. Probes
Cryostat Internal Gauge Pressure
Condenser LN2 Level
Cryostat LAr Level
LN2 runs dry, lose control of internal pressure. Vents at 1.5 psi, LAr Level begins to drop

## Slide 24: Vapor Pump Failure
Pump Inlet
From Cryostat

## Slide 25: Vapor Pump Failure
Gas analyzers sampling from ullage space
Filters saturate, see response from analyzers within 20 minutes
Complete loss within 30 minutes…ending the run

## Slide 26: Recommendations to Reduce Major Contamination
A panel was formed to investigate the failure of the pump
Conclusion was that design was done to acceptable codes and standards, also considering the short expected runtime
Panel suggested that analyzers be installed at locations where contamination can be caught prior to entering cryostat
Location is likely to be at liquid return line to cryostat where nitrogen (or oxygen) analyzer can be utilized
Upon detection of impurities, controls need to isolate cryostats – total shutdown mode

## Slide 27: Outline
Background information
  Overview of system and key components
Operating modes of 35T
  Piston Purge
  Cooldown & Filling
Issues seen during operation
  Liquid Pump Seizure
  Power outage
  Loss of LN2 cooling to condensers
  Vapor pump failure leading to gross contamination of liquid argon
Future plans for 35T system

## Slide 28: Future Plans of 35T
Two runs planned for 2016-2017, with preparation work being performed now
  High voltage test of ProtoDUNE style field cage assembly
  Beam plug test to investigate effects of reducing amount of liquid argon between detector and cryostat wall
Modifications to liquid argon return line in 35T cryostat
  Investigate effects on the purity gradient issue
Addition of heating elements in 35T cryostat to increase boil-off between runs, also useful for boil-off of pooling liquid during the spray cooling phase

## Slide 29
Questions?