## Slide 1: Cryogenics and Plan for Instrumentation
Alan A. Hahn
FNAL

## Slide 2: Outline
Intro to PC4
Cryo Schedule/Timeline
  Filling
  Running
  Emptying
Cryo Instrumentation
  Purity Monitors
  “Precision” Temperature Measurement
  Gas Analyzers
  “Standard” Cryo Devices

## Slide 3: 35-ton Prototype and LAPD Located in PC4, former Fixed Target Proton Beamline
35-ton Cryostat uses
  the LAPD Filters
  Gas Analyzers
  Instrumentation
LAr can be shifted back & forth between LAPD & 35-ton
  However there is a “tax”
    ~10% from LAPD->35 Ton
    ~30% from 35T -> LAPD
    Function of pump head of two cryostats

## Slide 4: 35T Phase 2 cryo timeline
76 days
Want to minimize the time between installation of apparatus and HV Commissioning

## Slide 5: 35T Cryo Timeline
35T Phase 2
  From Detector Installation Finish to Running– 76 days
  “Running” = start of HV Commissioning
  At least one reason is that LAPD Commissioning started after the the Detector Installation and not before.
This time want LAPD filled and ready at time of Apparatus installation
  Want Gas operation permit of 35 Ton (for purging and recirculation also on hand at/near time of installation.

## Slide 6: Stage 1 Cryo Run Timeline

## Slide 7: Length of HV Run Period
Assume HV Commissioning starts as soon as Cryostat is filled and pump is recirculating LAr to filters.
It takes ~2 weeks to get to a e- lifetime of ~2-3 ms.
  See next slide.
Assume one week at Design Field & 2-3 ms e- lifetime (impurities ~100 ppt) demonstrates “Holding” HV.
Adding extra week (or so) for contingency or higher HV fields gives ~1 month.
Length of Stage 1 run is limited by desire to get Stage 2 run finished in a “timely” manner

## Slide 8: e- Lifetime over Phase 2 Run
PrM2 & 3, both long PrMs at midway heights in Cryostat.
Sitewide Power Outage
LN2 Cooling Loss
Tubing Break
Pump Start
3 ms lifetime
~ 2 weeks
~ 3 days
Pump Restart

## Slide 9: 35T Emptying Timeline
Reduction in time due to
  No need to vaporize LAr as in last run– pump directly to LAPD
  Plan to add electric heaters (6 kW) to boil off remaining LAr
    ~2-3x faster than bubbling N2 as done before
  Combine warmup (after LAr gone) with dry air purge
    Need Cryo temps in 200 K range first.
Last Run

## Slide 10: Cryo Instrumentation

## Slide 11: Purity Monitors
Cryostat: String of 4 PrMs, Short-Long—Long-Short
Inline –1 Long PrM mounted after filters
  Always measures extremely Pure LAr
PrM 1
PrM 2
PrM 3
PrM 4
RTDs
Inline PrM
Flow from filters
Flow back to 35 T
E.Voiron Drawing

## Slide 12: Precision Temperature Measurement
RTD Readout—Lakeshore 218
  Have achieved ~ 2 mK rms resolution (statistical) on 35T Phase 2 RTD readout.
  These are the RTDs shown in previous slide.

## Slide 13: Gas Analyzers
Can “sniff” liquid or gas regions of cryostat
  O2, N2, H2O units, with sensitivities from % to sub-ppb levels
  Useful for both debugging during initial phase of gas purging and during the run to help identify contamination

## Slide 14: More Cryo Instrumentation
RTDs, Pressure Transducers, LAr Level Indicators…
Monitored and archived by the Cryo Controls System

## Slide 15: Other Cryo Motivation for this Run
35T Phase 2 (and perhaps Phase 1) showed that the Purity in Cryostat was vertically stratified.
  See slides 12 and 18
  Erik Voiron theorizes this is due to the filtered highly pure LAr that is returned to the Cryostat is slightly cooler than the bulk LAr.
    Seems to be backed up by PrM mounted RTDs (slide 12)
  LAr is being introduced at bottom of Cryostat and tends to stay there since it is cooler
  The pump intake is also at the bottom.
    We are filtering the LAr that is already pretty pure.
We plan to alter the return point to be near the surface layer of the Cryostat
  See if this avoids the previous stratification by better mixing the return LAr into the bulk LAr.
  Plan to place a “Precision” RTD near the return manifold to sample the temperature.
  As a side, with better mixing we might improve the overall purity of LAr.
See Erik’s DuneDocDB1156 for more info

## Slide 16: Xtra slides

## Slide 17: Phase 2: 35T Cooldown and Filling
LAPD Pump trip
LAr Xfer
from LAPD
GAr xfer from LAPD (over weekend) ~1.5 tons
Top-Off from TRAILER
Membrane
Temperature
sensors
~60 hrs
Cooldown rate
Mfg Specs <~ 10K/hr
~16hrs

## Slide 18: Immediately Notice Purity Stratification after pump start and purification
PrM 1 (near Cryostat bottom)
PrM 4 (near surface)
O2 Gas Analyzer, ~(~80 cm from bottom)