## Slide 1: ADC Transmission
Mauricio Villa
GEM Fellow, Duke University
16 August, 2016

## Slide 2: Outline
ADC Introduction
ADC Transmission Options
  BRECOFlex Belts
  NEXEN Roller Pinion Gear
Conceptual Design of Motor Platform
Questions

## Slide 3: ADC Introduction
DESI -- Dark Energy Spectrographic Instrument
  Mayall Telescope at Kitt Peak, AZ
ADC -- Atmospheric Dispersion Corrector lens
  There are 2 lenses used to correct for light dispersion caused by the atmosphere

## Slide 4: ADC Introduction: Specifications
Bearing
Bearing
Lens
Lens
Harmonic Drive

## Slide 5: ADC Introduction: Specifications
Description | Value
Rotational Load (ADC lens, cell, bearing and fasteners) | Weights = 214 kg (472LBS) wr^2 = 40.8 kg-m^2 (968 Lb-ft^2)
Starting Torque | 100 Nm (885 LB-in)
Running Torque | 80 Nm (708 LB-in)
Driven Gear Rotation Speed range | 0 to 10 degrees/sec (0 to 1.7 rpm)
Rump-up/down | Max 3 seconds
Lifetime | 155,000 duty cycles over a 5 year period
Idle time | months
Accuracy on the rotation | +/- 0.1 degree
Tangential Backlash | NONE if possible
Gear centers distance | 584mm +3mm/0 or (23.000” +0.125”/0)
Driven gear addendum diameter | 951mm 0/-3mm or (37.450” +/-0.125”)
Material | Steel/Plastic
Protection from corrosion | Required
Minimum safety factor against Yield, fracture or buckling | 4.0
Optimization of weights | required
Maintenance/cleaning of gears from outside of barrel | No lubrication is HIGLY preferable

## Slide 6: Outline
ADC Introduction
ADC Transmission Options
  BRECOFlex Belts
  NEXEN Roller Pinion Gear
Conceptual Design of Motor Platform
Questions

## Slide 7: Mating specifications for Pinion and Gear
From HarmonicDrive SHA series manual
  SHA25A-SG : bolt circle, bolt sizing and tolerances
From KAYDON Spec Drawing
  DWG 19691A01: bolt circle, bolt sizing and tolerances

## Slide 8: BRECOFlex Transmission System Components

## Slide 9: BRECOFlex Considerations
Allowable Tensile Strength
Tooth Shear Strength
Belt stiffness
  Backlash, Slipping teeth
  Pretension
Space constraints
  Size of idlers
  Width of belt

## Slide 10: BRECO Flex Drive: ADC layout Configuration based on the std sizes.
allow for a belt length of 3350mm

## Slide 11: BRECO Flex Drive: ADC layout Configuration

## Slide 12: NEXEN Transmission System

## Slide 13: NEXEN Transmission System
NEXEN catalog options do not satisfy our size requirements.
Worked with NEXEN to develop a custom design

## Slide 15: NEXEN and BRECOflex General Comparisons
NEXEN
  Advantages
    No middle section changes
    High positional accuracy (up to 30 µm)
    “Near zero backlash”
    Easy to assembly and to do maintenance
    Pinions can be replaced easily
    Lifetime of 200 years of operation
  Disadvantages
    Gear will be done in segments
      Addressed by reconfiguring oil can.
    Rely on a specific company
BRECOflex
  Advantages
    No middle section changes
    Multiple manufacturers
  Disadvantages
    Belt replacement poses challenges
    No solid information on
      Re-tensioning frequency
      Belt lifetime
    More components
      Installation more difficult

## Slide 16
NEXEN |  |  |  | 
Number of Sets | Price Per set | Total Price wo extra pinion | Price Extra Pinion per Unit | Total for this item ($k)
2 |  |  |  | 
3 |  |  |  | 
10-12 week delivery |  |  |  | 
Fermilab Designer – ~2 weeks (1 FTE) |  |  |  | 
Two motor and two controllers |  |  |  | 
Total |  |  |  | 
BRECO flex |  |  |  | 
Number of Sets | Price Per set | Total Price wo extra pinion | Price Extra Belt | Total for this item ($k)
3 |  |  |  | 
8 week delivery |  |  |  | 
Fermilab Designer – ~6 weeks (1 FTE) |  |  |  | 
Three motor and two controllers |  |  |  | 
Total |  |  |  | 
Prices are confidential!
Costs for these system are about that for a luxury car, with a difference of about $15-20k

## Slide 17: Outline
ADC Introduction
ADC Transmission Options
  BRECOFlex Belts
  NEXEN Roller Pinion Gear
Conceptual Design of Motor Platform
Questions

## Slide 18: Definitions
Need to align Motor/Pinion within parallelism specifications required by Nexen
(see “How to align Nexen gear and pinion system” from Aug 1,2016 )
Within this presentation, the “Fixed plate” refers to the block already welded to the Middle Section

## Slide 19: Conceptual Idea #1: Pros and Cons
Pros:
  Mating with Fixed plate is straightforward
  Plenty of space for connecter bolts and precision pins
Cons:
  Structural strength/rigidity a concern if tack welding shims/Blue plate/Yellow bracket
  Aligning Pitch and Roll simultaneously be distributing shims may prove time consuming/difficult

## Slide 20: Conceptual Idea #2: Pros and Cons
Pros:
  Improved structural rigidity between Yellow bracket and Blue Plate
  Adjusting Pitch and Roll simple with bolt method
Cons:
  Need to reduce length and width of blue plate (and Yellow bracket) to facilitate tack welding the blocks*
  Determining height of blocks will require a potentially iterative process

## Slide 21: QUESTIONS?
Special Thanks to…
Giuseppe Gallo and Gaston Gutierrez
Sandra Charles and Judy Nunez
The National GEM Consortium
Joe Howell and Dave Pushka
Skills Learned this summer
CAD modeling proficiency improved
Understanding of mechanical design of precision parts improved
Professional skills required in working with engineers improved
Learned how to make not-soggy (instant) oatmeal

## Slide 22: http://www.gemfellowship.org/
The National GEM Fellowship…
  For underrepresented US citizens in STEM
  PAYS FOR GRADUATE SCHOOL
  Apply during your senior year of undergrad
  Valuable internships with Employer Sponsors
  PAYS FOR GRADUATE SCHOOL
  Fellowships for both Masters and PhD students
  Living stipend during the school year
  Your Fermilab network has plenty of GEM people you can talk to!
  (Sandra Charles, Maurice Ball, Charlie Orozco, Chris Cameron, Miguel Marchan, Me and many other alums/employees)
  mauricio.villa@duke.edu