## Slide 1
Wall Current Monitor (WCM) and Fast Current Transformer (FCT)
at J-PARC Main Ring
Hironori Kuboki
Takeshi Toyama, Dai Arakawa (PF), Koutoku Hanamura,
Yoshinori Hashimoto, Masashi Okada, Ken-ichiro Satou,
Bruce Yee Rendon, Masaki Tejima
MR Monitor Gr.

## Slide 2
WCMs and FCTs at J-PARC Main Ring
WCM for feedforward (2015)
FCT for abort beam line (2016)
Summary & prospects
Wall Current Monitor (WCM) and Fast Current Transformer (FCT)
at J-PARC Main Ring

## Slide 3
Main Ring (MR)
500 m
Injection
from RCS
Fast extraction (FX)
Slow extraction
(SX)
RF cavity
WCMs | 1 (measurement) 2 (feedforward) | 
FCTs | 3-50BT | 5
 | MR | 6+1(Abort)
Total length | 1567.5 m | 
Energy | 3 - 30 GeV | 
b | 0.9715 - 0.9995 | 
Lorentz g | 4.22 - 33.21 | 
Harmonic | 9 | 
No. Bunches | 8 | 
Revolution periods | 5.38 - 5.23 msec | 
RF freq. | 1.67 - 1.72 MHz | 
Bunch length (time) | 70～200 nsec | 
Bunch length (space) | 20～60 m | 
Tune | FX: nx=22.40, ny=20.75 SX: nx=22.30, ny=20.78 | 
Neutrino
Hadron
3-50BT
Abort

## Slide 4
Main Ring (MR)
500 m
Injection
from RCS
Fast extraction (FX)
Slow extraction
(SX)
RF cavity
WCMs | 1 (measurement) 2 (feedforward) | 
FCTs | 3-50BT | 5
 | MR | 6+1(Abort)
Total length | 1567.5 m | 
Energy | 3 - 30 GeV | 
b | 0.9715 - 0.9995 | 
Lorentz g | 4.22 - 33.21 | 
Harmonic | 9 | 
No. Bunches | 8 | 
Revolution periods | 5.38 - 5.23 msec | 
RF freq. | 1.67 - 1.72 MHz | 
Bunch length (time) | 70～200 nsec | 
Bunch length (space) | 20～60 m | 
Tune | FX: nx=22.40, ny=20.75 SX: nx=22.30, ny=20.78 | 
Neutrino
Hadron
3-50BT
Abort

## Slide 5
Main Ring (MR)
500 m
Injection
from RCS
Fast extraction (FX)
Slow extraction
(SX)
RF cavity
WCMs | 1 (measurement) 2 (feedforward) | 
FCTs | 3-50BT | 5
 | MR | 6+1(Abort)
Total length | 1567.5 m | 
Energy | 3 - 30 GeV | 
b | 0.9715 - 0.9995 | 
Lorentz g | 4.22 - 33.21 | 
Harmonic | 9 | 
No. Bunches | 8 | 
Revolution periods | 5.38 - 5.23 msec | 
RF freq. | 1.67 - 1.72 MHz | 
Bunch length (time) | 70～200 nsec | 
Bunch length (space) | 20～60 m | 
Tune | FX: nx=22.40, ny=20.75 SX: nx=22.30, ny=20.78 | 
Neutrino
Hadron
3-50BT
Abort

## Slide 6
Main Ring (MR)
500 m
Injection
from RCS
Fast extraction (FX)
Slow extraction
(SX)
RF cavity
WCMs | 1 (measurement) 2 (feedforward) | 
FCTs | 3-50BT | 5
 | MR | 6+1(Abort)
Total length | 1567.5 m | 
Energy | 3 - 30 GeV | 
b | 0.9715 - 0.9995 | 
Lorentz g | 4.22 - 33.21 | 
Harmonic | 9 | 
No. Bunches | 8 | 
Revolution periods | 5.38 - 5.23 msec | 
RF freq. | 1.67 - 1.72 MHz | 
Bunch length (time) | 70～200 nsec | 
Bunch length (space) | 20～60 m | 
Tune | FX: nx=22.40, ny=20.75 SX: nx=22.30, ny=20.78 | 
Neutrino
Hadron
3-50BT
Abort

## Slide 7
Wall Current Monitor (WCM)
Time structure of the beam
Beam load compensation (feedforward)
Beam (+)
Wall current (-)
Voltage
Induced charge (+)
Ceramic insulator
Resistance
Output signals from WCM (#019)

## Slide 8
Core
Shield
Beam (+)
Wall current (-)
Voltage
Induced charge (+)
Ceramic insulator
Resistance

## Slide 9
Pipe ID | 165 mm (before 130 mm)
Resistance | 104 mW (13 W×130)
Core | Finemet core FT-3M ID: 207.4 mm, OD: 285.8 Thick: 35.3
New WCM for feedforward
at the 1st straight line
Replacement due to the size enlargement of the pipe
500 m
Injection
from RCS
Fast extraction (FX)
Slow extraction
(SX)
RF cavity
Neutrino
Hadron
3-50BT
Abort

## Slide 10
Pipe ID | 165 mm (before 130 mm)
Resistance | 104 mW (13 W×130)
Core | Finemet core FT-3M ID: 207.4 mm, OD: 285.8 Thick: 35.3
Signal
Shield
Finemet core (FT-3M)
Heat sink
Insulation ceramic
Replacement due to the size enlargement of the pipe
New WCM for feedforward
at the 1st straight line

## Slide 11
Frequency characteristic measurement
Network Analyzer
ROHDE & SCHWARZ
ZVT 8
Measured:
scattering matrix components
S11, S12, S21, S22
300 kHz～2 GHz
Signal
Inner pipe (conductor)
(a) WCM
(b) Reference pipe

## Slide 12
S21
New (Oct15)
Before (Oct07)
(a) WCM
(b) Reference
Straight pipe (L=359)
S21 amplitude
S21 phase

## Slide 13
Normalization of the WCM data
by the reference (straight pipe)
WCM:
Straight pipe:
New (Oct15)
Before (Oct07)
amplitude
300 kHz – 2 GHz (log)
phase
(enlarged)
300 kHz – 50 MHz
(enlarged)
300 kHz – 50 MHz
This WCM has been used for
feedforward (beam loading
compensation) without problem
(for now)

## Slide 14
Fast Current Transformer (FCT)
Beam
Shield
Finemet core
(FT-3M)
Insulation ceramic
Coil (N turns)
Signal
 | 3-50BT | MR | Abort (New)
Inner diameter | 200 mm 230 mm | 134 | 250
Finemet core FT-3M dimension | ID: 390 OD: 490 Thick 50 (25×2) | ID: 216 OD: ～310 Thick: 25 | ID: 375 OD: 500 Thick: 35
Turn Num. of coil | 25 | 50 | 10 × 4 = 40 turns
Time structure of the beam
Beam phase measurement, feedback
Get timing for FX/Abort kicker

## Slide 15
New FCT for abort line
Time structure of the beam
Beam phase measurement, feedback
Get timing for FX/Abort kicker
500 m
Injection
from RCS
Fast extraction (FX)
Slow extraction
(SX)
RF cavity
Neutrino
3-50BT
Abort
 | 3-50BT | MR | Abort (New)
Inner diameter | 200 mm 230 mm | 134 | 250
Finemet core FT-3M dimension | ID: 390 OD: 490 Thick 50 (25×2) | ID: 216 OD: ～310 Thick: 25 | ID: 375 OD: 500 Thick: 35
Turn Num. of coil | 25 | 50 | 10 × 4 = 40 turns

## Slide 16
New FCT for abort line
Time structure of the beam
Beam phase measurement, feedback
Get timing for FX/Abort kicker
Particle counting for abort line
500 m
Injection
from RCS
Fast extraction (FX)
Slow extraction
(SX)
RF cavity
Neutrino
3-50BT
Abort
 | 3-50BT | MR | Abort (New)
Inner diameter | 200 mm 230 mm | 134 | 250
Finemet core FT-3M dimension | ID: 390 OD: 490 Thick 50 (25×2) | ID: 216 OD: ～310 Thick: 25 | ID: 375 OD: 500 Thick: 35
Turn Num. of coil | 25 | 50 | 10 × 4 = 40 turns

## Slide 17
10-turn×4
Uniform 40-turn
Actually, the photo is not for 40 turns.

## Slide 18
Frequency characteristic measurement
Network Analyzer (Agilent E5071)
Measured: S11, S12, S21, S22
300 kHz～1 GHz
usable
amplitude

## Slide 19
40-turn winding uniformly
4 × 10 turns locally
300kHz-100MHz
amplitude
phase

## Slide 20
40-turn winding uniformly
4 × 10 turns locally
300kHz-100MHz
amplitude
phase
4×10-turn type has been adopted

## Slide 21
Beam signal
(sec)
(V)
(sec)
(V)

## Slide 22
(sec)
(V)
(sec)
(V)
(Hz)
(Hz)
FFT Beam signal
Raw beam signal
Further data measurement and analysis will be planned.
Data acquisition and interlock system will be prepared for operation.
Df=1.7 MHz (RF fundamental)

## Slide 23
Summary & prospects
WCMs and FCTs are used to obtain beam bunch and timing information.
Upgraded year by year for higher beam intensity operation.
New WCM for feedforward was installed in the summer 2015.
Localization of resistance: not yet understood.
New FCT for MR abort line was installed in Oct. 2016.
Optimization with data measurement and analysis are going from now.
Data acquisition and interlock system will be also prepared for beam operation.

## Slide 24
Additional

## Slide 25
Main Ring (MR)
500 m
Injection
from RCS
Fast extraction (FX)
Slow extraction
(SX)
RF cavity
WCMs | 1 (measurement) 2 (feedforward) | 
FCTs | 3-50BT | 5
 | MR | 6+1(Abort)
Total length | 1567.5 m | 
Energy | 3 - 30 GeV | 
b | 0.9715 - 0.9995 | 
Lorentz g | 4.22 - 33.21 | 
Harmonic | 9 | 
No. Bunches | 8 | 
Revolution periods | 5.38 - 5.23 msec | 
RF freq. | 1.67 - 1.72 MHz | 
Bunch length (time) | 70～200 nsec | 
Bunch length (space) | 20～60 m | 
Tune | FX: nx=22.40, ny=20.75 SX: nx=22.30, ny=20.78 | 
Neutrino
Hadron
3-50BT
Abort
WCM: 014, 163 (feedforward),
019 (observe)
FCT: 070, 086, 091, 142, 158, 214

## Slide 26
 | FCT#4 | FCT#5
a(slope) 0-500 V | 2.00594e-10 | 2.00283e-10
b(const.) 0-500 V | 0.66261 | 0.12221
FCT for proton counting
15/09/07
Input: 100 V
200 V
300 V
400 V
500 V
Calibration coefficients
(Proton#)=(A/1000*(Counts)+B)*5e+9
 | FCT#4 | FCT#5
A | 981 | 1007
B | 3 | 4
Input
voltage
t (msec)
Area
=V×t

## Slide 27
wfm150701_03_ch2.trc

## Slide 28
Turn 130 (8 bunches are injected)
Current (A) vs time (ch)
Baseline offset is determined for each plot range by Gaussian+const. fitting

## Slide 29
Turn 1000

## Slide 30
Turn 1500

## Slide 31
締め締め部材を買いに行く前に、もしかすると信号読み出しの場所によってスペクトルが変わるかもしれない、
と思い調べてみた。
① スタート: 9/30のベスト
15/10/01

## Slide 32
⑥ 一番初めの位置
明らかに値が大きい

## Slide 33
15/10/01
WCM lower frequency cut off
Freq (Hz)
Amplitude (abs(V(w)))
Cut off frequency
178.08 Hz