## Slide 1: pCM Status at Fermilab
Fermilab pCM Team
Cryogenic systems/SRF weekly meeting
25 October 2016

## Slide 2: CMTS Testing Progress – since previous meeting
  18 Oct – LLRF Microphonics study
  18 Oct – Coupler thermal measurements, #4
  19-20 Oct – Cavity processing for higher gradient
  21 Oct – access including #8 piezo wiring repair
  21-25 Oct – Coupler thermal measurements, #5, 6, 7
  24 Oct – Cavity #5 gradient investigation
  Nikolay will speak to coupler thermal measurements

## Slide 3: Summary of Processing (Genfa)
  #1, #2, #4, #7, #8 see big improvement on gradient
  #3 did not see improvement due to cryogenic limited processing as the pulsed quenches drive up the helium pressure
  #6 improved, but could have been better. Processing was limited when helium pressure increased
  #8 sees vacuum spikes between 17.5-20 MV/m. Has room to improve if vacuum can be improved, as the vacuum trip prevented continued processing. Cavity is close to the pumping port and vacuum gauge. Other cavities potentially has the same gas burst, but may be cryo-pumped by cavities before reaching the vacuum gauge.

## Slide 4: Prototype Cryomodule Latest Preliminary Results
*Usable Gradient: demonstrated to stably run CW, FE < 50 mR/h, no dark current
**Fast cooldown from 45K, >40 g/sec, extrapolated from 2.11K
 | VTS |  | pCM after RF_Conditioning |  |  | 
Cavity | Max Gradient [MV/m] | Q0 @16MV/m | Max Gradient*** [MV/m] | Usable Gradient* [MV/m] | FE onset [MV/m] | Q0 @16MV/m 2K** extrapolated
TB9AES021 | 23 | 3.1E+10 | 19.6 | 18.2 | 14.6 | 2.6E+10
TB9AES019 | 19.5 | 2.8E+10 | 19 | 18.8 | 15.6 | 2.6E+10
TB9AES026 | 21.4 | 2.6E+10 | 17.3 | 17.2 | 17.4 | 2.7E+10
TB9AES024 | 22.4 | 3.0E+10 | 21 | 20.5 | 21 | 2.5E+10
TB9AES028 | 28.4 | 2.8E+10 | 14.9 | 14.2 | 13.9 | 2.4E+10
TB9AES016 | 18 | 2.8E+10 | 17.1 | 16.9 | 14.5 | 2.9E+10
TB9AES022 | 21.2 | 2.8E+10 | 20 | 19.4 | 12.7 | 3.2E+10
TB9AES027 | 22.5 | 2.8E+10 | 20 | 17.5 | 20 | 2.5E+10
Average | 22.1 | 2.8E+10 | 18.6 | 17.8 | 16.2 | 2.7E+10
Total Voltage | 183.1 MV |  | 154.6 (141.4) | 148.1 (137.5) |  | 
***Also limited by administrative limit ~20MV/m
Notes: usable gradient 16% above spec

## Slide 5: Plans for rest of the week/next week
  Complete coupler thermal measurements (#8)
  Ongoing powered magnet studies
  Ongoing microphonics investigation
  ‘Slow’ cooldown from 15 Kelvin
  Cryo system modifications
    warm up and multi-week shutdown as early as next week
  Actively investigating 8-cavity in-phase operation (modified shielding assessment)

## Slide 6: Minimum Acceptance Criteria
✔
✔
✔
Beamline vacuum ~5E-10, unbaked spools
✔
✔
✔
✔
✔
Cavity #5 is only cavity not meeting spec
✖
✔

## Slide 7: Potential Delays
  Microphonics
  Cryo system improvements

## Slide 8: Summary
  Completing(ed) Minimum Acceptance test program
  Microphonics investigation (& remediation) ramping up
  Nearly ready to institute changes to cryo system for stable 2 K operation