## Slide 1: Low-background radio-assay capability in the UK
Paul Scovell – DMUK
UCL – 18/01/17

## Slide 2: Introduction
Last DMUK meeting was Oxford in Dec 2014
In the past few years, low background screening in the UK has matured significantly
Between Boulby, UCL and MSSI we now posses an enviable screening capability
Boulby
  4 ULB HPGe detectors – soon to become 7
UCL
  Agilent 7900 ICP-MS – with almost universal sample digestion capability
MSSI
  Radon emanation chamber – running successfully for several years for SuperNEMO and LZ

## Slide 3: HPGE @ Boulby
  Status in Dec 2014:
    We had two detectors, Chaloner (BEGe 5030) and Lunehead (Ortec p-type) installed under a cleanroom tent in the Boulby Palmer Laboratory
    Represented the lowest background HPGe detectors run in the UK
    We didn’t stop there!
  Sep 2015:
    Additionally added a SAGe well detector and a BEGe 2825
    ”Lumpsey” and “Wilton”

## Slide 4: HPGe – The Past
Well Detector - Lumpsey
Pre-Screening Detector - Wilton

## Slide 5: HPGE – the present
In March 2016, things changed again
  Ageing(!) Palmer Laboratory vacated
  Detectors moved to New BUGS laboratory
  Boulby Underground Germanium Suite
    Class 1000 cleanroom
    HEPA & A/C
    Environmental Monitoring
    LOTS of space
  3/4 of the castles purged with 3 LPM of N2 gas

## Slide 6: HPGe Backgrounds - Lunehead
40K
60Co
137Cs
212Pb
208Tl
214Pb
214Pb
214Bi
208Tl + An

## Slide 7: HPGE BackgroundS - All
Wilton has much smaller (thinner) shield
  Hence much more prominent 40K
Lumpsey not quite ULB
  To get sufficient resolution, some non-LB components must be used
Chaloner and Lumpsey show great low-E response
  210Pb @ 46.5 keV
  234Th @ 63.3 keV

## Slide 8: HPGe Characterisation
IAEA 385 used to characterise detectors
  Irish Sea sediment sent to ~100 labs for analysis (HPGe + others)
Used to characterise detectors for extended and well samples

## Slide 9: Comparison to Literature Values
&
=>
Efficiency
Coincidence Summing
Bq/kg
609
1120
1764
2447
2204
1238
934
768
1847

## Slide 10: HPGE – Screening efforts
We are currently screening components for a number of projects
The majority of samples currently come from LZ including:
  Hamamtsu R11410 3” PMTs & PMT bases
  Individual components for the above
We also screen a smaller number of components for SuperNEMO
We contribute to the SuperKamiokande gadolinium (EGADS) screening programme
We also assay samples for industry and potential new collaborators
However this is now. Looking toward the future…

## Slide 11: HPGe – The future
We continue to screen for LZ through 2017/18
The gadolinium screening programme extends beyond this
Additionally, BUGS is growing
We will add 3 new detectors in the next few months
All in S-ULB configuration in collaboration with Canberra Lingolsheim
  New “revolutionary” BEGe detector (everyone will want one) - Roseberry
  100% p-type - unnamed
  160% p-type – unnamed
We also plan to upgrade the well detector to S-ULB (money permitting!)

## Slide 12: ICP-MS @ UCL
Similarly to HPGe, ICP-MS has moved on in leaps and bounds since Dec 2014
Dedicated ISO Class 6 facility
Operated by Jim Dobson (the other 100% of his time - see LZ talk)
Agilent 7900 ICP-MS
  Integrated auto-sampler
  HF ability
  Able to digest a wide range of materials
Currently dedicated to LZ screening

## Slide 13: ICP-MS
ICP-MS probes the top of the U/Th chain
  Unlike HPGe which mostly probes lower daughters
  Two techniques are complementary
To improve sensitivity we use radiological tracers
  U233 and Th230 – same atomic number, different mass number
Tracking of known contamination of tracer
  Significantly reduces uncertainty of measurement
ICP-MS now hits ultimate sensitivity of 0.05 ppt
  Although this is equivalent to ~10 ppt (g/g) due to dilution

## Slide 14: Radon Emanation @ MSSI

## Slide 15: Radon Emanation
Electro-polished 70 l vessel
Silicon PiN diode at -1500 V
Calibrated using a 222Rn flow-through source
  Gives the standard Po alpha-decay signature
Credit X.-R. Liu (UCL)

## Slide 16: Radon Emanation – The Future?
Many low-background cryogenic detectors expect background from radon to dominate other sources
Radon emanation studies are mostly carried out at room temperature
It is known that radon emanation has a strong dependence on temperature
  Diffusion from materials doesn’t happen so readily at low temperature
Current background predictions for such experiments cannot yet take credit for this reduction
  Leads to unnecessary conservatism
Plans in place to create a low temperature radon emanation facility in the UK
  This will place the UK at the forefront for G3 DM and next-gen 0𝜈𝛽𝛽 detectors

## Slide 17: Conclusions
The UK is positioning itself as a world leader in low background screening
  Very few nations have the breadth of facilities we will have in the next year
Whilst making key contributions to current generation detectors, we have an eye on the future
The is particularly timely with R&D for future low-background experiments on the horizon
Through these facilities we have the opportunity to gain strong leadership in these experiments
  Results will shape the design and construction
  These facilities will be used to prove that we can source and measure materials to the required sensitivity
  It is a low-risk and high-reward!