## Slide 1: pCM Cavity Testing
Genfa Wu, Elvin Harms
pCM test workshop

## Slide 2: Outline
Introduction
Minimum Acceptance Tests Updated
  Cavity multipacting processing
  Unit Testing
  Heatload measurement
Extended Tests in Progress
  Fast cool down strategy

## Slide 3: Introduction
pCM at Fermilab was installed in July 2016.
Warm conditioning was conducted on all couplers in August 2016.
pCM cooled down in September 2016
pCM started cavity testing at 1.9K briefly under cold compressor mode
pCM cavity testing was completed at 2.1 K under warm pumping mode (Kinney pump) in October 2016
Cryoplant was shutdown for an upgrade.
pCM cooled down to 2K in December.

## Slide 4: Minimum Acceptance Criteria
✔
✔
✔
Beamline vacuum ~5E-10, unbaked spools
✔
✔
✔
✔
✔
Cavity #5 is only one not meeting spec
✔
✔
✔
pCM + FeedCap + EndCap ~13 W Static 2K
64 W 2K dynamic of 128 MV

## Slide 5: Multipacting Processing
Short RF pulsing to avoid the cryogenic instability
  1-10 Hz
  20-50 mS (Qext ~ 1e7 - 4e7)
  2-3 kW

## Slide 6: Multipacting Processing – CAV1
CAV1 was previously limited at 17.2 MV/m due to multipacting. RF pulsed processing improved the gradient to 19.6 MV/m in pulsed mode.

## Slide 7: MP in HOM of cavity#1 (zoom)
Gradient
RF power
HOM can Temp.
Radiation
MP starts at ~12MV/m
HOM can T quickly goes up above threshold gradient
Processed in short pulses (seconds) until T< 20K
After processing no MP up to max gradient
Radiation reduced after processing,
Exist w/o MP
Before processing
processing
After
processing
Slide made by N. Solyak

## Slide 8: Multipacting Processing – CAV2
CAV2 was previously limited at 16.8 MV/m due to multipacting. Previously FE onset was 15.6 MV/m. RF pulsed processing improved the gradient to 19 MV/m in pulsed mode. FE onset was the same. After 19 MV/m, quench was persistent. Similar to VTS result.
CAV2 ran 18.8 MV/m stably.

## Slide 9: Multipacting Processing – CAV3
Processing has little effect (~2 hours). Cavity gradient was stable around 17 MV/m. (Exceeds specification)
Processing was limited due to pressure cannot keep up at 2.1K. Processing created pressure instability (>36 torr)
Additional processing at 2.0K may be achievable.

## Slide 10: Multipacting Processing – CAV4
CAV4 Qext was adjusted to 4.18e7 40ms 2Hz processing was successful.
CAV4 maximum gradient was >20 MV/m Parked at Eacc_ACNET = 19.2 MV/m. for 2 min. Pressure increased until cavity quenched. Very small FE at this gradient CW.
Processing was around 30 min. There was radiation during processing.

## Slide 11: Multipacting Processing – CAV6
CAV6 was previously limited at 16.3 MV/m due to multipacting. Previously FE onset was 14.5 MV/m. RF pulsed processing improved the gradient to >17 MV/m in pulsed mode. FE onset was the same. Continued pulse processing brought helium pressure too high.
CAV6 ran 16.9 MV/m stable. More processing after cryogenic shutdown may benefit.

## Slide 12: Multipacting Processing – CAV7
Cavity initially quenched at 17.4 MV/m, the same gradient as before 100K warm cycle. Pulsing was set at 40 ms 2Hz (thanks to Mike), after pulsed processing, cavity went to 20 MV/m without pulsed quench.
Changing to CW, cavity was parked at 19.4 MV/m. FE was still below 50 mR/h. As the helium pressure started to increase, eventually cavity quenched after 2 min at 19.4 MV/m. I consider this limit is due to pressure increased to around lambda point.

## Slide 13: Multipacting Processing – CAV8
CAV8 was previously limited at 16.8 MV/m due to multipacting. Previously FE onset was 13.8 MV/m. coupler Qext to 5e7. RF pulsed processing improved the gradient to >20 MV/m in pulsed mode. No radiation was detected. After parking at 19.3 MV/m for one or two minutes, cavity beam line vacuum experienced a vacuum spike which SSA8 also tripped. Radiation was moderate. Cryo was stable. Cavity temperatures/coupler temperatures were all normal.
During the subsequent pulsed mode or CW mode, CAV8 experienced vacuum/radiation spikes at an arbitrary gradient between 17.5 MV/m and 18.5 MV/m. See attached screen capture. CAV8 was speculated to have multipacting that drove the residual gases around. It is to be verified after 300K warm up and the cryogenic shutdown.

## Slide 14: Multipacting Determination
Consistent gradient limit (almost all cavities)
No or very low radiation (except CAV5 FE)
Quench field can be conditioned
No electron signals in FP
Two cavities experienced HOM can heating
  One processed away (CAV1)
  One still present (CAV5 with high FE)
VTS experience was similar

## Slide 15: Possible Causes of Multipacting
Non baked beam line components that release H2
Long period (5-6 month) of cavity filled with N2
Coupler outgassing after exposed to air briefly
Active pumping is supplied by an ion-pump (Not a significant source)
  Turbo pump was valved out before cool down
1-cell cavity study is in progress

## Slide 16: Multipacting Mitigation
RF pulse conditioning
  Successful in FNAL/pCM
Warm up and pump longer time
  No coupler conditioning

## Slide 17: Cavity Limitations
Summary as of 2016-10-31 |  |  | at 2.1K |  |  | 
Cavity | Name | Eacc @<50mR/h [MV/m] | Max Gradient [MV/m] | Gradient limitation | FE onset [MV/m] | NOTE
1 | TB9AES021 | 18.2 | 21.2 | quench (cryo, admin limit) | 14.6 | Multipacting improved, <15min
2 | TB9AES019 | 18.8 | 19.0 | quench | 15.6 | Multipacting improved, <15min
3 | TB9AES026 | 17.2 | 17.3 | Quench (Multipacting, cryo, no x-ray) | 17.3 | No improvement, ~2h, limited by cryogenics
4 | TB9AES024 | 20.5 | 21.0 | Quench (Multipacting, admin limit) | 21.0 | Multipacting improved, <30 min
5 | TB9AES028 | 14.2 | 14.9 | Field Emission | 13.9 | 
6 | TB9AES016 | 16.9 | 17.1 | Quench (Multipacting, cryo) | 14.5 | Multipacting improved, <15 min
7 | TB9AES022 | 19.4 | 20.0 | Quench (Multipacting, admin limit) | 12.7 | Multipacting improved, <15 min
8 | TB9AES027 | 17.5 | 20.0 | Quench (Multipacting, cryo) | 20.0 | Multipacting improved, <10 min, limited by vacuum interlock
TOTAL VOLTAGE |  | 148.1 | 156.2 | Average Eacc @FE onset = | 16.2 | 
CAV3, CAV6 and CAV8 all exceeds specification. Additional RF processing at 2.0K may further improve cavity gradients.

## Slide 18: Multipacting Processing Improved the Gradients at 2.1K
 | pCM (DOE Review) |  |  | pCM after RF Conditioning* |  |  | 
Cavity | Max Gradient [MV/m] | Usable Gradient* [MV/m] | FE onset [MV/m] | Max Gradient [MV/m] | Usable Gradient* [MV/m] | FE onset [MV/m] | Q0 @16MV/m 2K** extrapolated
TB9AES021 | 19.6 | 18.2 | 14.6 | 19.6 | 18.2 | 14.6 | 2.6E+10
TB9AES019 | 17.0 | 16.8 | 15.6 | 19.0 | 18.8 | 15.6 | 2.6E+10
TB9AES026 | 17.3 | 17.2 | No FE | 17.3 | 17.2 | No FE | 2.7E+10
TB9AES024 | 16.5 | 16.0 | No FE | 21.0 | 20.5 | No FE | 2.5E+10
TB9AES028 | 14.9 | 13.8 | 11.5 | 14.9 | 14.2 | 13.9 | 2.4E+10
TB9AES016 | 16.7 | 16.7 | 14.5 | 17.1 | 16.9 | 14.5 | 2.9E+10
TB9AES022 | 17.4 | 17.1 | 12.7 | 20.0 | 19.4 | 12.7 | 3.2E+10
TB9AES027 | 16.8 | 16.6 | 13.8 | 20.0 | 17.5 | No FE | 2.5E+10
Average | 17.0 | 16.6 | 14.7 | 18.6 | 17.7 | 16.2 | 2.7E+10
Total Voltage | 141.4 MV | 137.5 MV |  | 154.6 MV | 148.1 MV |  | 
*Usable Gradient: demonstrated to stably run CW, FE < 50 mR/h, no dark current
**Fast cooldown from 45K, >40 g/sec, extrapolated from 2.11K
***Also limited by administrative limit ~20MV/m

## Slide 19: Cavity Performance at 2K (in progress)
Summary as of 2017-01-07 |  |  |  |  | 
 | pCM after RF Conditioning*** |  |  |  | 
Cavity | Max Gradient [MV/m] | Usable Gradient* [MV/m] | FE onset [MV/m] | Q0 @16MV/m 2K** extrapolated | Q0 @16MV/m 2K
TB9AES021 | 21.2 | 18.2 | 14.6 | 2.6E+10 | 2.7E+10
TB9AES019 | 19.0 | 18.8 | 15.6 | 2.6E+10 | 3.1E+10
TB9AES026 | 17.3 | 17.2 | 17.3 | 2.7E+10 | 
TB9AES024 | 21.0 | 20.5 | 21.0 | 2.5E+10 | 
TB9AES028 | 14.9 | 14.2 | 13.9 | 2.4E+10 | 
TB9AES016 | 17.1 | 16.9 | 14.5 | 2.9E+10 | 
TB9AES022 | 20.0 | 19.4 | 12.7 | 3.2E+10 | 3.3E+10
TB9AES027 | 20.0 | 17.5 | 20.0 | 2.5E+10 | 
Average | 18.8 | 17.8 | 16.2 | 2.7E+10 | 
Total Voltage | 154.6 MV | 148.1 MV |  |  | 
*Usable Gradient: demonstrated to stably run CW, FE < 50 mR/h, no dark current
**Fast cooldown from 45K, >40 g/sec, extrapolated from 2.11K
***Also limited by administrative limit ~20MV/m

## Slide 20: Field Emission
Arrays of radiation detectors along the CM and both ends.
Average Eacc@FE-onset >16 MV/m
Notes: 15 radiation detectors. Maximum was used to plot.

## Slide 21: FE Monitoring
Ion chamber detectors

## Slide 22: Field Emission - Potential Causes
Coupler was used on CAV5 in HTS and cycled through HTS preparation.
Other couplers can have potential sources of particulates
Copper plated bellows are new compared to ILC cryomodule
pCM had to dis-assemble parts that was considered containing poor copper plating

## Slide 23: Field Emission - Potential Mitigations
RF conditioning is limited by RF power
  High power can allow strong Qext and high peak field
  SSA power combiner may be an option
Helium processing is risky
High power pulsed processing
  Requires high power klystron
CAV5 FE processing is not planned as other cavities can make up the total voltage.

## Slide 24: Cryomodule Unit Testing (in progress)
Combined radiation < 34 mR/h (not phase synchronized)
115 MV/m
120 MV
Notes: Liquid level can be maintained.

## Slide 25: Heat Load Measurement - Static
Measured at 4K
Total ~ 13 W. pCM + FeedCap + EndCap
Courtesy of R. Wang

## Slide 26: Heat Load Measurement – 2K Dynamic (Very Preliminary)
Measured at 2K
Total ~ 64 W@128 MV (extrapolated from 7 Cavity heatload)

## Slide 27: Fast Cool Down
 | >40 g/s * | 32 g/s
Cavity | delta T(K) | delta T(K)
CAV1 | 3.8 | 3
CAV4 | 5.6 | 2
CAV5 | 8.3 | 5
CAV8 | 4.7 | 2.5
* Mass Flow out of range
Fast cool down in a cryomodule demonstrated
pCM Test Review January 11, 2017
Notes: Flux expulsion is to be tested with external magnetic field

## Slide 28: Cool Down Strategy (in progress)
To be developed
  Compare to slow cool down when remnant magnetic field is at minimal.
  Evaluate thermal current induced remnant field.
  Evaluate flux expulsion efficiency when ambient field is raised above specification.

## Slide 29: Summary
Minimum tests completed and acceptance criteria were mostly met except microphonics.
  Cryomodule exceeds total voltage specification
  Cryomodule exceeds heatload specification
  Radiation acceptable, no dark current measurement until microphonics is suppressed to below 10Hz.
Extended tests is in progress
  Fast cool down strategy is to be developed.
Notes: Through pcm testing, CM02 string assembly, single cavity test

## Slide 30: Extra slides

## Slide 31: Cavity Performance Measurement - Calibration
RF calibration
  Forward power at SSA output before the circulator
    Minimize the directional coupler signal deviation due to reflected power
    Calibration using calorimetric load
  Transmitted power and HOM power
    Network analyzer for warm and cold transmission lines
  Qext1
    Fast oscilloscope for decay time
    RF power decay
Heatload
  Massflow vs. known heater power

## Slide 32: LLRF Power Measurement Linearity Measurement
Chart: Labview LLRF Linearity Check
| 46.36 | 87.5 | 133.7 | 203.0 | 279.7 | 358.8 | 402.5 | 421.2
 | 46.7 | 88.3 | 136 | 207 | 286 | 366 | 412 | 428

## Slide 33: Directional Coupler Forward Power
Chart: Directional Coupler Power vs Sliding Short Position
| 0.0 | 1.0 | 2.0 | 3.0 | 4.0 | 5.0 | 6.0 | 7.0 | 8.0 | 9.0 | 10.0
SPA Single directional coupler at output of SSA - Forward Power | 462 | 463 | 466.4 | 467 | 466 | 463.6 | 461.3 | 462 | 465.3 | 467.2 | 467
SPA Dual directional coupler at output of circulator - Forward Power | 455 | 442 | 449 | 467 | 479 | 477 | 461 | 444 | 445 | 460 | 476
Mega Dual directional coupler at sliding short - Forward Power | 459 | 488 | 486 | 457 | 432 | 427 | 448 | 481 | 490 | 469 | 440
SPA Dual directional coupler at output of circulator - Reflected Power | 372 | 373 | 412 | 448 | 453 | 423 | 381 | 367 | 397 | 438 | 455
Mega Dual directional coupler at sliding short - Reflected Power | 395 | 406 | 385 | 351 | 338 | 353 | 386 | 408 | 395.4 | 362 | 340
Data from John Reid and Co.

## Slide 34: Cavity Performance Measurement - Gradient
Method A – Forward Power (Production method)
  Measure FPC Qext and forward power
Method B – Transmitted Power
  Measure transmitted power and field probe Qt (using VTS Qt)

## Slide 35: Cavity Gradient Measurement - Comparison
Method A – Forward Power (Production method)
  Waveguides assembly does not change
  Calibration is stable
  Less affected by environment
Method B – Transmitted Power
  Transmission line less rigid
  May require configuration changes due to signal levels
  More connections/disconnections
  VTS to cryomodule changes can happen
  VTS measurement error propagation

## Slide 36: Gradient Measurement
Gradient comparison as of 2016-09-19 |  | 
CAVITY | Name | % of Forward method over Pt method
1 | TB9AES021 | 4.2%
2 | TB9AES019 | 5.8%
3 | TB9AES026 | -5.8%
4 | TB9AES024 | -0.8%
5 | TB9AES028 | -4.5%
6 | TB9AES016 | 1.1%
7 | TB9AES022 | -3.3%
8 | TB9AES027 | 7.2%
 | Average | 0.5%
Note: | Mostly compared at low field | 

## Slide 37: Coupler Heating Effect
The FPC antenna moves into the cavity as the coupler warms up.
Pfor increased by 4%-7% for the same gradient.
Two hours to stabilize.

## Slide 38: Error Analysis - Eacc
Pt Method |  | 
_Eacc | _Pt | _Qt
7% | 8% | 12%
Pt Method |  | 
7% | 8% | 12%
Pf Method similar to Pt Method |  | 
_Eacc | _Pf | _Q1
3% | 5% | 2%
Pf Method similar to Pt Method |  | 
3% | 5% | 2%
Pf calibrated by calorimetric load
Q1 by 100%-50% transmitted power decay
Pt calibrated by insertion loss
Qt using VTS measured data

## Slide 39: Gradient Measurement during Production Testing
Measure decay time (Qext1) at low power (6-8 MV/m) in the beginning of the power rise.
Calibrate the transmitted power to the Eacc determined by the Pf.
Measure Pt to determine the Eacc. (Avoids the coupler heating effect).
Verify the Eacc is consistent at 16 MV/m between LLRF power measurement and agilent power meter

## Slide 40: Cavity Performance Measurement - Heatload
Method A – Mass flow change (Fermilab Production)
  Measure mass flow delta to extrapolate heat load
Method B – Measure compensation heater
  Maintain constant mass flow
  Measure RF compensation heater power
Method C – dP/dT
  Close JT valve to allow pressure rise due to known heater powerdynamic heating
  Close JT valve to allow pressure rise due to dynamic heating
  Measure pressure rise dP/dT to calculate the heat load

## Slide 41: Q0 Measurement
Record the static mass flow the day before and the day after.
Record the cryogenic parameters with known heater powers
  (3 hours for each of 2,4,6,8,10,12,14,16,18,20 W)
Bring cavity to 16 MV/m and wait 1-2 hours or until cryogenic parameters are stable.
Measure mass flow or pressure at the stable condition
Determine the dynamic heatload from mass flow or pressure of 16 MV/m from heater calibration.
Calculate the Q0 based on dynamic heatload and cavity gradient.

## Slide 42: Mass Flow Method (Mass flow calibrated with a known heater)
Forward Power and Gradient
Liquid level, pressure and JT opening
Gradient and Mass Flow
Liquid Level Correction
Delta Mass flow = 1 g/s -> 13.7 W at supply mass flow
1 g/s -> 20.64 W at return mass flow

## Slide 43: Q0 Accuracy
Both pressure and mass flow has standard deviation of <2%. 8% peak to peak deviation.
Heat load has 2.8% standard deviation, 11% peak to peak deviation
Gradient has 3% error bar.
Q error bar is 4.1% (std deviation) and 11.5% (based on peak to peak deviation).

## Slide 44: CAV1 Temperature Evolution
Coupler 5K
Mag shield outer
Mag shield inner
HOM coupler
Cavity T