## Slide 1: PIP-II Beam Instrumentation
Vic Scarpine
PIP-II Machine Advisory Committee Meeting
10-12 April 2017

## Slide 2: Outline
PIP-II Requirements
Review of selected instruments
Risks and R&D plan
Summary

## Slide 3: PIP-II Requirements
Accelerator instrumentation
sections:
  Ion source & LEBT
  MEBT
  Superconducting linac
  Transport lines
  Rings
  Present focus is development of instrumentation to test at PIP2IT
  PIP-II focus on pulsed operation with an eye toward CW
    CW operation has impact on instrumentation choices
The scope of beam diagnostics are to identify and provide the instrumentation systems necessary to successful commission, characterize and operate all PIP-II sub-accelerators.

## Slide 4: Requirement Highlights from CDR
Additional beam diagnostics:
Transverse emittance
Longitudinal bunch profiling
Transverse beam halo
Bunch-by-bunch chopping efficiency
Minimum set of beam measurements:
Beam intensity
Beam position / orbit
Transverse beam profiles
Beam phase / timing
Beam loss
Operational parameters:
When possible the beam instrumentation and diagnostics systems must be able to operate over a wide bandwidth from few microsecond long pulses for machine commissioning and tuning to long operational pulses. In addition, these systems must be compatible with CW beam operation.
Beam instrumentation is required to be integrated into reliable, fail-safe machine protection system

## Slide 5: From the CDR
Table 3.40 presents an estimate of quantities of different beam instrumentation systems by location for the linac and the 800 MeV beam transport line. In the table the SC linac also includes the linac extension capable to include 4 additional HB650 cryomodules.

## Slide 6: Beam Position and Phase – BPM System
PIP-II requires 89 warm BPMs and 37 cold BPM for the linac and 800 MeV transfer beam line
  capable of measuring horizontal and vertical position and beam phase relative to the RF, as well as a relative measure of the beam intensity
  Single four-button BPM located near each beam focusing element (solenoid, quad, doublet, triplet)
  The geometry of all BPMs will be similar with their aperture slowly increasing with acceleration
BPM Measurement modes
  pulsed beam at a single trajectory
  pulsed beam split into two trajectories – in MEBT chopper section
 | One-trajectory Mode | Two-trajectory Mode
Transverse position, mm | 10 | 100
Phase, degrees of 162.5 MHz | 0.05 | 0.5
Relative intensity, % | 1 | 2
BPM Pulsed Beam Measurement Requirements

## Slide 7: Warm and Cold BPM Designs
Section view of a PIP2IT warm MEBT BPM and BPM mounted to a MEBT quadrupole and benchtop stretched wire mapping
Section view of cold BPM and a SSR1 cold BPM assembly and benchtop stretched wire mapping

## Slide 8: MEBT BPM Development at PIP2IT
Status:
BPMs operational
  Pulsed beam only (<0.5 ms)
  Average position, phase, intensity per pulse into ACNET
  Position resolution ~10 mm
  Phase resolution ~0.2 deg
Time-of-flight BPM installed and tested
  BPM on movable stage
  Energy resolution ~0.3%
Warm 4-buttom BPMs imbedded in quad doublets
Narrow-band DAQ with FPGA-based electronics for CW and pulsed beam operation
12 channel, 14 bit, 125 MHz boards
Analog filter and amplifier boards

## Slide 9: Beam Current Measurement Systems
Primary Systems
DCCT
  Bergoz DCCT
  Tested in PIP2IT LEBT
Toroids
  Pearson split-core toroids
    0.1 V/A ±1%
    ~10 Hz to ~4 MHz
    7% per ms droop
  Testing in PIP2IT LEBT and MEBT
  Studying limitations to long-pulse measurements
  Input for MPS
    Places tight requirements on toroid system performance
Secondary Systems
Relative measurements
  Sensitive to beam effects
    Bunch position, velocity and shape
Ring pickups
  Testing in PIP2IT MEBT
  Systematic effects being studied
  Possible relative measurement for MPS
BPMs
  Testing in PIP2IT MEBT
Resistive Wall Current Monitor

## Slide 10: Transverse Beam Profiling Systems
Wire scanner system
MEBT and transfer beamline
Laserwire system
Linac
Parameter | Value
Spatial resolution | < 0.1 mm
Signal dynamic range | > 103
Spatial range | ± 15 mm
Single wire scan – step transversely through beam
Test in diagonal (45o) port of PIP2IT MEBT scraper
Horizontal and vertical transverse profiles in one scan
50 micron tungsten-rhenium
Prototype Wire scanner
H- + g  Ho + e-
psec mode-locked laser (MML) used to measure both transverse and longitudinal profiles
Narrow-band lock-in amp detects modulated signal
Measure profiles by either:
  Collection of electrons
  reduced-beam current
    Allows laser monitor to fit between cryomodules
Prototype Laserwire

## Slide 11: Allison Scanner for MEBT Transverse Emittance Measurements
Design a water cooled Allison-style MEBT emittance scanner based on LEBT scanner
Gives faster phase-space measurement
Reuse LEBT electronics
2.1 MeV  requires longer deflector plates  requires more beam line space
  ± 12 mrad angular range
Tight alignment requirements
Higher beam power  no CW operation
Installed in Dec 2016 for horizontal emittance measurements
Commissioned, calibrated and operational

## Slide 12: MEBT Chopper Extinction Measurement – Under Design
Use downstream Resistive Wall Current Monitors (RWCM)
Extinction
  Can average over many bunches or take a single shot measurement
  Measurement BW < 1 Hz
  Fits to bunch shape
  Measure impact on adjacent bunches

## Slide 13: Summary
Specifications of required beam measurements solidifying
CDR giving better picture of beam instrumentation choices
PIP2IT allowing for a majority of beam instruments to be tested up real PIP-II beam conditions
  Many systems operational or being tested
Continued R&D at PIP2IT essential to mitigate risks for PIP-II development

## Slide 14
Backup slides

## Slide 15: PIP-II Risks to Address and R&D Plans
Beam current measurement systems
  Short pulses to CW operation adds difficulties
  Integration into MPS may require changes in technology choices
    Tight tolerances required to protect SCF cryomodules
  R&D at PIP2IT
    Testing hardware and algorithms for wide beam pulse range
Choice of laser profiling technology
  Choice of technology impacts early in PIP-II design
    Infrastructure for laser transport must be part of PIP-II building/tunnel layout
    Profiling technique effects beamline space requirements
  R&D at PIP2IT
    Installation of prototype laserwire system
    Allows for optional laser choices
    Allows for opportunity to study different signal detection options
    Allows for study of systematic effects

## Slide 16: PIP-II Risks to Address and R&D Plans (continued)
Integration of instrumentation into MPS
  Choice of MPS architecture effects development of instrumentation DAQ system design
    Major architectural changes will need time to develop, prototype and test
  R&D at PIP2IT
    Testing of different MPS options
Beam loss monitor system
  Choice of BLM technology for low-energy measurements not clearly defined
  R&D at PIP2IT
    Testing of fast, sensitive PMT based x-ray detectors
    Testing of rad-hard, sensitive diamond detectors in vacuum
      Possible halo detectors
Longitudinal beam profiling
  Very small (4 ps) beam longitudinal size is difficult to measure
    No clear choice of technology
      OTR + streak camera?
      fs laser?
    Can not be tested at PIP2IT