## Slide 1
protoDUNE – single-phase Instrumentation - Gas Analyzers
Stephen Pordes, Fermilab

## Slide 2
protoDUNE – single-phase Instrumentation - Gas Analyzers
Motivation:
Check the quality of the argon delivered for water, oxygen and nitrogen content
Monitor the progress of the purging of the cryostat atmosphere
Look for the existence/appearance of leaks during the gas phase of recirculation and purification
Ensure that the ultimate level of atmospheric contamination is low enough not to be a significant load on the purification system
Monitor the effectiveness of the liquid purification in its early stages before the purity monitors come on line.
Help diagnose problems should they occur during operation, including assessing contamination introduced in case of emergency removal/replacement of equipment.
This list may not be complete but is based on our experience at Fermilab. The first use was motivated by trying to understand the purging process for removing the atmosphere from a cryostat.

## Slide 3
protoDUNE – single-phase Instrumentation - Gas Analyzers
Experience
Have (1 in ~10) rejected trailer-load of argon for high oxygen (7 ppm vs 5 of spec; company initially denied and then agreed)
Have detected leaks during purging and recirculation process
Have used to understand major sources that affect electron drift-lifetime
Have used to monitor purification process in cryostat before purity monitors are sensitive
Have used to compare contaminants in ullage with electron drift-lifetime to understand contaminant flow in ullage.
Have used as diagnostic when incidents (e.g., pump restart) have affected electron drift-lifetime.

## Slide 4
protoDUNE – single-phase Instrumentation - Gas Analyzers
Example
2 days
O2 analyzer (ppb)
2
1
3
Pump turn on produces a spike in Oxygen

## Slide 5
protoDUNE – single-phase Instrumentation - Gas Analyzers
Performance:
Measure one type of contaminant only and over limited range
Sensitivities for oxygen and water range down to fractions of a ppb
Nitrogen devices are less sensitive and in our experience much less stable – they require weekly calibration to operate at the ~ 0.3 ppm accuracy and should be calibrated before any crucial test
Need a switchyard to allow sampling from multiple sources
Moderate and high sensitivity analyzers available to the first atmosphere-to-purity experience at Fermilab

## Slide 6
protoDUNE – single-phase Instrumentation - Gas Analyzers
The Menorah switchyard (7 in, 8 out)

## Slide 7
protoDUNE – single-phase Instrumentation - Gas Analyzers
Proposal – top of range:
Vaisala or equivalent Dew Point Monitor
Medium range (DF310E) and high sensitivity (DF550) oxygen analyzers
Nitrogen analyzer (Servomex Servopro or LDetek8000)
High sensitivity (0.6 ppb) (but not highest sensitivity) H20 analyzer (Tiger Optics Halo 3)
Modest proposal
Omit the H20 analyzer

## Slide 8
protoDUNE – single-phase Instrumentation - Gas Analyzers
Response to charge:
In response to the charge:
1. The set of instrumentation proposed will cover the needs for gas analysis from delivery to purge to operations
2. It is envisaged that the same type of instrumentation is appropriate for DUNE. The length of DUNE may require several sets of instrumentation distributed along the length of the cryostat.
3. The major item that is not a stock commercial device is the valve switchyard. There is experience both in building such a device in-house and in purchasing one so the times to procure or construct are known. The analyzers come with commercial warranties and the switchyard can be constructed and qualified at Fermilaband checked at CERN. The items can be shipped to CERN as ordinary commercial objects. Calibration/span gases will be ordered and used for initial commissioning. It is assumed that installation of the piping to the analyzer rack will be done by the Neutrino Platform.
4. There is no installation plan as yet. (April 2017). This clearly needs to be resolved with the Neutrino platform.
5. The connections to the cryogenics system are defined in the Piping and Instrumentation Diagram (EDMS). The location of the analyzer rack and switchyard are yet to be determined.
6. The slow control system needs to record the readings from these devices and possibly exercise a calibration cycle. It is expected that the measurements from the analyzers will be recorded in a data-base and accessible in some historian. The readout interfaces have not been discussed with the slow control system and this needs to be understood and agreed in the next 6 months.

## Slide 9
protoDUNE – single-phase Instrumentation - Gas Analyzers
Response to charge (2)
7. We expect that the electrical grounding issues presented by the external piping used for sampling will be resolved by the electrical engineers when they define the grounding for the external piping in general.
8. The equipment is all external to the cryostat and is not expected to interfere with detector operations. We will check if the pump which may be used to boost the sample line pressures produces any electrical noise that affects the detector readout. The system needs to be operating particularly during the gas purge and recirculation phases.
9. safety of installation and operation The analytic instrumentation carries electrical certification. Installation of the switchyard and the instrumentation requires careful leak-checking but is otherwise straight forward. Once installed and commissioned, the operations burden is limited to setting the appropriate sources on the appropriate devices and so recording, and periodic calibrations of the N2 sensor. When vacuum pumps are used, procedures will be followed to ensure that we do not pull a vacuum where it is not wanted.
10. The proposal is based on experience at Fermilab where these instruments have been used extensively. We have consulted the people who built the systems at Fermilab and intend to continue this consultation and to incorporate the results of their and local (CERN) experience.