## Slide 2: NP02 -Proximity Cryogenics -Internal piping
Johan Bremer & Michel Chalifour on behalf of Cryogenics team

## Slide 3: Proto DUNE-NP02 status
Requirements
Schedule
Installation sequences
Chimneys interfaces
Internal piping
Return of experience

## Slide 4
 | Parameter | Value | Notes
 | GAr purge flow rate | 88 m3/hr | From 1.2 m/hr
 | Maximum cool-down rate TPC | 40 K/hr 10 K/m | T sensors to be defined and placed by Internal Cryogenics/Detector
 | Maximum Delta T between any two points in the detector | 50 K | T sensors to be defined and placed by Internal Cryogenics/Detector
 | Design Pressure | 1345 mbar | Design Pressure
 | Operating gas pressure | 1000 mbar | Operating gas pressure with a cold compressor
 | LAr filling flow rate | 18 l/min | Assuming 2 trucks/day
 | Cryostat static heat leak | 3.0 kW | GAr boil-off (18 g/s)
 | Other heat loads (estimate) | 5.0 kW | Total estimate is ~ 8 kW (50g/s)
 | LAr circulation (5 days turnover) | 1.63 kg/s | Nominal value (1 pump)
 | Max emptying (w both LAr pumps) | 3.2 kg/s | Limited by size of tank/truck
 | Condenser size | 16 kW | 
Requirements for NP.02

## Slide 5: NP02-Schedule Demaco

## Slide 6: Installation Sequence: Overview
NP-02
NP-04
H2 beam
H4 beam

## Slide 7: Installation Sequence: NP-02
March 2017- installed cryostat

## Slide 8: Installation Sequence: NP-02
July 2017- platforms installed
platforms

## Slide 9: Installation Sequence: NP-02
September 2017- racks & cable trays
October 2017 – cabling to field boxes
cable trays and field boxes

## Slide 10: Installation Sequence: NP-02
August-December 2017- valve boxes
valve boxes are delivered and installed in a staggered schedule

## Slide 11: Installation Sequence: NP-02
December 2017-January 2018 - cold lines
Cold lines

## Slide 12: Installation Sequence: NP-02
January 2018 - warm lines
Warm valve panel and warm piping

## Slide 13: CHIMNEYS manifold Interfaces (artistic view)
Manifold layout and interfaces are still under discussion with the collaboration

## Slide 14: Internal Cryogenics – Info
The Internal Cryogenics is a set of pipes and manifolds located inside the cryostat.
Its goal is:
  To distribute the LAr inside the cryostat during fill and LAr recirculation
  To cool down the cryostat and the detector (possible according to the requirements?) with two set of sprayers (LAr/GAr mixture and GAr momentum)
  To remove air from the cryostat prior to the start of the cryogenics operations (argon purge)
  To send the boil-off GAr to the Proximity Cryogenics
  To send the GAr to the Pressure Safety Valve (PSV) in case of a release event

## Slide 16
From yesterdays presentation from Adamo:
to discuss how to make the cool-down within the limits given?

## Slide 17: Internal Cryogenics – Status
Conceptual Design and location of pipes approved by NP-04. ✓
Contract awarded to Kriosystem: Feb 17, 2017. ✓
Kick-off meeting: Feb 27, 2017. ✓
Preliminary Design in progress.
Current estimated delivery: Jun 16, 2017.
Current estimated duration of installation: 3 weeks.

## Slide 18: Platforms for the NP02 boxes
For the CVB and PSVB
For the filters boxes
Design is approved
Cable trays and cable laying is following the schedule

## Slide 19: Return of Experience of WA 3*1*1
  To go to welds and limit the quantity of gaskets (e.g: interfaces with the chimneys)
  For convenience in EHN1 there are three LAr filtration vessels per cryostat (SP, DP), 1 filled with Mole Sieve and 2 filled with Copper beds in series.
  The initial plan was to regenerate them as a single unit, but it is now possible to regenerate them independently and then place the EH outside the boxes (risk mitigation).
  Consolidate the choice to place a LN2 phase separator at the inlet of the HEX condenser.
  Consolidate the choice to place the cool-down nozzles at the top part of the cryostat.
  Carefully check if there is any obstacle in front of the cool-down nozzles location.