## Slide 1: MQXFAP2 Electrical Acceptance Tests LBNL
Jordan Taylor
December 07, 2017

## Slide 2: Test Parameters
Coils Tested: QXFA101, QXFA102, QXFA104, QXFA105

## Slide 3: Electrical Test Data
Coil 101
RLQ |  | 
 |  | R (mΩ)
 |  | Ls (mH)
 |  | 20 Hz
 |  | 100 Hz
 |  | 1 kHz
 |  | Q
 |  | 20 Hz
 |  | 100 Hz
 |  | 1 kHz
QH R (Ω) |  | 
 | NT | A01
 | T | A02
 | T-P | B01
 | NT-P | B02
 | T-M | B03
 | NT-M | B04
VT R (mΩ) |  | 
 |  | A1
 |  | A2
 |  | A3
 |  | A4
 |  | A5
 |  | A6
 |  | A7
 |  | A8
 |  | B8
 |  | B7
 |  | B6
 |  | B5
 |  | B4
 |  | B3
 |  | B2
 |  | B1
604.07 |  | 616.6
4.848 |  | 4.834
3.376 |  | 3.307
1.878 |  | 1.867
0.81 |  | 0.79
1.6 |  | 1.6
1.99 |  | 2.02
15 |  | 16
Pre-Shipment |  | Post-Shipment
In Ship Crate |  | In Ship Crate
OD Up |  | OD Up
15 |  | 16
Pre-Shipment |  | Post-Shipment
In Ship Crate |  | In Ship Crate
OD Up |  | OD Up
3.466 |  | 3.8
3.514 |  | 3.8
1.871 |  | 2.3
1.897 |  | 2.1
1.910 |  | 2.3
1.862 |  | 2.1
0.034 |  | 0.1
0.087 |  | 0.1
195.01 |  | 199.1
254.04 |  | 259.3
254.25 |  | 259.6
259.97 |  | 265.5
260.11 |  | 265.6
265.38 |  | 270.9
265.79 |  | 271.4
266.04 |  | 271.6
271.58 |  | 277.3
271.91 |  | 277.7
277.44 |  | 283.3
419.65 |  | 428.4
603.93 |  | 616.6
604.01 |  | 616.6
Electrical Acceptance Tests Pass

## Slide 4: Electrical Test Data
Coil 102
RLQ |  | 
 |  | R (mΩ)
 |  | Ls (mH)
 |  | 20 Hz
 |  | 100 Hz
 |  | 1 kHz
 |  | Q
 |  | 20 Hz
 |  | 100 Hz
 |  | 1 kHz
QH R (Ω) |  | 
 | NT | A01
 | T | A02
 | T-P | B01
 | NT-P | B02
 | T-M | B03
 | NT-M | B04
VT R (mΩ) |  | 
 |  | A1
 |  | A2
 |  | A3
 |  | A4
 |  | A5
 |  | A6
 |  | A7
 |  | A8
 |  | B8
 |  | B7
 |  | B6
 |  | B5
 |  | B4
 |  | B3
 |  | B2
 |  | B1
15 |  | 16
Pre-Shipment |  | Post-Shipment
In Ship Crate |  | In Ship Crate
OD Up |  | OD Up
15 |  | 16
Pre-Shipment |  | Post-Shipment
In Ship Crate |  | In Ship Crate
OD Up |  | OD Up
609.08 |  | 615.4
4.825 |  | 4.941
3.220 |  | 3.338
1.812 |  | 1.8542
0.78 |  | 0.796
1.52 |  | 1.54
2.02 |  | 2.02
3.331 |  | 3.2
3.320 |  | 3.2
1.883 |  | 1.9
1.907 |  | 1.9
1.915 |  | 1.9
1.900 |  | 1.8
0.019 |  | 0.1
0.069 |  | 0.1
196.01 |  | 198.1
255.34 |  | 259.2
255.57 |  | 259.3
261.29 |  | 265.1
261.43 |  | 265.2
266.74 |  | 270.4
267.15 |  | 271.0
267.39 |  | 271.3
272.94 |  | 277.5
273.26 |  | 277.6
278.86 |  | 282.4
421.75 |  | 427.1
606.91 |  | 615.3
606.96 |  | 615.4
Electrical Acceptance Tests Pass

## Slide 5: Electrical Test Data
Coil 104
RLQ |  | 
 |  | R (mΩ)
 |  | Ls (mH)
 |  | 20 Hz
 |  | 100 Hz
 |  | 1 kHz
 |  | Q
 |  | 20 Hz
 |  | 100 Hz
 |  | 1 kHz
QH R (Ω) |  | 
 | NT | A01
 | T | A02
 | T-P | B01
 | NT-P | B02
 | T-M | B03
 | NT-M | B04
VT R (mΩ) |  | 
 |  | A1
 |  | A2
 |  | A3
 |  | A4
 |  | A5
 |  | A6
 |  | A7
 |  | A8
 |  | B8
 |  | B7
 |  | B6
 |  | B5
 |  | B4
 |  | B3
 |  | B2
 |  | B1
15 |  | 16
Pre-Shipment |  | Post-Shipment
In Ship Crate |  | In Ship Crate
OD Up |  | OD Up
15 |  | 16
Pre-Shipment |  | Post-Shipment
In Ship Crate |  | In Ship Crate
OD Up |  | OD Up
595.73 |  | 592.3
4.851 |  | 4.918
3.362 |  | 3.374
1.907 |  | 1.8664
0.81 |  | 0.816
1.63 |  | 1.61
2.05 |  | 1.99
3.63 |  | 3.7
3.65 |  | 3.7
1.95 |  | 2.2
1.98 |  | 2.2
2.01 |  | 2.2
1.99 |  | 2.1
0.07 |  | 0.1
0.14 |  | 0.1
192.31 |  | 191.1
250.52 |  | 249.1
250.74 |  | 249.2
256.38 |  | 254.9
256.54 |  | 255.1
261.71 |  | 260.1
262.11 |  | 260.6
262.37 |  | 260.8
267.91 |  | 266.3
268.19 |  | 266.7
273.62 |  | 272.1
413.86 |  | 411.5
595.53 |  | 592.1
595.61 |  | 592.2
Electrical Acceptance Tests Pass

## Slide 6: Electrical Test Data
Coil 105
RLQ |  | 
 |  | R (mΩ)
 |  | Ls (mH)
 |  | 20 Hz
 |  | 100 Hz
 |  | 1 kHz
 |  | Q
 |  | 20 Hz
 |  | 100 Hz
 |  | 1 kHz
QH R (Ω) |  | 
 | NT | A01
 | T | A02
 | T-P | B01
 | NT-P | B02
 | T-M | B03
 | NT-M | B04
VT R (mΩ) |  | 
 |  | A1
 |  | A2
 |  | A3
 |  | A4
 |  | A5
 |  | A6
 |  | A7
 |  | A8
 |  | B8
 |  | B7
 |  | B6
 |  | B5
 |  | B4
 |  | B3
 |  | B2
 |  | B1
15 |  | 16
Pre-Shipment |  | Post-Shipment
In Ship Crate |  | In Ship Crate
OD Up |  | OD Up
15 |  | 16
Pre-Shipment |  | Post-Shipment
In Ship Crate |  | In Ship Crate
OD Up |  | OD Up
592.7 |  | 598.2
4.695 |  | 4.801
3.146 |  | 3.246
1.812 |  | 1.812
0.76 |  | 0.783
1.54 |  | 1.578
2.01 |  | 2.024
3.610 |  | 3.9
3.597 |  | 3.9
2.012 |  | 2.3
2.036 |  | 2.4
2.028 |  | 2.3
2.029 |  | 2.3
0.008 |  | 0.1
0.055 |  | 0.1
190.21 |  | 193.1
247.79 |  | 251.6
248.02 |  | 251.8
253.58 |  | 252.5
253.74 |  | 257.6
258.86 |  | 262.8
259.29 |  | 263.1
259.52 |  | 263.3
264.94 |  | 269.0
265.27 |  | 269.3
270.67 |  | 274.6
409.36 |  | 415.3
589.08 |  | 598.0
589.13 |  | 598.1
Electrical Acceptance Tests Pass

## Slide 7: Impulse Test Data
Coil 101
Chart: Coil 101 Direct Polarity
Chart: Coil 10 Reverse Polarity
Electrical Acceptance Tests Pass

## Slide 8: Impulse Test Data
Coil 102
Chart: Coil 102 Direct Polarity
Chart: Coil 102 Reverse Polarity
Electrical Acceptance Tests Pass

## Slide 9: Impulse Test Data
Coil 104
Chart: Coil 104 Direct Polarity
Chart: Coil 104 Reverse Polarity
Electrical Acceptance Tests Pass

## Slide 10: Impulse Test Data
Coil 105
Chart: Coil 105 Direct Polarity
Chart: Coil 105 Reverse Polarity
Electrical Acceptance Tests Pass

## Slide 11: Impulse Test Data
Chart: Coil Comparison 1000V
Electrical Acceptance Tests Pass

## Slide 12: Hi Potential Tests
Coil 101 measures 10.7kΩ to RE OL Endshoe. Hi-Potential Test Fail: 1) Coil 101 → Coil 101 RE OL Endshoe 2) Coil 101 RE OL Endshoe → Coil 101 RE IL Endshoe
ALL COILS | Coil |  |  |  | Hipot Checks | Date |  |  | 
PHA 01 | 3200V/3200V (0.0μA) |  |  |  | Actual / Target (Leakage) |  |  |  | 
PHA 02 |  | PHA01 | PHA02 |  | (< 10 μA leakage) |  |  |  | 
PHB01 |  |  |  |  |  |  |  |  | 
PHB02 |  |  |  | PHB01 | PHB02 | PHB03 | PHB04 |  | 
PHB03 |  |  |  |  |  |  |  |  | 
PHB04 |  |  |  |  |  |  |  | LE IL Endshoe | RE IL Endshoe
LE IL Endshoe | 1000V/1000V (0.0μA) | 2500V/2500V (0.0μA) |  |  |  |  |  |  | 
LE OL Endshoe |  |  |  | 2500V/2500V (0.0μA) |  |  |  | 1000V/1000V (0.0μA) | 
RE IL Endshoe |  | 2500V/2500V (0.0μA) |  |  |  |  |  |  | 
RE OL Endshoe |  |  |  | 2500V/2500V (0.0μA) |  |  |  |  | 1000V/1000V (0.0μA)
Pole | 500V/500V (0.0μA) |  |  |  |  |  |  |  |