## Slide 1: JLAB QA/QC for Cryomodule
LCLS-II QA/QC Workshop
Anne McEwen / Valerie Bookwalter
25- Jan- 2017 (R11)
Notes: DRAFT

## Slide 2: Outline & Introduction
Introduction
  JLAB overall Quality Management System managed by EHS&Q
    Significant Non Conformances are reported by EHS&Q (S. Smith) per LCLS-II-1.1-QA-0563
  SRF Quality is reviewed monthly with SRF Management
Travelers / NCR Status , Pareto on Parts & Assembly
  Traveler Status – number Open/Closed
  NCR (Non Conformance) – number Open/Closed
  NCRs Pareto of types of quality issues
  Data Mining Tools & New Reports available in Pansophy
      Example CM-04 documentation Status
Note : Specific examples of Lessons Learned will be covered later in the Workshop to show how issues are closed and lessons learned integrated to Procedures & Processes for specific Parts/Issues

## Slide 3: Total * NCR Status by month (22 Jan2018) * all Cryomodules as of date of report
pCM – no change since November
Production NCRs
% NCRs closed continues to increase
97%
Closed
83%
Closed

## Slide 4: L2 pCM NCR Open/Closed by Work center
L2PRO-CM-ACTS
(CMTF Test)
L2 pCM - 8 Open NCRs (details in Appendix)
Cavity FE Onset at 12 MV/m
Couplers cannot be adjusted to proper Qext setting (4.0E7)
HOM A Couplers < 5E11
HOM B Couplers < 5E11
CM Performance

## Slide 5: pCM NCR Status (listed by Traveler )

## Slide 6: L2 PROD’N Open/Closed by Work center * all Production Cryomodules as of date of report
L2 Prd’n NCR * details in Appendix
Cavity 14 NCRs, Shock Watch
Cavity 14 NCRs various
Cavity VTRF - 64 NCRs
FPC – 6 NCRs
FPC Warm– 22 NCRs
Feed Thru – 3 NCRs
BLBP – 22 NCRs
Bellows - 68 NCRs
Tuners -7 NCRs
Instrumentation (seal) 1 NCR
* Note that a part can have more than one NCR

## Slide 7: Prod’n NCR (listed by Traveler )

## Slide 8: Open Production NCRs – Blue are on Hold
Parts “on Hold” (Blue) are typically often waiting until needed for re-work

## Slide 9: Example – Production NCRs Initiated to First Response
24
129
54
16
24
28
Last Wk
9
129
54
16
24
28
24
9
3
3

## Slide 10: NCRs Time to Closure (includes NCRs on HOLD)
33%
15%
26%
14%
Last Month
8%
4%

## Slide 11: Traveler Status
LCLS-II JLAB Cryomodule QA/QC Workshop 25Jan2018

## Slide 12: pCM Travelers - by Month &work Center
pCM Travelers closed except
ACTS = Acceptance Test (CMTF) , no change since November

## Slide 13: Production Travelers by month& work center
Travelers being closed as work is completed

## Slide 14: Pansophy Tools & Reports
LCLS-II JLAB Cryomodule QA/QC Workshop 25Jan2018

## Slide 15: Drill Down & New reports in Pansophy
Preparation for shipping
  Key Performance Summary Reports (similar to XFEL)
  By Cavity
  By String Assembly
  By Cryomodule
Tool for Traveler & NCR status by CM (example CM04)
  Status with “traffic Lights “ and option to drill down
Continuing Supporting Production, SOTRs and Quality
  Cryomodule Location in JLAB
  Vendor Hold point Data and Cavity Processing
  Cavity Performance listing
  CM Construction Status Board
  NCR Status and drill down reports

## Slide 16: Drill-Down : CMTF Test Report J1.3-04
Where to find L2 Prod’n Test Report https://pansophy.jlab.org/pansophy/
Main Menu
Data Mine
Drilldown
L2PRD
ACTS
  Select Cryomodule of interest
Cryomodule J1.3-04
Not yet shipped
Data shown in process
As of 22Jan2018

## Slide 17: Drill-Down: J1.3-04 (Part 1)
Cryomodule J1.3-04
Not yet shipped
Data shown in process
As of 22Jan2018
Shows Serial number traceability & status where available at the Cryomodule level

## Slide 18: Drill-Down: J1.3-04 (Part 2)
Shows Drill down with Weld Maps - Material Certs in SOTR folders (M Drive)
Cryomodule J1.3-04
Not yet shipped
Data shown in process
As of 22Jan2018

## Slide 19: Drill-Down: Cavity String used in J1.3-04
Cryomodule J1.3-04
Not yet shipped
Data shown in process
As of 22Jan2018
Report shows Cavity String
Serial No. , Traceability
Cavity Data & Status

## Slide 20: Drill-Down Cavity (Example Cavity 021 from J1.3-04)
Cryomodule J1.3-04
Not yet shipped
Data shown in process
As of 22Jan2018
Report Shows Cavity level traceability serial number & status

## Slide 21: Drill Down by Cryomodule Traveler /NCR Status
J1.3-04 | Open | Closed | Total | % closed | % open
Travelers * | 9 | 243 | 252 | 96% | 4%
NCRs * | 3 | 22 | 25 | 88% | 12%
Traveler/NCR status – Example for CM04 (Planned to ship soon )
Total in table above does not include the following
Travelers still in “Draft”
  Revised HOM Traveler & Procedure (in Draft)
  Shipping & Re-Work Travelers for CMs (in Draft)
Additional Supplier docs. from SOTRs on the M Drive (i.e.. Certs)
Also not included are following (no data in travelers )
  Tuners (spread sheet)
  JT Valves (10% are inspected)
  FPC Cold
  HOM (short & long vs A&B)
Through Interactive “traffic lights” in the documents Status Report , the user can “drill down” to the documents in question.
Example for J1.3-04 in the 3 pages to follow…….
Where to find L2 Traveler & NCR Status https://pansophy.jlab.org/pansophy/
Main Menu
Data Mine
Reports
L2PRD
CM Traveler 2
  Select Cryomodule of interest
* Quantity of Travelers /NCRs listed is the number instantiated
Cryomodule J1.3-04
Not yet shipped
Data shown in process
As of 22Jan2018

## Slide 22: Drill Down - CM04 Travelers/NCRs Status (Page 1)
This interactive tool is intended to assist the Project in identifying what remains to be closed for any one CM
KEY::
XX (YY )
Where
XX = Traveler sequence number
YY = NCR sequence number
: Traveler/ NCR OPEN
: Traveler/NCR CLOSED
XX & YY are links to Travelers and NCRs

## Slide 23: Drill Down - CM04 Travelers/NCR Status (Page 2)
KEY::
XX (YY )
Where
XX = Traveler sequence number
YY = NCR sequence number
: Traveler/ NCR OPEN
: Traveler/ NCR CLOSED
XX & YY are links to Travelers and NCRs
Cryomodule J1.3-04
Not yet shipped
Data shown in process
As of 22Jan2018

## Slide 24: Drill Down - CM04 Travelers/NCR Status (Page 3)
KEY::
XX (YY )
Where
XX = Traveler sequence number
YY = NCR sequence number
: Traveler/ NCR OPEN
: Traveler/ NCR CLOSED
XX & YY are links to Travelers and NCRs
Cryomodule J1.3-04
Not yet shipped
Data shown in process
As of 22Jan2018

## Slide 25: Drill Down - CM04 Travelers/NCR Status DEMO
Demonstration by Valerie
(Time permitting)

## Slide 26: Tools & Reports to Support Production
LCLS-II JLAB Cryomodule QA/QC Workshop 25Jan2018

## Slide 27: Pansophy Tools & Reports (incl. all CMs for Project)
Cryomodule Location in JLAB
Vendor Hold point Data and Cavity Processing
Cavity Performance listing
CM Construction Status Board
NCR Status and drill down reports

## Slide 28: CM Locations and Work Centers
Where to find L2 CM Locations Report https://pansophy.jlab.org/pansophy/
Main Menu
Facilities
CM Locations

## Slide 29: Vendor Hold Point Data & Cavity Processing Status
Note: blank space usually means this cavity went to FNAL

## Slide 30: Cavity Performance Listing
Where to find L2 Cavity Performance Report https://pansophy.jlab.org/pansophy/
Main Menu
Data Mine
Reports
Cavity Specification

## Slide 31: CM Construction Status Board
Where to find L2 CM Construction Status https://pansophy.jlab.org/pansophy/
Main Menu
Data Mine
Drilldown
SRF Reports
L2 CM PRD
Will be adding a column for ship traveler & date

## Slide 32: Project Overview (part 1- all Cryomodules to date)

## Slide 33: Project Overview (part 2 all cryomodules to date)
Status by Author &Traveler

## Slide 34: Summary
LCLS-II JLAB Cryomodule QA/QC Workshop 25Jan2018

## Slide 35: Summary
Quality Tools are supporting the Project
  Continuously improving the Reporting functions
QC Results
  NCRs are being addressed and closed

## Slide 36: Drill Down - CM04 Travelers/NCR Status DEMO
Demonstration by Valerie
(Time permitting)

## Slide 37: Back – UP Slides
LCLS-II JLAB Cryomodule QA/QC Workshop 25Jan2018

## Slide 38: For reference - List of Travelers
 | Traveler Name | Traveler ID | AUTHOR | REVIEWER | PROJ MNGR |  | DRAFT DUE | FINAL DUE | DRAFT DONE | FINAL DONE | NCR list | Traveler (Hold point) Closer list | NOTES
 | Inspection travelers |  |  |  |  |  |  |  |  |  |  |  | 
 | Cavity |  |  |  |  |  |  |  |  |  |  |  | 
 | Helium Vessel Inspection | L2PRD-CAV-INSP-CAV | Marhauser | Dickey | Daly |  | 1-May-16 | 15-May-16 | 9-May-16 | 9-May-16 | marhause,edaly,carpente,forehand,overtonr | marhause,edaly,carpente,forehand,overtonr | 
 | Coupler (pairs) | L2PRD-CAV-INSP-FPC | Stirbet | Reilly | Daly | Wilson | 16-Jul-16 | 30-Jul-16 | 9-Sep-16 | 9-Sep-16 | areilly,edaly,macha,kwilson,stirbet |  | 
 | Cavity Field Probe Feedthru Receiving Inspection | L2PRD-CAV-INSP-FPFT | Park | Carpenter | Daly |  | 26-Feb-16 | 11-Mar-16 | 26-Feb-16 | 26-Feb-16 | hkpark | hkpark,areilly | 
 | Cavity HOM Feedthru Receiving Inspection | L2PRD-CAV-INSP-HMFT | Park | Carpenter | Daly |  | 26-Feb-16 | 11-Mar-16 | 26-Feb-16 | 26-Feb-16 | hkpark | hkpark,areilly | 
 | Cavity RF Inspection | L2PRD-CAV-RFIN | Marhauser | Forehand | Davis |  | 10-Jun-16 | 24-Jun-16 | 17-Jun-16 | 17-Jun-16 | marhause,forehand,overtonr,edaly,carpente |  | 
 | Cavity String |  |  |  |  |  |  |  |  |  |  |  | 
 | Cavity Beam Line Bellows PRCM | L2PRD-CST-INSP-BLBP | Wilson | Carpenter | Daly |  | 6-May-16 | 20-May-16 | 6-May-16 | 6-May-16 | kwilson | kwilson | 
 | Cavity Beam Line Bellows Short | L2PRD-CST-INSP-BLBS | Wilson | Carpenter | Daly |  | 26-Apr-16 | 10-May-16 | 26-Apr-16 | 26-Apr-16 | kwilson | kwilson | 
 | String Beam Line Extension Downstream | L2PRD-CST-INSP-BLXD | Wilson | Carpenter | Daly |  | 12-Apr-16 | 26-Apr-16 | 12-Apr-16 | 12-Apr-16 | kwilson | kwilson | 
 | String Beam Line Extension Upstream | L2PRD-CST-INSP-BLXU | Wilson | Carpenter | Daly |  | 12-Apr-16 | 26-Apr-16 | 12-Apr-16 | 12-Apr-16 | kwilson | kwilson | 
 | BPM FeedThrus Receiving Inspection | L2PRD-CST-INSP-BPMFT | Park | Carpenter | Daly | Wilson | 24-Jun-16 | 8-Jul-16 | 21-Jun-16 | 24-Jun-16 | hkpark |  | 
 | 300K Current Leads CF Flange Inspection | L2PRD-CST-INSP-CFL | Hiatt | Wilson | Daly |  | 22-Aug-16 | 5-Sep-16 | 26-Aug-16 | 26-Aug-16 | hiatt |  | 
 | End Lever Tuner Mechanical Frame Inspection | L2PRD-CST-INSP-TUNMC | Huque | Carpenter | Daly |  | 15-Jun-16 | 29-Jun-16 | 15-Jun-16 | 15-Jun-16 | huque,edaly |  | 
 | End Lever Tuner Moto | L2PRD-CST-INSP-TUNMT | Huque | Carpenter | Daly |  | 15-Jun-16 | 29-Jun-16 | 13-Sep-16 | 13-Sep-16 | huque,edaly |  | 
 | Tuner Piezo Actuatir Assembly Inspection | L2PRD-CST-INSP-TUNPZ | Huque | King | Daly |  | 18-Feb-16 | 3-Mar-16 | 18-Feb-16 | 18-Feb-16 | huque,edaly |  | 
 | Qquadrapole Magnet SPQA002 Inspection | L2PRD-CST-MAG-QUAD | Hiatt | Beck | Daly |  | 29-Aug-16 | 12-Sep-16 | 7-Sep-16 | 7-Sep-16 | hiatt,edaly |  | 
 | Cryomodule |  |  |  |  |  |  |  |  |  |  |  | 
 | Bellows 2-Phase Pipe | L2PRD-CM-INSP-BP2PH | Cheng | Carpenter | Daly |  | 11-May-16 | 25-May-16 | 11-May-16 | 13-May-16 | cheng,fischer,areilly,edaly | cheng,edaly,areilly | 
 | Cryomodule Cool Down Cryogenic Valve (CDV) | L2PRD-CM-INSP-CDV | Cheng | Carpenter | Reilly |  | 14-Feb-17 | 28-Feb-17 | 20-Feb-17 | 20-Feb-17 | cheng,fischer,areilly,edaly |  | 
 | Receiving Inspection CM Disk 2-Phase Support | L2PRD-CM-INSP-DK12P | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection CM Disk 2-Phase Support | L2PRD-CM-INSP-DK22P | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Shipping End Cap Inspection | L2PRD-CM-INSP-EDCP | Huque | Fischer | Reilly |  | 2-Nov-16 | 16-Nov-16 | 3-Nov-16 | 4-Jan-17 | huque,edaly |  | 
 | Shipping Feed Cap Inspection | L2PRD-CM-INSP-FDCP | Huque | Fischer | Reilly |  | 2-Nov-16 | 16-Nov-16 | 3-Nov-16 | 4-Jan-17 | huque,edaly |  | 
 | Warm FPC Inspection | L2PRD-CM-INSP-FPCW | Stirbet | Reilly | Daly | Fischer | 28-Sep-16 | 12-Oct-16 | 28-Sep-16 | 28-Sep-16 | areilly,edaly,stirbet,fischer |  | 
 | FPC Waveguide Inspection | L2PRD-CM-INSP-FPCWG | Stirbet | Reilly | Daly | Fischer | 28-Sep-16 | 12-Oct-16 | 28-Sep-16 | 28-Sep-16 | edaly,stirbet,fischer,mdickey |  | 
 | Receiving Inspection CM Magnetic Shield Assembly | L2PRD-CM-INSP-IMAG | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection CM Rod Invar 2-Phase Long | L2PRD-CM-INSP-IROD2L | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection CM Rod Invar 2-Phase | L2PRD-CM-INSP-IROD2P | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection CM Assembly, JT Cryogenic Valve | L2PRD-CM-INSP-JTV | Cheng | Carpenter | Reilly |  | 10-Oct-16 | 24-Oct-16 | 10-Oct-16 | 10-Oct-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection CM Assembly, JT Valve Tube Kit | L2PRD-CM-INSP-JTVTK | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection CM Assembly, Liquid He Level-DS | L2PRD-CM-INSP-LLDS | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection of CM Weldment Line Liquid Level DS | L2PRD-CM-INSP-LLNDS | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection CM Weldment Line Liquid Level US | L2PRD-CM-INSP-LLNUS | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Receiving Inspection CM Assembly, Liquid He Level-US | L2PRD-CM-INSP-LLUS | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | LCLS-II Cryomodule MC Coupling Line Assembly Inspection | L2PRD-CM-INSP-MCCLA | Cheng | Carpenter | Reilly |  | 27-Jun-16 | 11-Jul-16 | 27-Jun-16 | 27-Jun-16 | cheng,fischer,areilly,edaly | cheng,edaly,areilly | 
 | Receiving Inspection CM Weldment 2-Phase Pipe End Cap | L2PRD-CM-INSP-PEC2P | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 
 | Upper Cold Mass Inspection | L2PRD-CM-INSP-UCM | Cheng | Carpenter | Reilly |  | 30-Sep-16 | 14-Oct-16 | 13-Sep-16 | 13-Sep-16 | cheng,fischer,areilly,edaly |  | 
 | Receiving Inspection CM Vacuum Vessel | L2PRD-CM-INSP-VV | Cheng | Carpenter | Reilly |  | 12-Sep-16 | 26-Sep-16 | 12-Sep-16 | 13-Sep-16 | cheng, fischer, areilly, edaly |  | 

## Slide 39: For Reference List of Travelers (cont)
 | Traveler Name | Traveler ID | AUTHOR | REVIEWER | PROJ MNGR |  | DRAFT DUE | FINAL DUE | DRAFT DONE | FINAL DONE | NCR list | Traveler (Holdpoint) Closer list | NOTES
 | Processing Travelers |  |  |  |  |  |  |  |  |  |  |  | 
 | Cavity |  |  |  |  |  |  |  |  |  |  |  | 
 | Cavity Heat Treatment & Nitrogen Doping | L2PRD-CAV-ASSY-HEAT | Forehand | Palczewski | Davis |  |  |  |  |  |  |  | 
 | Cavity Cleaning / Ultrasonic Cleaning | L2PRD-CAV-CHEM-USC | Zhao | Follkie | Reilly |  |  |  |  |  |  |  | 
 | First HPR | L2PRD-CAV-CHEM-HPR1 | Follkie | Harris | Reilly |  |  |  |  |  |  |  | 
 | Second HPR (COMBINE WITH HPR1) | L2PRD-CAV-CHEM-HPR2 | Follkie | Harris | Reilly |  |  |  |  |  |  |  | 
 | Cavity String |  |  |  |  |  |  |  |  |  |  |  | 
 | Cryomodule |  |  |  |  |  |  |  |  |  |  |  | 
 | Upper Cold Mass Magnetic Hygiene Control | L2PRD-CM-DMAG-UCM | Cheng | Legg | Reilly |  | 1-Feb-17 | 15-Feb-17 | 8-Feb-17 | 8-Feb-17 |  |  | 
 | Cryomodule Vacuum Vessel Demagnetization | L2PRD-CM-DMAG-VV | Cheng | Legg | Reilly |  | 3-Jan-17 | 17-Jan-17 | 10-Jan-17 | 10-Jan-17 |  |  | 
 | Assembly Travelers |  |  |  |  |  |  |  |  |  |  |  | 
 | Cavity |  |  |  |  |  |  |  |  |  |  |  | 
 | Cavity 1st Assembly | L2PRD-CAV-ASSY-FRST | Forehand | Palczewski | Reilly |  |  |  |  |  |  |  | 
 | Cavity 2nd Assembly | L2PRD-CAV-ASSY-SCND | Forehand | Palczewski | Reilly |  |  |  |  |  |  |  | 
 | Cavity Evacuation & Leak Test | L2PRD-CAV-ASSY-LEAK | Forehand | Palczewski | Reilly |  |  |  |  |  |  |  | 
 | Cavity String |  |  |  |  |  |  |  |  |  |  |  | 
 | Cavity String Assembly | L2PRD-CST-ASSY | Macha | Forehand | Reilly |  | 3-Oct-16 | 17-Oct-16 | 4-Oct-16 | 4-Oct-16 | forehand,kdavis,areilly,edaly,fischer |  | 
 | BPM Magnet sub-assembly | L2PRD-CST-ASSY-BPM | Macha | Forehand | Reilly |  | 3-Oct-16 | 17-Oct-16 |  |  |  |  | 
 | String Pump down, leak test and transfer to phase 1 | L2PRD-CST-ASSY-LEAK | Macha | Forehand | Reilly |  | 3-Oct-16 | 17-Oct-16 | 27-Feb-17 | 27-Feb-17 |  |  | 
 | Cryomodule |  |  |  |  |  |  |  |  |  |  |  | 
 | CM Assembly First | L2PRD-CM-ASSY-FRST | Fischer | Worland | Reilly |  | 1-Dec-16 | 15-Dec-16 | 11-Nov-16 | 11-Nov-16 |  |  | 
 | CM Assembly Second | L2PRD-CM-ASSY-SCND | Fischer | Worland | Reilly |  | 1-Dec-16 | 15-Dec-16 | 10-Feb-17 | 10-Feb-17 |  |  | 
 | CM Assembly Final | L2PRD-CM-ASSY-FINL | Fischer | Worland | Reilly |  | 1-Dec-16 | 15-Dec-16 | 8-Mar-17 | 8-Mar-17 |  |  | 
 | Cryomodule Vacuum Vessel Assembly | L2PRD-CM-ASSY-VV | Fischer | Worland | Reilly |  | 1-Dec-16 | 15-Dec-16 | 1-Mar-17 | 1-Mar-17 |  |  | 
 | Testing Travelers |  |  |  |  |  |  |  |  |  |  |  | 
 | Transfer of L2PRD 9-cell cavities to Test Stands | L2PRD-CAV-TRANS-TSTD | Forehand | Davis | Daly |  | 23-Jul-16 | 6-Aug-16 | 23-Jul-16 | 25-Jul-16 | forehand,marhause,kdavis,areilly |  | 
 | Cavity Cooldown | L2PRD-CAV-VTA-COOL | Kushnick | Reece | Reilly | Davis | 26-Jul-16 | 9-Aug-16 | 26-Jul-16 | 26-Jul-16 | kdavis,areilly,edaly |  | 
 | Cavity Production VTA HOM Survey | L2PRD-CAV-VTA-HOM | Stirbet | Reilly | Daly | Marhause | 19-Aug-16 | 19-Aug-16 | 8-Sep-16 | 8-Sep-16 |  |  | 
 | Cavity VTA Test | L2PRD-CAV-VTRF | Davis | Ari | Reilly | Marhause | 22-Jul-16 | 5-Aug-16 | 22-Jul-16 | 22-Jul-16 | kdavis,marhause,ari | marhause,ari,kdavis,stirbet,edaly,areilly | 
 | Cryomodule Acceptance Test | L2PRO-CM-CMTF-ACTS | Drury | Humphry | Reilly |  |  |  |  |  |  |  | 
 | Cryomodule Acceptance Test | L2PRO-CM-CMTF-LPRF | Drury | Humphry | Reilly |  |  |  |  |  |  |  | 
 | Cryomodule Acceptance Test | L2PRO-CM-CMTF-HPRF | Drury | Humphry | Reilly |  |  |  |  |  |  |  | 
 | Installation Travelers |  |  |  |  |  |  |  |  |  |  |  | 
 | Standard Project Travelers |  |  |  |  |  |  |  |  |  |  |  | 
 | D3 Report | L2PRD-D3 | Bookwalter | Reilly | Daly |  | 1-Jan-15 | 15-Jan-15 | 16-Feb-15 | 16-Feb-15 | bookwalt,areilly,edaly | bookwalt,areilly,edaly | 
 | Cavity Non-conformance Report | L2PRD-NCR | Bookwalter | Carpenter | Daly |  | 1-Jan-15 | 15-Jan-15 | 16-Feb-15 | 16-Feb-15 | bookwalt,areilly,edaly | bookwalt,areilly,edaly | 
 | Inspection Summary Report | L2PRD-INSR | Carpenter | DeKerlegand | Reilly |  | 21-Mar-16 | 4-Apr-16 | 21-Mar-16 | 21-Mar-16 | areilly,carpente |  | 

## Slide 40: Level of Inspection by Traveler
Traveler ID | Inspection Travelers Started/Complete |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 | No Data | Started | Complete |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 | pCM | CM2 | CM3 | CM4 | CM5 | CM6 | CM7 | CM8 | CM9 | CM10 | CM11 | CM12 | CM13 | CM14 | CM15 | CM16 | CM17 | CM18 | CM19 | CM20 | CM21 | 
L2PRD-CAV-INSP-CAV |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CAV-INSP-FPC |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CAV-INSP-FPFT |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CAV-INSP-HMFT |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CAV-RFIN |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CST-INSP-BLBP |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CST-INSP-BLBS |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CST-INSP-BLBU |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CST-INSP-BLXD |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CST-INSP-BLXU |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | no longer used starting with CM7
L2PRD-CST-INSP-BPMFT |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 3 strings worth in CR
L2PRD-CST-INSP-CFL |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CST-INSP-TUNMC |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CM-INSP-BP2PH |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CM-INSP-CDV | no pCM traveler |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-DK12P |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-DK22P |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-EDCP | no pCM traveler |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CM-INSP-FDCP | no pCM traveler |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CM-INSP-IMAG |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CM-INSP-IROD2L | Done as one traveler |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-IROD2P |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-JTV |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-JTVTK |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-LLDS |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-LLNDS |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-LLNUS |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-LLUS |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-MCCLA |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-PEC2P |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10% needed
L2PRD-CM-INSP-UCM |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
L2PRD-CM-INSP-VV |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Courtesy Brian Carpenter

## Slide 41: Back-Up pCM Detailed reports
LCLS-II JLAB Cryomodule QA/QC Workshop – 25Jan2018

## Slide 42: Cavity Status Board
Cavity String Status

## Slide 43: pCM- Open NCRs
Status Report for L2PRO-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
324 | 24-Jul-17 | CAVITY | Field Emission Onset at 12.0 MV/m Document Number: LCLSII-4.5-PP-0670-R1, 1.3 GHz Cryomodule Performance Requirements and Minimum Acceptance Criteria states Onset must be 14.0 MV/m or greater | TB9AES034 | L2PRO-CM-ACTS | 1 |  | 
325 | 24-Jul-17 | CRYOMODULE | Coupler for Cavity 2 cannot be adjusted to proper setting (4.0E7) | CM01 | L2PRO-CM-ACTS | 1 |  | 
326 | 24-Jul-17 | CRYOMODULE | Cavity 6 coupler cannot be adjusted to proper Qext setting (4.0E7) | CM01 | L2PRO-CM-ACTS | 1 |  | 
327 | 24-Jul-17 | CRYOMODULE | Cavity 7 coupler cannot be adjusted to proper Qext setting (4.0E7) | CM01 | L2PRO-CM-ACTS | 1 |  | 
328 | 24-Jul-17 | CRYOMODULE | Cavity 86 coupler cannot be adjusted to proper Qext setting (4.0E7) | CM01 | L2PRO-CM-ACTS | 1 |  | 
329 | 24-Jul-17 | CRYOMODULE | HOM A couplers Qext's on cavities 1, 5, 6 and 7 were measured to be lower than 5E11. | CM01 | L2PRO-CM-ACTS | 1 |  | 
330 | 24-Jul-17 | CRYOMODULE | HOM B couplers Qext's on cavities 5, 7 and 8 were measured to be lower than 5E11 | CM01 | L2PRO-CM-ACTS | 1 |  | 
331 | 25-Aug-17 |  | This NCR is intended to be a summary NCR listing out the major performance issues found on the LCLSII prototype cryomodule (serial # J1.3-01). This NCR could be used for submission to the SLAC Project QA as part of the significant NCR Dashboard reporting. Performance testing of the pCM was conducted in the JLab CMTF using Traveler ID L2PRO-CM-ACTS. Seven itemized issues (shown below) were reported, each having its own regular NCR recorded in the traveler database. Cavities 1, 5, 6 & 7 HOM A lower than 5E11 (NCR 329) Cavities 5, 7 & 8 HOM B lower than 5E11 (NCR 330) Cavity 6 FE Onset @ 12MV/m (NCR 324) Cavity 6 coupler could not be adjusted to proper setting 4.0E7 (NCR 326) Cavity 2 coupler could not be adjusted to proper setting 4.0E7 (NCR 325) Cavity 7 coupler could not be adjusted to proper setting 4.0E7 (NCR 327) Cavity 8 coupler could not be adjusted to proper setting 4.0E7 (NCR 328) | J1.3-01 | L2PRO-CM-ACTS | 1 |  | 

## Slide 44: Back Up Prod’n Detailed reports
LCLS-II JLAB Cryomodule QA/QC Workshop 25Jan2018

## Slide 45: Cavity Status Board – RI

## Slide 46: Cavity Status Board – RI

## Slide 47: Cavity Status Board – RI

## Slide 48: Cavity Status Board – RI

## Slide 49: Cavity Status Board - EZ

## Slide 50: Cavity Status Board - EZ

## Slide 51: Cavity Status Board - EZ

## Slide 52: Status Board - CM

## Slide 53: Open NCRs: Cavity
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
522 | 14-Oct-16 |  | On bellows restraint block, input coupler side, 1 bolt is stripped and hole needs to be re-tapped. This can be ID'd by the missing bolt. All HOMs, FP, and BL flanges are under torqued. Right angle valve rotation 50 degrees beyond what is shown on sketch in insp. traveler. | 202 | L2PRD-CAV-INSP-CAV | 9 | Proceed to next WorkCenter | 
528 | 18-Oct-16 |  | Torque less than 30 Nm on beam tube flange (short side). Sharp edge on weld near dish. Beam tube flange (long side) exceeds 30 Nm. HOM flange bolts (short side) less than 5 Nm. HOM flange bolts (long side) less than 5 Nm. Input antenna flange bolts torqued more than 10 Nm. 12 of 24 washers present. One side missing washers on beam tube (long side). | 200 | L2PRD-CAV-INSP-CAV | 10 | Proceed to next WorkCenter | 
583 | 18-Nov-16 |  | HV demag un -successful, 4G on one location | 210 | L2PRD-CAV-INSP-CAV | 14 | Proceed to next WorkCenter | 
614 | 2-Dec-16 |  | Threads bad on bellows restraint bracket. Please see attached. | 24 | L2PRD-CAV-INSP-CAV | 18 | Proceed to next WorkCenter | 5-Dec-16
617 | 5-Dec-16 |  | Nicks on corner of helium vessel tank bracket. Please see attached. | 28 | L2PRD-CAV-INSP-CAV | 19 | Proceed to next WorkCenter | 
618 | 5-Dec-16 |  | Threads bad on bellows restraint bracket. | 28 | L2PRD-CAV-INSP-CAV | 19 | Proceed to next WorkCenter | 
704 | 27-Jan-17 | CRYOMODULE | (1) attaching bolt for the bellows restraint bracket sheared off as it was being removed for a CMM measurement of the tuner attachment block. It appears that the bolt was rubbing on one side of the slot in the bracket. | Â 0022 | L2PRD-CAV-INSP-CAV | 23 | Proceed to next WorkCenter | 
801 | 4-Apr-17 |  | No burst disk installed | 55 | L2PRD-CAV-INSP-CAV | 30 |  | 
802 | 4-Apr-17 | Cavity | Burst Disc not present. | 54 | L2PRD-CAV-INSP-CAV | 31 |  | 
810 | 7-Apr-17 | Cavity | 650 +/-1 checks 651.5. | 73 | L2PRD-CAV-INSP-CAV | 34 |  | 
881 | 10-May-17 |  | Reminate magnetic field is 2.0G. Demagnetized (14) cycles, no change. | 61 | L2PRD-CAV-INSP-CAV | 45 | Proceed to next WorkCenter | 
876 | 10-May-17 |  | Reminate magnetic field is 6G. De-magged (12) cycles. Field at start was 14G | 62 | L2PRD-CAV-INSP-CAV | 46 | Proceed to next WorkCenter | 
890 | 17-May-17 | Cavity | Small damage spot on Support Block. | 53 | L2PRD-CAV-INSP-CAV | 48 |  | 
917 | 12-Jun-17 |  | Bolt cross threaded in bellows restraint bracket. Bolt won't come out of bracket. | 264 | L2PRD-CAV-INSP-CAV | 56 | Proceed to next WorkCenter | 
548 | 21-Oct-16 | LCLS-II PRODUCTION CAVITY ASSEMBLY | External shock watch activated or damaged (See photos in traveler). No Accelerometer. Extra parts missing. Documentation missing. | 200 | L2PRD-CAV-RECV-CAV | 9 |  | 
549 | 21-Oct-16 | LCLS-II PRODUCTION CAVITY ASSEMBLY | External shock watch activated or damaged (See photos in traveler). No Accelerometer. Extra parts missing. Documentation missing. | 201 | L2PRD-CAV-RECV-CAV | 10 |  | 
550 | 21-Oct-16 | LCLS-II PRODUCTION CAVITY ASSEMBLY | External shock watch activated or damaged (See photos in traveler). No Accelerometer. Extra parts missing. Documentation missing. | 202 | L2PRD-CAV-RECV-CAV | 11 |  | 
551 | 21-Oct-16 | LCLS-II PRODUCTION CAVITY ASSEMBLY | External shock watch activated or damaged (See photos in traveler). No Accelerometer. Extra parts missing. Documentation missing. | 203 | L2PRD-CAV-RECV-CAV | 12 |  | 
582 | 17-Nov-16 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer LED blinking triple RED, which indicated memory full. The last mission record was time stamped at 11/14/2016 00:53:59 UTC. Note: This NCR applies to this shipment of cavities from RI. Serial numbers are 0021, 0023, 0024, & 0028. | 21 | L2PRD-CAV-RECV-CAV | 17 |  | 
1150 | 23-Oct-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer is flashing red. | 138 | L2PRD-CAV-RECV-CAV | 115 | Proceed to next WorkCenter | 
1151 | 23-Oct-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer is flashing red. | 137 | L2PRD-CAV-RECV-CAV | 116 | Proceed to next WorkCenter | 
1153 | 23-Oct-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer is flashing red. | 135 | L2PRD-CAV-RECV-CAV | 117 | Proceed to next WorkCenter | 
1152 | 23-Oct-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer is flashing red. | 134 | L2PRD-CAV-RECV-CAV | 118 | Proceed to next WorkCenter | 
582 | 17-Nov-16 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer LED blinking triple RED, which indicated memory full. The last mission record was time stamped at 11/14/2016 00:53:59 UTC. Note: This NCR applies to this shipment of cavities from RI. Serial numbers are 0021, 0023, 0024, & 0028. | 21 | L2PRD-CAV-RECV-CAV | 17 |  | 
1150 | 23-Oct-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer is flashing red. | 138 | L2PRD-CAV-RECV-CAV | 115 | Proceed to next WorkCenter | 
1151 | 23-Oct-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer is flashing red. | 137 | L2PRD-CAV-RECV-CAV | 116 | Proceed to next WorkCenter | 
1153 | 23-Oct-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer is flashing red. | 135 | L2PRD-CAV-RECV-CAV | 117 | Proceed to next WorkCenter | 
1152 | 23-Oct-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Accelerometer is flashing red. | 134 | L2PRD-CAV-RECV-CAV | 118 | Proceed to next WorkCenter | 

## Slide 54: Open NCRs: FPC & RFIN
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
413 | 9-Sep-16 | FUNDAMENTAL POWER COUPLER PAIR | Only found packing list / inventory sheet, no certification or acceptance report. | CPI-LPA-003 | L2PRD-CAV-RECV-FPC | 1 |  | 
414 | 9-Sep-16 | FUNDAMENTAL POWER COUPLER PAIR | Only found packing list / inventory sheet, no certification or acceptance report. | CPI-LPA-013 | L2PRD-CAV-RECV-FPC | 2 |  | 
415 | 9-Sep-16 | FUNDAMENTAL POWER COUPLER PAIR | Only found packing list / inventory sheet, no certification or acceptance report. | CPI-LPA-014 | L2PRD-CAV-RECV-FPC | 3 |  | 
416 | 9-Sep-16 |  | Only found packing list / inventory sheet, no certification or acceptance report. | CPI-LPA-015 | L2PRD-CAV-RECV-FPC | 4 |  | 
417 | 9-Sep-16 |  | Only found packing list / inventory sheet, no certification or acceptance report. | CPI-LPA-016 | L2PRD-CAV-RECV-FPC | 5 |  | 
1221 | 11-Dec-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Field flatness 12.17% FM: This is not field flatness, but M-parameter in kHz, <= 10 kHz is specified. Wrong reference data have been used for vendor. M-Parameter in fact 3.47 kHz | 115 | L2PRD-CAV-RFIN | 106 | Proceed to next WorkCenter | 

## Slide 55: Open NCRs: FPCW
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
499 | 10-Oct-16 |  | Missing sapphire window. | CPI-LPA-020 | L2PRD-CM-INSP-FPCW | 17 | Proceed to next WorkCenter | 
643 | 6-Jan-17 |  | SMA electron probe instead of optical view port near to warm ceramic window | CPI-WP-013 | L2PRD-CM-INSP-FPCW | 44 | Proceed to next WorkCenter | 
644 | 6-Jan-17 |  | SMA connector for electron probe instead of optical view port near to the warm ceramic. Will be replaced during assembly on cryomodule. | CPI-WP-007 | L2PRD-CM-INSP-FPCW | 45 | Proceed to next WorkCenter | 
645 | 6-Jan-17 |  | Feed through with SMA connector instead of optical view near to the warm ceramic. Will be replaced during warm coupler assembly on cryomodule. | CPI-WP-028 | L2PRD-CM-INSP-FPCW | 46 | Proceed to next WorkCenter | 
657 | 18-Jan-17 |  | Electron probe installed on instrumentation ports near to the warm ceramic. It will be replaced with optical view poprt during CM assembly. | CPI-WP-015 | L2PRD-CM-INSP-FPCW | 47 | Proceed to next WorkCenter | 
658 | 18-Jan-17 |  | Electron probe near to the warm ceramic. Will be replaced with optical view port during CM assembly. | CPI-LPA-013 | L2PRD-CM-INSP-FPCW | 48 | Proceed to next WorkCenter | 
667 | 19-Jan-17 |  | Electron probe installed on instrumentation port near to the warm ceramic. Will e replaced with optical view port during CM assembly | CPI-LPA-013 | L2PRD-CM-INSP-FPCW | 49 | Proceed to next WorkCenter | 
668 | 19-Jan-17 |  | Electron probe installed on instrumentation port near to the warm ceramic. Will replace this with an optical view port during CM assembly. | CPI-LPA-016 | L2PRD-CM-INSP-FPCW | 50 | Proceed to next WorkCenter | 
669 | 19-Jan-17 |  | Electron probe installed on the instrumentation port near to the warm ceramic. Will be replaced with an optical view port during CM assembly. | CPI-LPA-016 | L2PRD-CM-INSP-FPCW | 51 | Proceed to next WorkCenter | 
670 | 19-Jan-17 |  | Electron probe installed on the instrumentation port near to the warm window. Will be replaced with optical view port during CM assembly. | CPI-LPA-015 | L2PRD-CM-INSP-FPCW | 52 | Proceed to next WorkCenter | 
671 | 19-Jan-17 |  | Electron probe installed on instrumentation port near to the warm ceramic. Will be replaced with an optical view port during CM assembly. | CPI-LPA-008 | L2PRD-CM-INSP-FPCW | 53 | Proceed to next WorkCenter | 
672 | 19-Jan-17 |  | Electron probe installed on instrumentation port near to the warm ceramic. Will be replaced with optical view port during CM assembly. | CPI-LPA-014 | L2PRD-CM-INSP-FPCW | 54 | Proceed to next WorkCenter | 
673 | 20-Jan-17 |  | Electron probe installed on the instrumentation port near to the warm ceramic window. Will be replaced with an optical view port during CM assembly. | CPI-LPA-007 | L2PRD-CM-INSP-FPCW | 55 | Proceed to next WorkCenter | 
674 | 20-Jan-17 |  | Electron probe installed on instrumentation port near to the warm window. Will be changed with an optical view port during CM assembly. | CPI-LPA-015 | L2PRD-CM-INSP-FPCW | 56 | Proceed to next WorkCenter | 
675 | 20-Jan-17 |  | Electron probe installed on instrumentation port near to the warm ceramic. Will be replaced with an optical view port during CM assembly | CPI-LPA-003 | L2PRD-CM-INSP-FPCW | 57 | Proceed to next WorkCenter | 
740 | 9-Feb-17 | UNKN | Electron probe installed on port near to the warm ceramic. Will be replaced with optical view port during CM assembly. | CPI-LPA-019 | L2PRD-CM-INSP-FPCW | 58 | Proceed to next WorkCenter | 
741 | 9-Feb-17 | UNKN | Brass nitrogen admission valve, non UHV compatible. | CPI-LPA-019 | L2PRD-CM-INSP-FPCW | 58 | Proceed to next WorkCenter | 
742 | 9-Feb-17 | UNKN | SMA eprobe installed on port near to the warm ceramic. Will be replaced wit optical view port during CM assembly | CPI-LPA-003 | L2PRD-CM-INSP-FPCW | 59 | Proceed to next WorkCenter | 
743 | 9-Feb-17 | UNKN | SMA feedthrough instead of optical view on port near to warm ceramic window. Will be replaced with optical view port during CM assembly. | CPI-LPA-020 | L2PRD-CM-INSP-FPCW | 60 | Proceed to next WorkCenter | 
745 | 9-Feb-17 | UNKN | eprobe with SMA connector instead of optical view port near warm ceramic. Will be replaced with optical view port during CM assembly. | CPI-LPA-004 | L2PRD-CM-INSP-FPCW | 61 | Proceed to next WorkCenter | 
752 | 13-Feb-17 | UNKN | eprobe with SMA connector installed on port near to the warm ceramic. Will e replaced with optical view port during CM assembly. | CPI-LPA-024 | L2PRD-CM-INSP-FPCW | 62 | Proceed to next WorkCenter | 
753 | 13-Feb-17 | UNKN | Found eprobe with SMA connector on the port near to the warm ceramic. Will replace with optical view port during CM assy. | CPI-LPA-024 | L2PRD-CM-INSP-FPCW | 63 | Proceed to next WorkCenter | 

## Slide 56: Open NCRs: FeedThrus
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
25 | 16-Mar-16 | Field Probe Feedthru | Roughness/machined tear out on probe surface. See traveler for photos. | 76 | L2PRD-CAV-INSP-FPFT | 54 |  | 
2 | 15-Mar-16 | HOM Feedthru | Surface roughness on probe diameter;RA measures 1.8 Surface nick on probe diameter See traveler for photos. | 61 | L2PRD-CAV-INSP-HMFT | 5 |  | 
419 | 13-Sep-16 | BPM Feedthru | Gap measurement slightly out of tolerance. | 134 | L2PRD-CST-INSP-BPMFT | 44 |  | 

## Slide 57: Open NCRs: VTRF
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
588 | 22-Nov-16 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Final 2K run, FE onset 15MV/m, lower than spec (17.5MV/m) | 201 | L2PRD-CAV-VTRF | 24 |  | 
589 | 22-Nov-16 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Final 2K run, Emax 16.6MV/m, lower than spec (19MV/m). | 201 | L2PRD-CAV-VTRF | 24 |  | 
604 | 1-Dec-16 |  | Initial 2K run, FE onset at 11.6MV/m, lower than spec(>17.5MV/m) | 210 | L2PRD-CAV-VTRF | 26 | Hold/Modify | 
605 | 1-Dec-16 |  | Initial 2K run, Q0 at 16MV/m was 2.41e10, lower than spec(>=2.5e10). | 210 | L2PRD-CAV-VTRF | 26 | Hold/Re-measure | 
622 | 12-Dec-16 |  | Initial 2K run, FE onset at 8.5 MV/m, lower than spec which is 17.5 MV/m. | 24 | L2PRD-CAV-VTRF | 30 | Hold/Modify | 
823 | 12-Apr-17 | CAVITY | FE onset at 6.5MV/m, lower than spec which is >17.5MV/m. This was the third attempt to pass Cavity-0024 at the VTA. After the field emissions in the first attempt, the cavity went through the UPW rinse to remove any potential FE sources. Following the second attempt, the cavity was disassembled and an internal inspection was conducted. A scratch on the iris at the end cell was discovered. A light EP was used to recover the cavity, which then proofed to be unsuccessful. The vendor has claimed responsibility of this cavity and as a result the cavity will be dis-positioned as Return-to-Vendor (RTV). | 24 | L2PRD-CAV-VTRF | 43 |  | 
916 | 2-Jun-17 |  | Initial Q0(16) measured 1.8700e+10, below the specification of 2.5e10 | 83 | L2PRD-CAV-VTRF | 55 |  | 
893 | 18-May-17 |  | Q0(16MV) was 2.1e10, below the specification of 2.5e10 | 63 | L2PRD-CAV-VTRF | 56 | Hold/Re-measure | 
928 | 15-Jun-17 |  | Field probe cable on test stand opened during cooldown. HOMs were mistuned. | 82 | L2PRD-CAV-VTRF | 65 | Hold/Modify | 
926 | 15-Jun-17 |  | Cavity FE onset was 15.5 MV/m, below the specification of 17.5 MV/m. | 225 | L2PRD-CAV-VTRF | 66 | Hold/Modify | 
927 | 15-Jun-17 | CAVITY | Cavity Q0(16MV) was 1.90e+10, below the specification of 2.5e10 | 225 | L2PRD-CAV-VTRF | 66 | Hold/Modify | 
930 | 16-Jun-17 |  | Q0(16)Measured 2.2000e+10, below spec 2.5e10 | 104 | L2PRD-CAV-VTRF | 67 |  | 
965 | 11-Jul-17 |  | Final FE onset was 11.7 MV/m, below the specification of 17.5 MV/m | 264 | L2PRD-CAV-VTRF | 73 | Hold/Modify | 
948 | 27-Jun-17 |  | Cavity Emax was 15.2 MV/m, below the specification of 19 MV/m | 265 | L2PRD-CAV-VTRF | 74 | Hold/Re-measure | 27-Jun-17
954 | 28-Jun-17 |  | FE onset at 8.67MV/m, lower than spec which was > 17.5MV/m. | 209 | L2PRD-CAV-VTRF | 76 | Hold/Re-measure | 25-Jul-17
955 | 28-Jun-17 |  | FE onset at 10.7MV/m, lower than spec which is > 17.5MV/m. | 211 | L2PRD-CAV-VTRF | 77 | Hold/Re-measure | 25-Jul-17
956 | 28-Jun-17 |  | Q0 at 16MV/m was 1.2800e+10, lower than spec which is 2.5e10. | 211 | L2PRD-CAV-VTRF | 77 | Hold/Re-measure | 25-Jul-17
950 | 28-Jun-17 |  | Q0 at 16MV/m was 2.38e10, lower than spec which is 2.5e10. | 109 | L2PRD-CAV-VTRF | 78 |  | 
974 | 19-Jul-17 |  | Q0(16) = 2.0000e+10, below spec of 2.5e10 | 209 | L2PRD-CAV-VTRF | 83 |  | 
976 | 19-Jul-17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | Emax = 17.9 MV/m, below spec of 19 MV/m. Also, Qo(16)=2.00e10, below spec of 2.5e10. This is captured under NCR974. | 209 | L2PRD-CAV-VTRF | 83 |  | 

## Slide 58: Open NCRs: VTRF (cont.)
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
983 | 24-Jul-17 |  | initial Q0(16) = 2.3000e+10, below spec | 268 | L2PRD-CAV-VTRF | 84 | Hold/Re-measure | 
985 | 24-Jul-17 |  | Initial Q0(16) = 2.3000e+10, below spec | 261 | L2PRD-CAV-VTRF | 85 | Hold/Modify | 
980 | 23-Jul-17 |  | Q0(16) = 1.8e10, below specification of 2.5e10 | 211 | L2PRD-CAV-VTRF | 87 |  | 
999 | 3-Aug-17 | UNKN | Onset of FE at 15.2MV/m, below specification of 18MV/m. | 272 | L2PRD-CAV-VTRF | 90 | Hold/Modify | 
1024 | 21-Aug-17 |  | Q0(16) 2.1700e+10, below spec of 2.5e10 | 281 | L2PRD-CAV-VTRF | 98 |  | 
1023 | 21-Aug-17 |  | Q0(16MV/m)= 2.1000e+10, below spec of 2.5e10 | 279 | L2PRD-CAV-VTRF | 99 |  | 
1027 | 22-Aug-17 |  | Q0(16) was 2.4000e+10, below spec of 2.5e10 | 118 | L2PRD-CAV-VTRF | 101 |  | 
1028 | 22-Aug-17 |  | FInal FE onset was 11 MV/m, below the spec of 17.5 MV/m | 118 | L2PRD-CAV-VTRF | 101 | Hold/Modify | 
1035 | 25-Aug-17 |  | Qextfp was <2.5e11 | 261 | L2PRD-CAV-VTRF | 103 |  | 
1064 | 1-Sep-17 |  | FE onset limit below 17.5MV/m. Field Emission onset at 10.2MV/m. | 238 | L2PRD-CAV-VTRF | 107 |  | 
1075 | 6-Sep-17 |  | Cavity Cavity Q below spec at 2.15e10 at 16MV/m | L2-133 | L2PRD-CAV-VTRF | 110 |  | 
1076 | 6-Sep-17 |  | Cavity had Low Qo | L2-131 | L2PRD-CAV-VTRF | 110 |  | 
1077 | 6-Sep-17 |  | Cavity had low Qo | L2-130 | L2PRD-CAV-VTRF | 110 |  | 
1079 | 7-Sep-17 |  | Qext2 was 1.6200e+11, lower than the spec of 2.5-7.0 e11 | 286 | L2PRD-CAV-VTRF | 111 |  | 
1093 | 13-Sep-17 |  | Failure in Field probe cable connection. | 288 | L2PRD-CAV-VTRF | 116 |  | 
1097 | 18-Sep-17 |  | Q2 is below specification. | L2-268 | L2PRD-CAV-VTRF | 117 |  | 
1099 | 18-Sep-17 | Cavity | Cavity max gradient 17.3MV/m. | L2-268 | L2PRD-CAV-VTRF | 117 |  | 
1098 | 18-Sep-17 | UNKN | Field Probe Q was low at 1.8e11 | 288 | L2PRD-CAV-VTRF | 118 |  | 
1116 | 26-Sep-17 |  | Field probe Q was below spec at 1.89e11 | 128 | L2PRD-CAV-VTRF | 124 |  | 
1117 | 26-Sep-17 |  | Quench at 17.85 MV/m | 128 | L2PRD-CAV-VTRF | 124 |  | 
1120 | 29-Sep-17 | Cavity | Cavity quenched at 15MV/m. | L2-284 | L2PRD-CAV-VTRF | 125 |  | 
1136 | 11-Oct-17 |  | Dewar pressure during test was 18 or 19 Torr | 276 | L2PRD-CAV-VTRF | 128 |  | 
1137 | 11-Oct-17 |  | field probe Q was low at 1.25e11 | 276 | L2PRD-CAV-VTRF | 128 |  | 
1161 | 2-Nov-17 |  | Field probe Q slightly high. We should edit the traveler if we are going to increase this value in the long term. | 299 | L2PRD-CAV-VTRF | 132 |  | 
1165 | 2-Nov-17 |  | FE onset at 3.5MV/m during decay measurement. | 304 | L2PRD-CAV-VTRF | 133 |  | 
1173 | 9-Nov-17 |  | Qext2 was 1.9500e+11, out of spec which is 2.5-7.0 e11 | 305 | L2PRD-CAV-VTRF | 136 |  | 
1174 | 9-Nov-17 |  | Q0 at 16MV/m was 2.28e10, out of spec which is >= 2.5e10 | 305 | L2PRD-CAV-VTRF | 136 |  | 
1172 | 9-Nov-17 |  | Qext2 was 1.9900e+11, out of spec which is 2.5-7.0 e11. | 308 | L2PRD-CAV-VTRF | 137 |  | 
1180 | 16-Nov-17 |  | Qext2 was 2.3800e+11, out of spec which is 2.5-7.0 e11. | 305 | L2PRD-CAV-VTRF | 140 |  | 
1181 | 16-Nov-17 |  | Final FE onset at 11.4MV/m, lower than spec which is > 17.5 MV/m. | 305 | L2PRD-CAV-VTRF | 140 |  | 
1203 | 30-Nov-17 |  | Field probe Q is low at 2.13e11 | 135 | L2PRD-CAV-VTRF | 143 |  | 
1212 | 5-Dec-17 |  | Emax was 13.6MV/m, lower than spec which is >19MV/m. | 293 | L2PRD-CAV-VTRF | 144 |  | 
1213 | 5-Dec-17 |  | Qext2 was 2.1700e+11, out of spec which is 2.5-7.0 e11. | 293 | L2PRD-CAV-VTRF | 144 |  | 
1214 | 5-Dec-17 |  | FE onset 11.2MV/m, lower than spec which is >17.5MV/m. | 314 | L2PRD-CAV-VTRF | 146 |  | 
1215 | 5-Dec-17 |  | Q0 at 16MV/m was 2.35e10, lower than spec which is 2.5e10 | 314 | L2PRD-CAV-VTRF | 146 |  | 
1216 | 5-Dec-17 |  | Qext2 was 2.4100e+11, out of spec which is 2.5-7.0 e11 | 314 | L2PRD-CAV-VTRF | 146 |  | 
1218 | 6-Dec-17 |  | Cavity pressure was 5.3e-5 mBar at the beginning of the test | 137 | L2PRD-CAV-VTRF | 147 |  | 
1219 | 6-Dec-17 |  | field probe Q was low at 2e11 | 137 | L2PRD-CAV-VTRF | 147 |  | 
1222 | 15-Dec-17 |  | KLorentz was -2.3828 Hz/(MV/m)^2, out of spec which is -0.6~-1.2 Hz/(MV/m)^2 | 328 | L2PRD-CAV-VTRF | 148 |  | 
1223 | 15-Dec-17 |  | Qext2 was 2.4300e+11, slightly lower than spec which is 2.5-7.0 e11 | 328 | L2PRD-CAV-VTRF | 148 |  | 
1231 | 21-Dec-17 |  | Qext2 was 2.4300e+11, slightly out of spec which is 2.5-7.0 e11. | 332 | L2PRD-CAV-VTRF | 149 |  | 
1232 | 21-Dec-17 |  | Qext2 was 2.1700e+11, out of spec which is 2.5-7.0 e11 | 331 | L2PRD-CAV-VTRF | 150 |  | 
1233 | 21-Dec-17 |  | Initial Q0 at 16MV/m was 1.7500e+10, out of spec which is >= 2.5e10 | 331 | L2PRD-CAV-VTRF | 150 |  | 
1240 | 19-Jan-18 |  | final FE onset 16 MV/m | L2-144 | L2PRD-CAV-VTRF | 156 | Hold/Modify | 

## Slide 59: Open NCRs: Cryomodule
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
1101 | 19-Sep-17 |  | Qext value for HOM1 (cplr side) Cavities 1, 2, 3, 5, 6 do not meet the Acceptance criteria | J1.3-02 | L2PRD-CM-ACTS | 1 | Hold/Modify | 
1102 | 19-Sep-17 | Cryomodule | Qext's of HOM2 for cavities 2 and 4 do not meet Acceptance criteria | J1.3-02 | L2PRD-CM-ACTS | 1 | Hold/Modify | 
1103 | 19-Sep-17 | LCLLS II 1.3 GHz Cavity | Cavity 4 in this cryomodule has a FE Onset gradient of 3.7 MV/m. This does not meet Acceptance criteria | J1.3-02 | L2PRD-CM-ACTS | 1 |  | 
1104 | 19-Sep-17 |  | Cavity 4 in this cryomodule is limited by quench at 8.0 MV/m and does not meet Acceptance criteria | J1.3-02 | L2PRD-CM-ACTS | 1 |  | 
1106 | 19-Sep-17 | HOM 1 Coupler | HOM1 Qext's for Cavities 2,4,5,7,8 do not meet Acceptance criteria | J1.3-03 | L2PRD-CM-ACTS | 2 | Hold/Modify | 
1107 | 19-Sep-17 |  | HOM 2 Qext's do not meet Acceptance criteria for cavities 2,4,5,6, and 8. | J1.3-03 | L2PRD-CM-ACTS | 2 | Hold/Modify | 
1138 | 11-Oct-17 | Upper Cold Mass Assembly | Flange "A" DS End: Small nick in conflat knife edge.(marked) | RFCG-UC-020 | L2PRD-CM-INSP-UCM | 6 |  | 
1134 | 10-Oct-17 | Vacuum Vessel | Paint residue/scratches found on flange face: Items listed below are marked with permanent marker where concerns exist for further review. F10018083:Top end ports light scratches on flange. F10020935 Centerline Port:Paint and minor nicks on knife edge. F10023641 Fixed Flange US: Paint in seal area. F10023655 Coupler Instrumentation Port:Paint/Over spray on flange face/most of entire surface. (8 flanges) F10023650 Coupler Ports:Minor scratches/oxidation,dings. Flange 1,5,6,7,8. F10023643 Sliding Flange US:Paint present in O-ring grooves. (both) F1006622 Body Side Port: Small paint spot on flange face. F10030285 Instrumentation Flange:Small paint spot on flange face 1,3. F10030301 Instrumentation Flange: Small paint spots on flange face Flange 1,2. F10026614 Tuner Motor Ports: Minor paint spots on flange face 6,7,8 See attached inspection report for detail of inspection results. | RFCG-VV-008 | L2PRD-CM-INSP-VV | 8 |  | 
1135 | 10-Oct-17 | Vacuum Vessel | F10023636 Fixed Flange (DS):Damage to edge of flange causing some surface irregularity to flange face. Some minor scratches.(See photo attached) | RFCG-VV-008 | L2PRD-CM-INSP-VV | 8 |  | 
1149 | 20-Oct-17 | Vacuum Vessel | See attached inspection report and inspection photos above for detail of inspection results. Paint residue/scratches found on flange face: Items listed below are marked with permanent marker where concerns exist for further review. F10018068 Middle Port: Scratches on flange(marked) F10020935 Centerline Port: Paint on flange and knife edge.(marked) F10023650 Coupler Port: Light scratches on flange. (marked) F10023643 Sliding Flange: Generally good. Two small areas of discoloration/paint residue in O-ring Grooves.(marked) F10030285 Instrumentation Flange: Paint on flange. (marked) F10030301 Instrumentation Flange: Scratches.(marked) F10026614 Tuner Motor Flanges: Light paint spots on flange.(marked) F10030811 Relief Port Flange: Paint spot in O-ring groove.(marked) | RFCG-VV-018 | L2PRD-CM-INSP-VV | 9 |  | 
1236 | 8-Jan-18 | Vacuum Vessel | See attached inspection report above for detail of inspection results. Paint residue/light scratches found on flange face: Items listed below are marked with permanent marker where concerns exist for further review. Fixed Flange (US):F10023641 Paint found on flange surface.(marked) Coupler Instrumentation Port: F10023655 Paint and minor scratches found on flange 4,5,6,8.(marked) Tuner Motor Ports: F10026614 Light scratches on flange face; 3,4,5,7. (marked) | RFCG-VV-022 | L2PRD-CM-INSP-VV | 10 |  | 

## Slide 60: Open NCRs: Bellows - BLBP
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
88 | 18-May-16 | unkn | Scratches present in seal groove. | 1010 | L2PRD-CST-INSP-BLBP | 10 | Hold/Modify | 
90 | 18-May-16 | Bellow Long | Scratches present in seal groove. | 1012 | L2PRD-CST-INSP-BLBP | 12 | Hold/Modify | 
92 | 18-May-16 | Bellow Long | Scratches present in seal groove. Tool chatter present in seal groove. Contamination present in seal groove. | 1013 | L2PRD-CST-INSP-BLBP | 13 | Hold/Modify | 
93 | 18-May-16 | Bellow Long | Contamination spots in seal groove. Scratches present in seal groove. | 1014 | L2PRD-CST-INSP-BLBP | 14 | Hold/Modify | 
95 | 18-May-16 | Bellow Long | Scratches present in seal groove. | 1015 | L2PRD-CST-INSP-BLBP | 15 | Hold/Modify | 
96 | 18-May-16 | Bellow Long | Scratches present in seal groove.Contamination spot in seal groove. | 1016 | L2PRD-CST-INSP-BLBP | 16 | Hold/Modify | 
98 | 18-May-16 | Bellow Long | Scratches present in seal groove.Contamination spots in seal groove.Tool chatter present in seal groove. | 1017 | L2PRD-CST-INSP-BLBP | 17 | Hold/Modify | 
99 | 18-May-16 | Bellow Long | Scratches present in seal groove. | 1018 | L2PRD-CST-INSP-BLBP | 18 | Hold/Modify | 
102 | 31-May-16 | Bellow Long | Scratches on rotatable flange at 11:00. Scratches on non-rotatable flange at 10:00. | 1020 | L2PRD-CST-INSP-BLBP | 20 | Hold/Modify | 
103 | 31-May-16 | Bellow Long | Scratches on non-rotatable flange at 12:00. | 1023 | L2PRD-CST-INSP-BLBP | 23 | Hold/Modify | 
219 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .040-.140 at welds. | 2002 | L2PRD-CST-INSP-BLBP | 37 | Hold/Modify | 
220 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .050-.190 at welds. | 2003 | L2PRD-CST-INSP-BLBP | 38 | Hold/Modify | 
221 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .035-.075 at welds. | 2004 | L2PRD-CST-INSP-BLBP | 39 | Hold/Modify | 
222 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .054-.080 at welds. | 2005 | L2PRD-CST-INSP-BLBP | 40 | Hold/Modify | 
225 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .065-.190 at welds. | 2010 | L2PRD-CST-INSP-BLBP | 45 | Hold/Modify | 
226 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .100 at welds. | 2011 | L2PRD-CST-INSP-BLBP | 46 | Hold/Modify | 
227 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .130-.145 at welds. | 2012 | L2PRD-CST-INSP-BLBP | 47 | Hold/Modify | 
228 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .075-.260 at welds. | 2013 | L2PRD-CST-INSP-BLBP | 48 | Hold/Re-measure | 
229 | 25-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .050-.145 at welds. | 2014 | L2PRD-CST-INSP-BLBP | 49 | Hold/Re-measure | 
230 | 26-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .075-.150 | 2015 | L2PRD-CST-INSP-BLBP | 50 | Hold/Modify | 
232 | 26-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .065-.255 | 2017 | L2PRD-CST-INSP-BLBP | 52 | Hold/Modify | 
234 | 26-Jul-16 | Bellow Long | Remnant magnetic field measurement exceeds requirement. Checks .070-.280 | 2019 | L2PRD-CST-INSP-BLBP | 54 | Hold/Modify | 

## Slide 61: Open NCRs: Bellows - BLBP (cont.)
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
253 | 1-Aug-16 | Bellow long | Remnant magnetic field checks .080-.280. | 2021 | L2PRD-CST-INSP-BLBP | 56 | Hold/Modify | 
258 | 1-Aug-16 | Bellow long | Remnant magnetic field checks .288-.298. | 2026 | L2PRD-CST-INSP-BLBP | 61 | Hold/Modify | 
1008 | 14-Aug-17 | Bellow PRD | Blister. Pitting. Dark spots/stains. | 2034 | L2PRD-CST-INSP-BLBP | 69 | Hold/Modify | 15-Nov-17
1036 | 25-Aug-17 | Bellow PRD | Blister,dark stains. | 2035 | L2PRD-CST-INSP-BLBP | 70 | Hold/Modify | 
1039 | 25-Aug-17 | Bellow PRD | Pitts in plating at cuff ID. | 2041 | L2PRD-CST-INSP-BLBP | 76 |  | 
386 | 19-Aug-16 |  | Weld protrudes out of cuff. Please see attached. Surface finish out of tolerance, NR flange measured 0.9 microns | 2047 | L2PRD-CST-INSP-BLBP | 82 | Hold/Modify | 
553 | 21-Oct-16 |  | Scratches present. | 2053 | L2PRD-CST-INSP-BLBP | 88 | Hold/Modify | 
554 | 21-Oct-16 |  | Scratches present in seal area. | 2059 | L2PRD-CST-INSP-BLBP | 96 | Hold/Modify | 
555 | 21-Oct-16 |  | Scratches present in seal area. | 2061 | L2PRD-CST-INSP-BLBP | 98 | Hold/Modify | 
423 | 19-Sep-16 | Bellow long | Part was returned to supplier for rework; these are re-inspection results after rework. Edge burrs corrected. Light scratches still present in seal groove NR flange. Slight tool chatter still present in seal groove NRF. | 2001 | L2PRD-CST-INSP-BLBP | 100 | Hold/Modify | 
424 | 19-Sep-16 | Bellow long | Part was returned to supplier for rework; these are re-inspection results after rework. Edge burrs corrected. Light scratches still present in seal groove NR flange. Slight tool chatter still present in seal groove NRF. | 2002 | L2PRD-CST-INSP-BLBP | 101 | Hold/Modify | 
427 | 19-Sep-16 | Bellow long | Part was returned to supplier for rework; these are re-inspection results after rework. Edge burrs corrected.Light scratches still present in seal groove NRF. | 2011 | L2PRD-CST-INSP-BLBP | 108 | Hold/Modify | 
429 | 19-Sep-16 | Bellow long | Part was returned to supplier for rework; these are re-inspection results after rework. Edge burrs corrected. Light scratches still present in seal groove NRF. | 2012 | L2PRD-CST-INSP-BLBP | 109 | Hold/Modify | 
435 | 19-Sep-16 | Bellow long | Part was returned to supplier for rework; these are re-inspection results after rework. Edge burrs corrected.Light scratches still present in seal groove NRF. | 2017 | L2PRD-CST-INSP-BLBP | 114 | Hold/Modify | 
436 | 19-Sep-16 | Bellow long | Part was returned to supplier for rework; these are re-inspection results after rework. Edge burrs corrected. Light scratches still present in seal groove NRF. | 2019 | L2PRD-CST-INSP-BLBP | 116 | Hold/Modify | 
443 | 21-Sep-16 | Bellow long | Part was returned to supplier for rework; these are re-inspection results after rework. Edge burrs corrected.Very light scratches still present RF. | 2026 | L2PRD-CST-INSP-BLBP | 122 | Hold/Modify | 
468 | 23-Sep-16 |  | Light scratches in seal groove on rotatable flange. Surface finish out of tolerance. Rotatable flange measured 1.0 microns | 2069 | L2PRD-CST-INSP-BLBP | 139 | Hold/Modify | 

## Slide 62: Open NCRs: Bellows - BLBP (cont.)
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
469 | 23-Sep-16 |  | Scratches in seal groove on NR flange. Surface finish measured 1.0 microns on RF. | 2070 | L2PRD-CST-INSP-BLBP | 140 | Hold/Modify | 
472 | 23-Sep-16 |  | Light scratches in seal groove on NR flange. Finish measured 1.0 microns on rotatable flange. | 2073 | L2PRD-CST-INSP-BLBP | 143 | Hold/Modify | 
473 | 23-Sep-16 |  | Light scratches in seal groove of rotatable flange. Surface finish = 1.0 rotatable flange. | 2074 | L2PRD-CST-INSP-BLBP | 144 | Hold/Modify | 
1155 | 24-Oct-17 | WELDMENT BELLOWS PRCM | Blisters.Irregular spot in plating.Light water spots. | 2063 | L2PRD-CST-INSP-BLBP | 145 | Hold/Modify | 
466 | 22-Sep-16 |  | Scratches on NR flange. Surface finish out of tolerance. Rotatable flange measured 1.0 microns | 2067 | L2PRD-CST-INSP-BLBP | 149 | Hold/Modify | 
574 | 14-Nov-16 | Bellow PRD | Contamination spots in seal groove. | 1001 | L2PRD-CST-INSP-BLBP | 173 | Hold/Modify | 5-Oct-17
1060 | 31-Aug-17 | Bellow PRD | Blisters,Dark stains/spots. | 2131 | L2PRD-CST-INSP-BLBP | 230 | Hold/Modify | 
1158 | 24-Oct-17 | Bellow PRD | One small blister. Otherwise good. | 2189 | L2PRD-CST-INSP-BLBP | 303 |  | 
1182 | 21-Nov-17 | Bellow PrCM | Multiple dark spots and stains. | 2191 | L2PRD-CST-INSP-BLBP | 305 | Hold/Modify | 
1183 | 21-Nov-17 | Bellow PrCM | Multiple dark spots and stains. | 2192 | L2PRD-CST-INSP-BLBP | 306 | Hold/Modify | 
1184 | 21-Nov-17 | Bellow PrCM | Multiple dark spots and stains. | 2193 | L2PRD-CST-INSP-BLBP | 307 | Hold/Modify | 
1185 | 21-Nov-17 | Bellow PrCM | Multiple dark spots and stains. | 2194 | L2PRD-CST-INSP-BLBP | 308 | Hold/Modify | 
1186 | 21-Nov-17 | Bellow PrCM | Multiple dark spots and stains. | 2195 | L2PRD-CST-INSP-BLBP | 309 | Hold/Modify | 
1194 | 27-Nov-17 | Bellow PRD | Dark spots,Pitting,irregular spot in plating, | 2197 | L2PRD-CST-INSP-BLBP | 311 |  | 
1195 | 27-Nov-17 | Bellow PRD | One blister,multiple dark spots and stains. | 2198 | L2PRD-CST-INSP-BLBP | 312 | Hold/Modify | 
1196 | 27-Nov-17 | Bellow PRD | Blister, minimum dark spots. | 2199 | L2PRD-CST-INSP-BLBP | 313 | Hold/Modify | 
1197 | 27-Nov-17 | Bellow PRD | Multiple dark spots and stains. | 2201 | L2PRD-CST-INSP-BLBP | 315 | Hold/Modify | 
1199 | 28-Nov-17 |  | Many dark spots present on plating. | 2207 | L2PRD-CST-INSP-BLBP | 321 | Hold/Modify | 
1200 | 29-Nov-17 |  | Dark spots on plating. | 2210 | L2PRD-CST-INSP-BLBP | 324 | Hold/Modify | 
1201 | 29-Nov-17 |  | Dark spots on copper plating. | 2211 | L2PRD-CST-INSP-BLBP | 325 | Hold/Modify | 
1171 | 7-Nov-17 | WELDMENT BELLOWS PRCM | Heavy oxidation on plated surfaces.Plating is dull as compared to earlier inspections. Stains present. This bellows has been issued and assembled to a cavity and HPR'd. Inspection is after dis-assembly. | 2068 | L2PRD-CST-INSP-BLBP | 345 |  | 

## Slide 63: Open NCRs: Bellows – BLBS, BLXU & BLXD
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
73 | 2-May-16 | unkn | Scratches present in seal groove. | 1002 | L2PRD-CST-INSP-BLBS | 2 | Hold/Modify | 
74 | 2-May-16 | unkn | Scratches present in seal groove. | 1003 | L2PRD-CST-INSP-BLBS | 3 | Hold/Modify | 
53 | 22-Apr-16 | unkn | scratch in seal area on rotatable flange. | 1003 | L2PRD-CST-INSP-BLXD | 3 | Hold/Modify | 
494 | 6-Oct-16 | Weldment Extension DS | Scratches present in seal groove NRF.Scratches present in seal area RF. Interior of pipe very rough surface.(Required 3.2 checks 2.5-2.7) Roughness required .8 in seal. (checks 1.3-2.0)NRF | 2003 | L2PRD-CST-INSP-BLXD | 10 | Hold/Modify | 
725 | 6-Feb-17 | Weldment Extension DS | Dark spot in plating.Minor pitting/roughness in plating surface. | 2007 | L2PRD-CST-INSP-BLXD | 15 |  | 
1010 | 14-Aug-17 |  | Weld splatter. Approximately 25 places. | 2008 | L2PRD-CST-INSP-BLXD | 16 |  | 
1126 | 3-Oct-17 | Weldment Extension DS | Dark stains. | 2020 | L2PRD-CST-INSP-BLXD | 43 | Hold/Modify | 3-Oct-17
1208 | 4-Dec-17 | Weldment Extension DS | Base metal defects. | 2023 | L2PRD-CST-INSP-BLXD | 45 |  | 

## Slide 64: Open NCRs- TUNMC & INSR
Status Report for L2PRD-NCR on 22-Jan-18 |  |  |  |  |  |  |  | 
TRAV_SEQ_NUM | TRAV_OPEN_DATE | PARTDESCRIPTION | DESCRIBECOMMENT | SN | TRAVELERID | TRAVSEQNUMBER | DISPOSITION | REWORKDATE
1224 | 18-Dec-17 | End lever tuner | Missing F10031141 Qty.2 Item 4 drawing F10008766. Note: Rod are being modified per Naeem. New configuration rods not available for this tuner at this time. | 138 | L2PRD-CST-INSP-TUNMC | 62 |  | 
1225 | 18-Dec-17 | End lever tuner | Missing F10031141 Qty.2 Item 4 drawing F10008766. Note: Rod are being modified per Naeem. New configuration rods not available for this tuner at this time. | 140 | L2PRD-CST-INSP-TUNMC | 63 |  | 
1226 | 18-Dec-17 | End lever tuner | Missing F10031141 Qty.2 Item 4 drawing F10008766. Note: Rod are being modified per Naeem. New configuration rods not available for this tuner at this time. | 142 | L2PRD-CST-INSP-TUNMC | 65 |  | 
1227 | 19-Dec-17 | End lever tunner | Missing F10031141 Qty.2 Item 4 drawing F10008766. Note: Rod are being modified per Naeem. New configuration rods not available for this tuner at this time. | 143 | L2PRD-CST-INSP-TUNMC | 66 |  | 
1228 | 19-Dec-17 | End lever tuner | Missing F10031141 Qty.2 Item 4 drawing F10008766. Note: Rod are being modified per Naeem. New configuration rods not available for this tuner at this time. | 144 | L2PRD-CST-INSP-TUNMC | 67 |  | 
1229 | 20-Dec-17 | End lever tuner | Angle 5+/-1 degree checks 6.2 on F10006870 Tuner support motor side. | 145 | L2PRD-CST-INSP-TUNMC | 68 |  | 
1230 | 20-Dec-17 | End lever tuner | Missing F10031141 Qty.2 Item 4 drawing F10008766. Note: Rod are being modified per Naeem. New configuration rods not available for this tuner at this time. | 145 | L2PRD-CST-INSP-TUNMC | 68 |  | 
1237 | 10-Jan-18 | UNKN | Lot 1: Sn: B01-B10 Flange seal groove inspection shows tool marks which may present a leak path. Drawing specifies concentric tool lay. Does not comply with this requirement. Appears to be a milled surface. See attachment for inspection results summary. Lot 2: Sn:B11-B20 Flange seal groove inspection shows scratches which may present a leak path. Sn: B11,B13,B15,B16,B18,B20 Surface finish requirement exceeds requirement of .4 RA. SN:B14,B17,B19 Scratch outside seal zone. OK See attachment for inspection results summary. | ?NA | L2PRD-INSR | 11 |  |