## Slide 1: F1.3-05 Beam Line Leak

## Slide 2: Beam Line Vacuum During Shipping
Actual
Scaled

## Slide 3: Beam Line Vacuum Leak
Cavity | Sound Detection | Note
CAV1 | Apparent leaking sound | Coupler Bellows area sounded louder than cavity bellows
CAV2 | Background | 
CAV3 | Background | 
CAV4 | Background | 
CAV5 | Background | 
CAV6 | Background | 
CAV7 | Background | 
CAV8 | Background | 
Down stream manifold has background sound level
Vacuum pressure was ~86 torr when sound detection was conducted.
  Pumping was through downstream shipping manifold
Vacuum pressure was stalled at 86 torr.

## Slide 4: Beam Line Vacuum Leak
CAV1
  When coupler antenna was pulled out ~2.5 mm (2 full turns), pressure started to increase. Sound level increased also.
  When coupler antenna was restored to original position, pressure resumed normal decreasing at a slow rate.
CAV5
  When coupler antenna was moved out and back in three full turns, pressure did not register any change as in CAV1.

## Slide 5
Picture from Rich Motill