## Slide 1: Dune Detector Safety System
Dune Collaboration Meeting
January 30, 2019
T. Shaw

## Slide 2: Dune Detector Safety System (DDSS)
The Dune Detector Safety System (DDSS) functions to protect experimental equipment.
The DSS is required to detect abnormal and potentially harmful operating conditions. It must recognize when conditions are not within the bounds of normal operating parameters and automatically take pre-defined protective actions.
The DSS must communicate to the Facility Information Reporting Utility System (FIRUS) and the Slow Controls System which monitors the detector status. Through communication links and working together, the three distinct systems provide the ability to monitor the status of the experiment, protect the equipment and provide life safety.

## Slide 3: Block Diagram DDSS Interaction

## Slide 4: HV Consortia Detector Safety
HV power supplies interlock with the liquid argon level (well known)
shut off HV power supplies when a sudden pressure drop inside the cryostat is detected (such as when the pressure relief value is open)
  A sudden pressure drop will likely cause sudden boiloff in the liquid which might trigger HV discharge
Keep HV power supplies off for a period of time (TBD) when starting the cryogenic recirculation
  We had experienced increase HV instabilities right after the start of LAr pumps in NP04. It is likely that debris was stirred up by the sudden change of flow pattern.
Plan to work with slow control to implement automatic streamer detection and suppression script.

## Slide 5: TPC Consortia Detector Safety
Hardware interlocks will be put in place such that detector components cannot be operated or even powered up unless the conditions are safe both for the detector and for personnel.
Interlocks will be used on all the low voltage power and on the bias voltage supplies, including inputs from environmental monitors both inside and outside the cryostat.
Examples of interlocks that will be used include
  turning off the power to the WIECs if the corresponding cooling fans are not operational or if the temperature inside the crates exceeds a pre-set value.
  Similar interlocks will be used for the low voltage power and the bias voltage supplies in the CE racks.
  Interlocks may be required connecting the value of the bias voltage on the field cage termination electrodes to the high voltage applied on the TPC cathode.
  Interlocks will turn off the transmitters on the WIECs if the readout fibers bundles are cut.
  Interlocks on local smoke detection at feedthroughs to power.

## Slide 6: Calibration Consortia Detector Safety

## Slide 7: Calibration Consortia Detector Safety

## Slide 8: Calibration Consortia Detector Safety

## Slide 9: Calibration Consortia Detector Safety

## Slide 10: Other Consortia – need input
APA
CISC
DAQ
  Eye Safety for fibers
  Smoke Detection in DAQ Racks
  Leak Detection in Rack Cooling System
  Temperature in Racks
PD
Cryogenics – See David Montanari’s presentation