## Slide 1
M. Marchevsky, E. Hershkovitz, S. Prestemon (LBNL)
J. Weiss and D. van der Laan (ACT)
Probing current distribution in CORC with Hall arrays
MDP 2019 Collaboration Meeting, Jan. 16 2019

## Slide 2: Outline
Current sharing and quench detection: two immediate problems to address for use of CORC in HEP and fusion applications
Hall array instrumentation for evaluating current sharing, measuring quench propagation velocity and detecting a quench
Experimental results
Conclusions

## Slide 3: Recent progress and challenges in developing CORC ® conductors for accelerator applications ®
  The first CORC® Canted Cosine Theta dipole was demonstrated at LBNL (1.1 T at 4.2 K)
Wang et al., A viable dipole magnet concept with REBCO CORC® wires and further development needs for high-field magnet applications, Supercond. Sci. Technol. https://doi.org/10.1088/1361-6668/aaad8f
  CORC® wires reach record in-field current densities beyond 400 A/mm2 at 20 T
Recovering and stationary normal zones have been experimentally demonstrated. Also significant inhomogeneity in current distribution between tapes was observed starting at I/Ic~0.3 (77 K, self-field)
M. Majoros et al., “Stability and normal zone propagation in
YBCO CORC® cables “, Supercond. Sci. Technol. 29 (2016) 044006, doi:10.1088/0953-2048/29/4/044006
Understanding stability limits and developing adequate quench detection/protection techniques is crucial for accelerator magnet applications of CORC ®

## Slide 4: Measuring current sharing and quench propagation with a Hall sensor
Normally, current sharing can be studied by measuring voltage drop across individual tapes in the cable. Alternatively one can measure magnetic field re-distribution.
Bs
Bx
Bx
z
Z-axis Hall sensor
Bs
Bx
Bx
z
Z-axis Hall sensor
Re-distribution of current flow away from its solenoidal path can be detected by appearance of the transverse field component in the vicinity of a normal zone boundary
As real CORC® conductor has equal number of tapes wound clockwise and counter-clockwise, the solenoidal component is largely compensated. But whenever current distribution between tapes becomes non-uniform (or just some of them quench), the “leaking” transverse field emerges, similar to the single tape scenario.
Single tape case:

## Slide 5: Hall sensor array design
Commercial GaAs ion-implanted Hall sensors: work well down to cryogenic temperatures (sensitivity actually increases with cooling)
6V dc / ~ 300 mA at 77 K (60 mW per probe)
A simple 2-layred PCB board designed to integrate sensors in an array
Yokogawa WE700 DAQ with 32 simultaneous 16-bit channels with hardware-selectable range down to 10 mV and speed up to 100 kHz used for data acquisition
Sensitivity is ~1.2 mV/mT

## Slide 6: Experimental arrangement
Piezo-transducer (small)
Piezo-transducer
(large)
Hall sensor array
Spot heater
Wire-wound heater
Linear period is 1.9 mm
16-tape CORC® wire; 20 cm long section with 15 cm long terminals

## Slide 7: I-V characteristic and raw Hall signals
120 mV criterion
IV curve
Straight line subtraction from each signal
Remove offsets
Signals are acquired at 1 kHz; then low-pass filtered at 100 Hz.
Post-processing:
As Hall array is not perfectly aligned with respect to the wire, signals corresponding to Bx are still present.
Linear current ramp to 1150 A.
(Full scale = 32768 units = 50 mV)
Liquid nitrogen bath conditions

## Slide 8: Current re-distribution towards the transition
Ramp at 5 A/s to 1115 A
Ramp at 4 A/s to 1159 A
#9
(120 mV criterion)
#10
Field “leakage”
0
30
Sensor #
Time (s)
50
312
(80 mV criterion)
0
30
Sensor #
Time (s)
50
417.6
Voltage
Current
~450 A
Inhomogeneous
current distribution at quench

## Slide 9: Heater-initiated quench (I)
#13
Heater location
~3.6 s
0
30
Sensor #
Time (s)
102
113.6
Heater
voltage
Voltage
850 A
6.9 cm/s
30
Sensor #
0
Wire heater (~ 4 Ohm, bifilar chromel wire)
is at sensor #4 location

## Slide 10: Heater-initiated quench (II)
2 cm/s
617 A
#9
585 A
#12
1.6 cm/s
~4 s
~3.4 s
Heater location
Heater location
0
Sensor #
Heater
voltage
0
Sensor #
Heater
voltage
Voltage
Time (s)
8
16.5
Time (s)
56
62.4
30
30

## Slide 11: Heater-initiated quench (III)
0.66 cm/s
400 A
#14
#15
~ 0 cm/s
~6.1 s
~10 s
300 A
Non-propagating quench
0
30
Sensor #
Heater
voltage
Voltage
0
30
Sensor #
Heater
voltage
Voltage
Time (s)
55
66.5
Time (s)
90
105.1

## Slide 12: Quench propagation velocity summary
QPV appears to be somewhat higher than in bare HTS tape at similar conditions. Potentially, current sharing and reduced cooling capacity within the CORC® wire play a role here; we will investigate it further.

## Slide 13: Active acoustic quench detection
A body is pulsed by a sender transducer and “ring-down” transient waveform reverberates across the sample multiple times
Transient oscillation is acquired by a receiver transducer; and stored as “reference” URef (t). Its shape is uniquely defined by the body geometry, density and elastic modulus E(T)
Transient acquisitions are repeated periodically; every new transient Ux(t) is compared to URef(t) using cross-correlation; the relative time shift Dt yielding maximal cross-correlation is measured.
When a hot spot develops, E(T) decreases locally, This proportionally increases Dt.
Sound velocity: , where Young’s modulus E exhibits the strongest temperature dependence: (s, t – adjustable parameters).
In metals the relative change is roughly at 77 K
DT~1.6 K
Dtrel~3.5 ns
DUmax ~ 7.6 mV (68 mV/cm)
ReBCO tape experiment

## Slide 14: Acoustic detection results
It turned out, that strong fluctuations of time shift are observed when CORC® wire is immersed in liquid nitrogen. This peculiar behavior is likely due to the intrinsic “multipath” configuration for the acoustic wave in CORC. Any minor change of a mechanical interface between neighboring tapes leads to a significant variation of the transient and corresponding fluctuation of measured time shift Dt.
Nevertheless, if standard deviation of Dt is plotted versus the applied current, a dependence closely resembling the I-V curve is recovered.
Reference waveform

## Slide 15: Conclusions
We have developed non-invasive magnetic and acoustic instrumentation to probe current sharing and quenching in CORC® conductors
Significant current flow inhomogeneity was detected in CORC® wire at current levels above ~ 0.3 I/Ic
Quench propagation velocity was measured as function of applied current at 77K; presence of non-propagating hot spots was confirmed for quenches initiated by an external heat pulse at I< 0.2 Ic
Active acoustic technique was proven to sense low-level heat dissipation in CORC® wires, and will be further developed in collaboration with ACT

## Slide 16: Future plans
We plan to develop flexible-board Hall arrays to perform de-convolution of current distribution in the CORC
We are thinking of ways of implementing Hall probes in CORC terminals for quench detection purposes
Hall arrays can be useful as a standard addition to LTS magnet instrumentation for measuring various aspects of remnant magnetization, creep, etc…