## Slide 1: LEM Production and Tests in Cold Box for ProtoDUNE-DP
E. Mazzucato (CEA Saclay/Irfu)
LEM Production
LEM HV Tests in Cold Box

## Slide 2: LEM Production

## Slide 3
LEM production
Cleaning + drying
+ polymerization
Soldering HV pins +
glueing MACOR insulation
Ultrasonic bath
LEM Survey
LEM manufacturing by ELTOS (Italy)
Characterization and tests @ CEA/Irfu
74 LEMs produced and tested for
CRP1 & CRP2 (Jan. – Oct./2018)

## Slide 4
LEM HV Tests in a High Pressure Chamber
Amplification gain inside the LEM depends on gas density  or P/T :
Geff = Ced with  = Ae-B/E
GAr @ 88°K and 1 bar  T = 20°C and P 3.3bar.
Stack of 6 LEMs
LEM + ANODE
CATHODE
collimated
 source

## Slide 5: LEM Design for ProtoDUNE-DP
CFR-34 – 311 prototype
CFR-35 – NP02
5 + 10 mm
clearance
2 + 2 mm
clearance
Not necessarily the final design for a 10kt detector!
 86% active area
 96% active area

## Slide 6: LEM HV Tests in Ar @ 3.3 bar
VLEM = 3.5 kV in Ar @ 3.3 bar
- gain > 100 in DLAr (before charging up)
- no PS trip for > 40h
- spark rate : 3/h
CFR-35 – NP02
CFR-34 – 311
- spark rate : 20/h
(> 45% of sparks near edges or corners)
multiple
sparks
VLEM = 3.3-3.5 kV in Ar @ 3.3 bar

## Slide 7: LEM QA/QC by ELTOS
LEM Specifications
RIM
HOLE
FR4
PCB

## Slide 8: LEM Survey @CEA/Irfu
t1
t2
t1
t2
Cu
Cu
FR4

## Slide 9: LEM HV training in Ar @ 3.3 bar
All LEMs in CRP1 and CRP2 tested up to 3.4 – 3.5 kV with < 1 spark / 20 minutes
LEM A143
LEM A143
Fully automated HV training up to 3.5kV
3 hours @ 3.5 kV
V
I

## Slide 10: LEM HV Tests in Cold Box

## Slide 11: CRP1 and CRP2 HV LEM Tests
Since last July, CRP test conditions with Cold Box have been improving (LV, TP, LAr level regulation, GRID, etc…)
LEM HV system modified (DB, HV connections) for
more reliable operation in cold GAr
HV PS system and SC monitoring still temporary
with limited number of channels (LEMs grouped
by 6; 14 available LEM HV channels instead of 42)
CAEN HV PS
Patch Panel
SCM
LEM DB
and HV
6 LEM TOP/DB
36 LEM BOT

## Slide 12: CRP1 and CRP2 LEM HV Tests
Will mainly address the CRP1 & CRP2 CB tests of Oct. and Nov. as CRPs were in their final configuration
No control on Ar purity (< 100 ppm) and GAr density (Patm)
Reminder : performance goals for ProtoDune-DP (G  20, stability) are based on the operation of a 3L TPC with an active LEM area 900 times smaller than one single CRP …
CRP1 and CRP2 operated for « long » periods of time at different HV settings with all 36 LEMs powered, exceeding by far LEM operation time with the 311
Still, some issues remain to be understood …

## Slide 13: Atmospheric Pressure
From ProtoDUNE-SP
SC web page
935 mbar
965 mbar
985 mbar
CRP2
CRP1
CRP1
CRP3
CRP1

## Slide 14: Dual Phase Principle
GND
-1kV
-4kV
-7kV
ANODE
LEM TOP
LEM BOT
Extr. Grid
Induction : 5kV/cm
Amplification : 30kV/cm
Extr. (GAr) : 3.6kV/cm
Extr. (LAr) : 2.4kV/cm

## Slide 15: CRP1 LEM HV Test Results (Oct. 15th - 30th)
Liquid level stable to within 250µm (2mm rapid increase of level occured during LAr delivery on Oct. 30th)
TLEM  91°K
VLEM-GRID = 3kV
As VTOP increases up to the 1kV nominal value, the field inside the LEM needs to be decreased in order to keep stable operation conditions.
* TRIP time set too short
VTOP (kV) | VBOT (kV) | ELEM (kV/cm) | Time (h) | Spark Rate (h-1) | Patm (mbar) | Geff (no ch. up)
0.25 | 3.35 | 31.0 | 12 | 1.3 | 968 - 972 | 20
0.50 | 3.55-3.60 | 30.5-31.0 | 13 | 1.3 | 962 - 966 | 24 - 31
0.75 | 3.70 | 29.5 | 42 | 0.6 | 943 - 953 | 20
1.00 | 3.80 | 28.0 | 18 | 2 trips* | 970 - 976 | 9
1.00 | 3.85 | 28.5 | 12 | 3 trips | 936 - 947 | 15

## Slide 16: Effective gain
VGRID-LEM (V)
Includes probability
for electrons to enter
LEM holes
VLEM_TOP (V)
Geff = extr × GLEM × ind
VLEM (V)
Gushchin et al., Sov. Phys.
JETP 55 (1982) 860-862.
Extraction efficiency
Collection efficiency

## Slide 17: First CRP2 Results (Nov. 15th – 26th)
LEM HV tests in DLAr mode started on Nov. 15th after CRP positioned w.r.t. liquid level
Liquid level stability : 200µm
VLEM-GRID = 3kV
VTOP (kV) | VBOT (kV) | ELEM (kV/cm) | Time (h) | Spark Rate (h-1) | Patm (mbar) | Geff (no ch. up)
0.10 | 3.15 - 3.20 | 30.5 - 31.0 | 17 | 0.8 | 969 - 973 | 9 - 11
0.25 | 3.34 | 30.9 | 16 | 1.3 | 968 - 970 | 19
0.50 | 3.55 | 30.5 | 11 | 0.9 | 957 - 965 | 24
0.50 | 3.555 | 30.55 | 42 | 0.5 | 962 - 964 | 25
Similar trend and HV limitation for CRP2 as CRP1. Being investigated. Not observed during source tests up to 35kV/cm in HP chamber with single LEM + anode.

## Slide 18: CRP2 : VTOP = 0.25kV and VBOT = 3.34kV
0 trip, 20 sparks @ 0.25kV/3.34kV in 16 hrs

## Slide 19: CRP1 : VTOP = 0.50kV and VBOT = 3.60kV
0 trip, 17sparks @ 0.50kV/3.60kV in 13 hrs

## Slide 20: CRP1 : VTOP = 0.75kV and VBOT = 3.70kV
4 trips @ 0.75kV/3.75kV in 24 hrs
0 trip, 26 sparks @ 0.75kV/3.70kV in 42 hrs

## Slide 21: CRP1 : VTOP = 1.00kV and VBOT = 3.85kV
3 trips @ 1.0kV/3.85kV
Very unstable conditions.
Coincides with LAr delivery
that caused 2mm rapid
increase of level in CB
Discharges

## Slide 22: Issues with CRP1
On Oct. 30th, found 3 problematic
LEMs (16,28,30) from 2 different HV
channels with similar discharging
behaviour
Needed to lower their VBOT value to
run in stable conditions
VTOP = 0.50kV and VBOT = 3.50kV

## Slide 23: Sparks, trips and continuous discharges
1 spark
Recovery time < PS TRIP time
2 sparks
Recovery time > PS TRIP time (120s)
No Trip
2 Channels
tripped
Recovery time > TRIP time
Continuously discharging
PS Trip
Can be detected
and stopped
with appropriate
SC monitoring
program
8 minutes

## Slide 24: Inspection of CRP1 after removal from the CB revealed dark spots on 4 LEMs
LEM 16
Points to a weaker regions of the LEM
Detailed E-field calculations needed
Issues with CRP1
Treated with potassium
permanganate @ CERN to
remove carbonized FR4 inside
holes. Copper surface
unaffected.
Cleaned at CEA/Irfu and tested up to 3.2kV @ 3.3 bar
Back on CRP1
LEM 30
LEM 28
LEM 14

## Slide 25: CRP2 : VTOP = 0.50kV and VBOT = 3.55kV
LEM 5 continuously discharging
for several minutes before HV
lowered
LEM 5
LEM 5 of CRP2 to be replaced
with a spare one

## Slide 26: Summary
LEMs manufactured by ELTOS for ProtoDUNE-DP meet the required specifications. However, the production for 2 CRPs took longer than expected (×2) due to manufacturing delays.
The two CRPs are now operational although some problems developped on a few LEMs (cured now) and some HV limitation observed. Nominal effective gain of 20 before charging up is within range.
Need « intelligent » SC program to better protect LEM against multiple discharges. Question being addressed with CERN-EP-DT-DI for ProtoDUNE-DP.
Ready now to use CRPs in ProtoDUNE-DP and operate LEMs in real conditions (LAr purity and stable thermodynamic conditions, drift region) to assess performance of a large scale DLAr TPC.
Understanding/improvement of LEM performance & design can be addressed concurrently.