## Slide 1: 3.9 GHz Cryomodule Testing
Elvin Harms
LCLS-II 3.9 GHz CM Delta Final Design Review
29-30 January 2019

## Slide 2: Outline
Status of and plans for infrastructure modifications
Cryomodule handling at CMTS
Status of test plan and traveler
Lessons learned from 1.3 GHz
Respond to charge questions:
  Q5: Is the design likely to meet performance expectations?
  Q8: Have lessons learned from the 1.3 GHz cryomodule quality assurance, assembly and test been applied?

## Slide 3: Test stand infrastructure
Move end cap girder
  ~half the distance to the feed cap
Replace cabling in cave as needed
Cryo unchanged
Reviewing/comparing instrumentation differences with 1.3 GHz
6-8 week changeover
Testing at ~1/month rate

## Slide 4: Test stand infrastructure – RF
3.9 GHz Amplifiers
  All SSA’s in hand
  One at CMTF and getting prepped for testing
  Acceptance data received from SLAC
LLRF will be different
  Use the LCLS-II production system
  Commissioning time will be needed
  Schedule, staff availability of some concern

## Slide 5: Test stand infrastructure – RF
All RF distribution hardware received
  Characterization in progress
  Waveguide through wall pre-installed
Courtesy Ding Sun
Part No | Insertion Loss (from Input to Output Port) db | Return Loss (Input Port) db | Forward Coupling db | Reflected Coupling db | Return Loss (Output Port) db | Return Loss (Forward Coupling Port) db | Return Loss (Reflected Coupling Port) db | Directivity at Forward Coupling Port db | Directivity at Reflected Coupling Port db
1845196 | 0.007 | 41.720 | 40.196 | 40.498 | 41.065 | 22.789 | 22.003 | 34.117 | 38.064
1845194 | 0.007 | 38.461 | 40.439 | 39.944 | 39.671 | 22.788 | 20.827 | 43.875 | 48.447
Mega WR284 Dual Directional Coupler
Jan, 2019 --- Feb, 2019
Data @ 3.9 GHz

## Slide 6: Cryomodule handling at CMTF
Vacuum and cryogenics connections are identical to 1.3 GHz cryomodules
  Use clean room practices where appropriate
1.3 GHz experience valuable
  Proven quality techniques and staff
  All aspects
    Arrival
    Installation
    Prep for test
    Cooldown
    Cold testing
    Warmup & removal

## Slide 7: Test stand operation – ESH&Q
ESH&Q engaged
  ORC for 3.9 GHz is being planned for
  Existing shielding and detectors appear adequate
    Do due diligence and confirm
  TLM detector already modified
    Runs the length of the cryomodule
    Split into 2 parts to ensure easy switchover for shorter length.

## Slide 8: Status of test plan & traveler
Acceptance criteria drafted
  Under initial review
Test plan under development
  Based on 1.3 GHz plan with appropriate changes
Traveler draft prepared and under review
  Based on 1.3 GHz Test Traveler, # 464421

## Slide 9: Acceptance criteria (draft)
LCLSII-4.1-PR-0097-R2: 3.9 GHz Physics Requirements (4/7/2016)
LCLSII-4.1-FR-0096-R1: 3.9 GHz Functional Requirements Specification (12/8/2016)

## Slide 10: Test stand operation
1.3 GHz cryomodule testing has already demonstrated mature and proven applications
Adaptable for 3.9 GHz testing
Techniques and test sequencing evolving for optimum throughput

## Slide 11: Testing differences/similarities with 1.3 GHz
Simplifications
QL set, fixed couplers
No magnet
Flux expulsion not needed
Fast cooldown not necessary
Cro vacuum connections identical
Radiation monitoring/faraday cups unchanged
Potential Challenges
LLRF
Blade vs. end lever tuners
Microphonics

## Slide 12: Lessons learned
CMTS1 is a proven, reliable test bed
  Necessary modifications understood
  Majority of the changes are mechanical
  8-cavity cryomodule makes change-over simpler
  Staff continuity
Learn from past experience
  Test stand already commissioned
  Alignment
  Cryo heat exchanger failure
  Microphonics mitigation
  Post-test leak check
  3.9 GHz for FLASH (and XFEL) experience
    don’t assume ‘success is in the bag’

## Slide 13: Response to charge
  Q5: Is the design likely to meet performance expectations?
    Early results from single cavity testing
    CMTS1 performance to date
    Strong, experienced team
  Q8: Have lessons learned from the 1.3 GHz cryomodule quality assurance, assembly and test been applied?
    Following on to experience of testing 19 modules over 3 years already

## Slide 14: Summary
Preparation for testing 3.9 GHz cryomodules begun
  Regular (at least bi-weekly) dialogue
  Test documents under development largely based on existing PRD and FRS
  ESH&Q already engaged
RF hardware in hand and being assessed
1.3 GHz testing a valuable experience
  Excellent results achievable on schedule
  Production mindset being embraced

## Slide 15
Thank you for your attention – Questions?

## Slide 17: Back up slides