## Slide 1: Quench Antenna Status and Plan in AUP MQXFAP 1b Magnet
H. Song, Sonny, J. Muratore, P. Joshi
For AUP quench antenna coordination efforts
6/10/2019

## Slide 2: Quench Antenna Orientation - 1
Orientation check
  Antenna coils are facing magnet midplanes (not magnet coil centers!) @ Maxim
  there was slight misorientation when magnet/quench antenna center vertically realigned by 100 mm in quench 23, but has been corrected in quench 24.
    Coil in 1b magnet is ~200 mm shorter than AP2

## Slide 3
4753.61 mm
[187.142 in]
Top of Magnet Shell
Bottom of Warm Bore Tube
Bottom of Magnet Shell
4563 mm
[179.646 in]
Inside of bottom of warm bore tube below end of magnet support rods = 20.4 mm
Top of Warm Bore Tube
6256.57 mm
[246.322 in]
903.86
[35.585]
Top hat to Lambda plate
Rotating coil in MAX UP Post
656
25.827
Rotating coil
1b Magnet center elevated/realigned by 100 mm in quench #23, because 1b coil is 200 mm shorter than AP2 coil.
Quench
Antenna
Elements
Location
-2
8 Elements in middle

## Slide 4: Quench Antenna Gain - 3
Quench antenna gain check
R = 49.9 ohm
Actual Gain = ~991.
Is this correct? As designed? @Maxim
There were paper document which indicate Gain of 50 or 650, not 991. not consistent.
How to check the interface box – to ensure it works functionally? @Maxim
And the power supply? @Maxim

## Slide 5: DAQ cards PXIe 4300
Confirm the channel numbering in DAQ and Antenna (BNL)
  Existing 4 slots of cards, 32 channels for 16 elements in the chassis (Quench Antenna)
  The acquisition voltage range was (+/- 10 V), changed to +/1 V - today as of 6/11/2019.
    Available range, +/- 10 V, 5 V, 2 V, 1 V, (no +/- 0.1 V)
  4 more slots of cards can be added
    Need formal approval? Just go ahead order? At $9036
    @AUP Team
http://www.ni.com/pdf/manuals/375476c.pdf
http://www.ni.com/en-us/support/model.pxie-4300.html
(lead time 5-10 days)

## Slide 6: Plan to add 9 more elements to the existing quench antenna
4 more PXIe 4300 cards will be ordered – Go ahead submit the order? 5 -10 days @AUP Team
When the 9 more elements will be delivered to BNL? (from LBNL?) @Maxim and Thomas
  Will the interface box included? @Maxim and Thomas
Additional labor/time
  More cabling (underground) and wiring to accommodate the 9 more elements – 10-15 days labor based Sonny’s feedback
Turn-around time for 2nd thermal cycle – 8-9 days – based on B. McKeon
  Shall we wait for the quench antenna modification to be completed? @AUP Team

## Slide 7: Conclusions
Orientation checked and corrected
Drawing and quench antenna elements –compared.
Gain checked = 991
PXIe 4300 voltage range changed to +/- 1 V - today
Plan to add 9 more elements to the existing quench antenna
  A few decisions and/or approvals are needed.

## Slide 8: Backup slides

## Slide 9: Meeting summary – quench antenna discussions meeting (6/7/2019) – T. Strauss
Top priority:
  Overlay drawing of the QA with mechanical structure is top priority to identify quench location on structure.
To Do
  Confirm the channel numbering in DAQ and Antenna (BNL)
  QA disk distance measurement with tape measure (BNL)
  Generate an assembly drawing (BNL)
  Check the gain settings (BNL)
  Determine the mechanical center of the magnet w.r.t. QA location
  Check interface box has capacitors or not, have them identical.
  Check orientation, idea is the coil looks at the inner layer
  Quench propagation in low field region estimate (Giorgio) Inner layer: 8mm/ms - done
  This will determine the QA distance needed

## Slide 10
Short term plan,
  Next test in next Thermal cycle with better coverage of QA. 8mm/ms with ~10ms: 16cm spacing or 25 elements needed, e.g. 50 channels.
Maxim
  Send existing PCB boards and parts to BNL, and schematics for assembly.
  Send interface board drawing to BNL
BNL
  Assemble PCB boards
  Assemble interface box
  Assemble a cable for 8 channels
  32 more channels available, cards would need to be bought, 4 more cards ($2.2k/card).
  Measure the extra resistors on the boards to ensure is everything is the same
  Change DAQ range from 1V to 100mV
Plastic spacer manufacturing and procurement - get the right length (maybe Maxim has spares).
Before next test
  Confirm QA gain settings, and test of full assembly with long cables before insertion.
  Fermi
  Investigate how more elements can be added. Check if alternative to PCB antenna exists, can be made or re-used (suggestion Stoyan/Joe DiMarco design).

## Slide 11
Medium Term goal
  Update the schematic drawings (Maxim)
  Goal a 64 channel QA - 32 elements
  64 channel interface box
Long term plans
  Plan for Test Stand 4 QA
  QA on trace on the coil, possibilities at BNL and FNAL

## Slide 12: Coil drawing
Jesse Schmalzle, Michael Anerella,

## Slide 13: Antenna configuration and positioning
QA1
QA2
By Maxim, Joe, Sonny,

## Slide 15
4753.61 mm
[187.142 in]
Top of Magnet Shell
Bottom of Warm Bore Tube
Bottom of Magnet Shell
4563 mm
[179.646 in]
Inside of bottom of warm bore tube below end of magnet support rods = 20.4 mm
Top of Warm Bore Tube
6256.57 mm
[246.322 in]
903.86
[35.585]
Top hat to Lambda plate
Rotating coil in MAX UP Post
656
25.827
Rotating coil
~220 mm
200.8 mm
Home
355.6 mm
[14’’]
MagCent  Coil220_Cent
2281.5 – 200.8 – 164.99
=1915.71 (mm) Half_L
Half_U = 2814.29 mm
4730mm
190.61mm
164.99mm
331.99mm
Warm Length
Switch_UL