## Slide 1: Development of J-PARC New Muon Monitor
Takashi Honjo
Osaka City University
on behalf of the T2K muon monitor group
October 25, 2019
Notes: Hello, Iʼm Takashi Honjo. Iʼm going to talk about “Development of J-Parc New Muon monitor”.

## Slide 2: Contents
Current muon monitor
Issue of current monitor for higher beam power
Candidates for the new monitor (PMT & EMT)
EMT beam test
Summary
Notes: 2. These are contents of my talk. First, I explain the current muon monitor. Next, I mention some issues of the current monitor for higher beam power. Then, I talk about the candidate for the new monitor, PMT and EMT, preparation status of EMT beam test, and summary.

## Slide 3: What is Muon Monitor (MUMON)
Monitor the direction and intensity of muons from pion/kaon decays.
This is only detector that can monitor the beam bunch-by-bunch after the hadron production target.
→It’s essential for T2K.
Composed of Si & ion chamber (IC) array.
μ
Ion
Chamber
Si
Array
Notes: 3. Let me start from the current muon monitor, MUMON. This detector monitors the direction and intensity of muons produced along with the neutrino.
It is located 118 m away from the target. This is the only detector that can monitor the beam bunch-by-bunch after the hadron production target. So, MUMON is essential for the beam operation. MUMON is composed of two arrays of sensors, one is the silicon PIN photodiodes and another is the ionization(アイニゼイション) chambers.

## Slide 4: Current Monitor (Si, IC)
Working since 2009.
Each Si & IC has 7 × 7 channels and covers 150 ×150 cm2.
Get the beam intensity and position by fitting 2D distribution of collected charges.
The intrinsic detector resolution (RMS).
  Beam center<0.3 mm
  Beam intensity<0.1%
yield
total 45 months
[mrad]
0
0.5
-0.5
0
0.5
-0.5
-1
1
x direction
y direction
[cm]
Notes: 4. This detector is working since 2009.
Each array covers a 150(one fifty)×150 cm2 region.
And there are 7×7 channels with a 25 cm pitch. And, get the beam intensity and center position by fitting 2D distributions of collected charges. Resolution is less than 0.3 mmfor the beam center, and less than 0.1% for the intensity measurement. Si & IC has shown excellent performance since the T2K started.

## Slide 5: J-PARC Upgrade
In the future, J-PARC will increase the beam power.
# of muons / cm2 / beam bunch
Operation cycle
Now (485 kW)
Future (1.3 MW)
2.48 s
1.16 s
3.5 × 106
4.9 × 106
2016
2018
2020
2022
2024
2026
2028
2030
MR Beam Power [kW]
0
200
400
600
800
1000
1200
1400
0
5
10
15
20
25
30
35
40
Integrated Delivered Protons [1021POT]
MR Power Supply upgrade
Candidates for new detector
PMT
EMT ->Replace photocathode of PMT with Al (detail later) ->Higher radiation hardness
-> Concerning some issue
as the beam power is increasing
Horn current
250 kA
320 kA
Notes: 5. We are planning to increase the T2K beam power in the future.
The bunch cycle will be doubled.
the T2K horn current will be increased to 320 kA from 250 kA.
The muon flux per bunch will increase by a factor of 1.4.
We are concerning some issues as the beam power is increasing.
So, we are considering new detectors.
There are the two candidates PMT and EMT.
I will explain them in detail later.

## Slide 6: Issue of Current Detector (Si)
Response degradation of the Si Detector.
Signal decrease due to radiation damage is seen (~1% / 5×1020P.O.T).
The requested value from T2K is within 3%
-> Replace the Si detector every half a year in the future
1%
Notes: 6. We have concern for using the current detector in the future. The first issue is response degradation of the silicon detectors. This figure shows the result of the last physics run at T2K. The signal from the silicon decreased by 1% during about 5 months for the current beam power. The requested value from T2K is within 3%. So, we would have to replace the Si detector every half a year for the higher beam intensity expected in the near future.

## Slide 7: Issue of Current Detector (Ion Chamber)
Signal response of IC
for higher beam power
IC-Ar yield/proton beam power
vs. proton beam power
IC-Ar
IC-He
400 kW
Non-linearities, at high intensity.
Using of He gas instead of Ar.
-> Pileup is due to faster ion drift speed.
ADC count
ADC count
FADC sample (~15 ns/sample)
FADC sample (~15 ns/sample)
[1012 protons/bunch]
IC Center/Reference
[nC/(1012 protons/bunch)]
5 μs
5 μs
~8%
Notes: 7. The second issue is the signal response of the ion chambers for higher beam powers. The left figure shows the normalized response of the IC using the Ar gas for eight bunches at various beam powers.
We see non-linearities here from 400 kW. One solution of this would be to use a lighter gas, such as He,
in which the number of generated electron/ion pairs is much smaller than Ar. However, as shown in the right figure,
He suffers from signal pileup caused by the fast drift speed of the lighter ion. So, this affects bunch-by-bunch monitoring.

## Slide 8: What is the Electron–Multiplier Tube (EMT)?
The components of EMT are the same as the standard PMT except the photocathode.
Photocathode material is replaced with Al for EMT -> Expect higher radiation hardness (low signal gains though).
We have some options for EMT.
PMT
EMT
The Development of Prototype Detectors -PMT&EMT-
Notes: 8. I’m going to talk about the development of prototype detectors
Let me explain the EMT.
This is candidate of new detector.
These figures are the schematic views of PMT & EMT. Components of the EMT are the same as the standard PMT except the photocathode.
Photocathode material is replaced with Al(アル-ミナル) which we expect has higher radiation hardness.
The electron emission rate of Al is smaller than standard photocathode but sufficient signal is expected for the high intensity beam. In addition to this, we prepare some EMTs.
These are different options.

## Slide 9
The Development of Prototype Detectors -PMT&EMT-
16 mm
PMT
EMT
14.6 mm
Notes: This is the picture of PMT and EMT. the diameter of effective area is 8 mm.

## Slide 10: The Development of Prototype Detectors -PMT&EMT-
Some candidate detectors were tested.
Installed behind the MUMON.
μ
Notes: 10. The performance of some prototype detectors was studied with the T2K muon beam. There are 2 PMTs and 6 EMTs like this picture. These were installed behind the MUMON.

## Slide 11: PMT’s Stability
The signal continued to drop for about 3 months.
Tried different voltage.
Stability is worse than the current detectors.
PMT #1
PMT #2
We tested two PMTs.
PMT #1/CT
about 3 months
type B/CT
about 3 months
Change HV
Change HV
15
20
25
30
35
40
45
50
55
60
15
20
25
30
35
40
45
50
55
60
10
Notes: 11. First, let me talk about the PMT. We installed two standard PMTs. We checked the linearity and stability.
The PMT signal continued to drift for three months.
So, the stability is worse than the current detector. The cause is thought to be damage of the photocathode.

## Slide 12: PMT’s Signal for 8 bunches
Measured the signal of PMT for different Intensities and HV.
ADC count
470 kW, -350V
40 kW, -500V
FADC samples (~15 ns/sample)
FADC samples (~15 ns/sample)
Expect the same magnitude
of response
220
ADC count
50
PMT does not seem to be a good candidate.
Signal was OK at 40 kW, but saturated at 470kW.
->Voltage drop by photocathode resistance.
gain at -500 V is ten times at -350 V
Notes: 12. We measured the signal of PMT for different Intensities and HV.
Under those two conditions, the signal is considered(d) to have the same magnitude.
However, the waveform is strange at high intensity.
PMT shows clear non-linearity due to voltage drop by photocathode resistance.
So, PMT does not seem to be a good candidate.

## Slide 13: Time Response of EMT
Smaller tail contribution is better. -> Affects bunch-by-bunch monitoring for T2K.
EMT #1
Ratio (Tail/1st bunch)
Si
ADC count
FADC sample (~15ns / sample)
ADC count
Tail
EMT performs better than Si and IC
EMT #1
EMT #2
Notes: 13. This is a result of time response of EMT. Smaller tail contribution is better because it affects (a) bunch-by-bunch monitoring.
The upper left figure shows EMT
The lower left figure shows Si.
The right figure shows the tail contribution obtained from the regions shown here.
EMTs perform better than Si and IC for bunch-by-bunch beam monitoring.

## Slide 14: Intensity Resolution of EMT
Intensity resolution of the EMT is better than a required value of ~1% . for beam position measurement with an accuracy < 0.3 mrad.
EMT #1
Counts
EMT #1
EMT #2
EMT/Reference[pC/(1012 protons/spill)]
Notes: 14. Next, Intensity resolution about EMT
Intensity resolution of the EMT is worse than that of Si nor IC,
but better than a required value of about 1% for beam position measurement with an accuracy <(less than) 0.3 mrad

## Slide 15: Linearity of EMT
Measured with two EMTs with different capacitances in divider circuit.
Large capacitance may give better response at higher intensity.
10
10
10
15
100
100
330
330
100
330
(nF)
R1 – R10 : 330 kΩ
R11 : 160 kΩ
small capacitance
large capacitance
Notes: 15. We measured linearity for two EMTs with different capacitances in divider circuit.
This is the divider circuit.
Large capacitance may give better response at higher intensity.

## Slide 16: Linearity of EMT
Large capacitance EMT shows better linearity as expected.
Type 1A
Type 1A
6.5 %
Type 1B
Type 1B
3.5 %
small capacitance
large capacitance
- 500 V
- 500 V
500 kW
500 kW
[1012 protons/spill]
[1012 protons/spill]
EMT #1/Reference
[pC/(1012 protons/spill)]
EMT #2/Reference
[pC/(1012 protons/spill)]
Notes: 16. These are the results.
The left figure for small capacitance, and right one for the large capacitance.
At the high intensity, the signal decreased by about 5 %.
Large capacitance EMT shows better linearity as expected.

## Slide 17: Linearity of EMT
EMT with the same capacitance was compared for different HV. -> Low HV is better (less electric field distortion).
- 500 V
small capacitance
small capacitance
- 450 V
6.5 %
3.5 %
[102 protons/spill]
[102 protons/spill]
EMT #1/Reference
[pC/(1012 protons/spill)]
EMT #1/Reference
[pC/(1012 protons/spill)]
Notes: 17. EMT with the same capacitance was compared for different HV.
The linearity is better for small voltage.
Higher voltages degrade the linearity due to electric field distortion.

## Slide 18: Stability of EMT
We see ~3% drift for the first 10 days since installation. -> After the drift, both the EMTs show good stability (~1%).
One of the four EMTs showed a large and long drift. -> Tube-to-tube variation large?
EMT yield
2/23
3/30
4/1
4/11
3 %
15 %
3 months
drift
drift
1 month
11 days
Notes: 18. This is the result of stability.
We measured the stability for four EMTs.
The results for one EMT are shown in the figure above.
We see about 3% drift for the first 10 days since installation.
After the drift, the EMTs show good stability.
However, One of the four EMTs showed(d) a large and long drift
So, Tube-to-tube variation could be large?

## Slide 19: Summary so far and future research
PMT is worse than current detectors.
EMT is a good candidate for new muon monitor.
However…
  As intensity increases, linearity deteriorates. -> Change the voltage ratio of the divider circuit.
  One of the four EMTs showed a large and long drift. -> We are going to measure many EMTs.
  Large drift at the beginning due to alkali metal on dynode. -> Test EMTs without alkali.
Notes: 19. This slide is summary so far and future research
PMT is worse than current detectors.
On the other hand, EMT is a good candidate for the new muon monitor
However, There are something to study more.
First, linearity gets worse at high intensity.
We will Optimize the divider circuit
One of the four EMTs showed a large and long drift.
So, we are going to measure many EMTs.
The signal drifts(s) are rather large due to alkali metal.
We will test EMT removing alkali metal from dynodes.

## Slide 20: EMT beam test
We study more details EMT by a beam test.
We are planning to measure the followings …
Linearity up to 1.4 times current intensity (4.9×106 muons/cm2/100 ns)
Stability corresponding to 10-years beam exposure with the planned intensity.
3. Others
Without alkali metal
EMT with different circuits
at the end of Nov. 2019
at Research Center for Electron Photon Science (ELPH)
at Tohoku University in Japan
Notes: 20. The EMT could be a good candidate for the new muon monitor. We study more details by a beam test at Research Center for Electron Photon Science (ELPH) at Tohoku University in Japan.
In particular, performance tests at future intensity are required. We will measure linearity up to the future intensity,
and stability corresponding to 10 years beam exposure with the planned intensity.
In addition, using EMT without alkali metal,
And, different divider circuits.

## Slide 21: Setup of beam test
Measure the beam profile with Si array.
Measure the beam intensity with CT.
Compare EMT with the current detector of Si.
beam
window
Si array
Si
EMT
CT
Setup1
Setup2
beam
window
e-
e-
Notes: 21. This is the beam test setup.
Setup 1 is used(ユーズド) to measure the profile of electron beam by using Si array.
In the setup2, measuring EMTs performances compare with the current detector of Si.
CT is put in the upstream to calibrate the beam intensity.

## Slide 22: Preparation for EMT beam test
Supporting structure for sensors have been designed.
Scattering effect is checked by FLUKA simulation. -> less than 1%
beam window
Setup 1
Setup 2
change
Si
EMT
Si array
e-
160 mm
270 mm
200 mm
180 mm
e & e-
beam window
supported Si
EMT
Notes: 22. Supporting structure for sensors have been designed.
These are the supporting structure for sensors for the beam test.
Scattering effect(イ) is checked by FLUKA simulation.
The support was designed to produce as little scattering as possible.
And confirmed that scattering by the support is less than 1%.

## Slide 23: Summary
MUMON is essential for the T2K experiment to ensure good neutrino beam quality.
We are developing a new MUMON detector for higher intensity beam operation.
EMT is a good candidate. -> However, we need to study more about linearity, large drift …
We are planning a beam test at the end of Nov. 2019.
Looking forward to EMT test results !!
Notes: 23.
This is summary.
MUMON is essential for the T2K experiment to ensure good neutrino beam quality.
We are developing a new MUMON detector for higher intensity beam operation.
EMT is a good candidate. However, there are some little problems.
We are planning to do a beam test at the end of Nov. 2019.

## Slide 24: backup

## Slide 25
Stability EMT Type 2B

## Slide 30: Prototype EMT
prototype detector that improved the circuit of Hamamatsu R9880
capacitance of decoupling capacitor change for corresponding 8 bunchs during 5 μs

## Slide 32: What is the EMT?
The component of EMT are the same as the standard PMT except the photocathode.
Photocathode material is replaced with Al for EMT →expect higher radiation hardness (low signal gains though).
Current option
Type of the divider circuit is difference of distribution ratio
expect Increase efficiency of secondary electron emission for include Sb
 | Dynode material (Sb) | Divider circuit | Expected gain
Type 1 | yes | nominal | 〜6×106
Type2 | yes | Type-A | 32% of Type 1
Type3 | no | nominal | 10〜103
Notes: radiation –resistant

## Slide 33: Setup of EMT Type-1
Two custom-made prototype detectors based on HAMAMATSU R9880.
Dynode decoupling capacitors between the types.
Readout : 65 MHz Flash-ADC (the same as the Si and IC)
PMT
Type A
PMT
Type B
EMT
Type 1A
EMT
Type 1B
μ
Notes: ダイノード後段での増幅電子の影響を見るため、2種類のHV値でテストした

## Slide 34: Setup of EMT Type-2
To improve linearity, we change the voltage ratio of the divider circuit. -> By reducing the potential difference between some dynodes in the first half, the space charge effect at the later stage was suppressed.
PMT
Type A
PMT
Type B
EMT
Type 2A
EMT
Type 2B
μ
Notes: ディバイダー回路の電圧比を調整した。線型性を改善するために
前半のダイノード間の電位差を小さくして後段で空間電荷効果を抑制させる。

## Slide 35: Result of EMT Type-2
Type-2A is stable within 1%, after drifting 3%.
Type-2B drifted 15% in a month, then became stable.
We did not measure linearity. -> The linearity is to be measured in the next beam test.
2.5 month
2.5 month
Type-2A Yield (with HPS correction)
Type-2B Yield (with HPS correction)
Notes: Type-2Aは始め数%減衰した後は、1%程度で安定
Type-2Bは1ヶ月かけて15%信号が減少したが、その後安定

## Slide 36: Setup of EMT Type-3
EMT Type-3 has no alkali-Sb on dynodes (expected gain is 10~103) →We expect higher radiation hardness than the Type1.
Unfortunately, we could not see any signal for 450V.
PMT
Type A
PMT
Type B
EMT
Type 3A
EMT
Type 3B
μ

## Slide 37: Measurement of EMT Type-3
We will check if we see signal in the beam test using higher beam intensities.
radioactive source
fluorescent light
with alkali-Sb
HV -800 V
HV -800 V
with alkali-Sb
without alkali-Sb
We tried other signal sources such as fluorescent light and a radioactive source(152Eu: gamma rays at ? eV)” .
Recommended voltage value is 500 V, but no signal even higher voltages.
HV -1300 V
without alkali-Sb
HV -800 V
Notes: 電圧の適正値は500 Vだがそれよりも高い電圧をかけても信号が見れなかった

## Slide 38
e
e
dynode 1
dynode 2
dynode 3
dynode n
anode
Al
w/o alkali-Sb
μ
e
e
dynode 1
dynode 2
dynode 3
dynode n
anode
Al
w/ alkali-Sb
μ
e
e
dynode 1
dynode 2
dynode 3
dynode n
anode
μ
Photocathode
e
e
dynode 1
dynode 2
dynode 3
dynode n
anode
Al
μ