## Slide 1: Muon Target Replacement
N. Kawamura, Muon Sci. Sec., MLF (Materials and Life Science Facility), J-PARC
NBI2019, 25/Oct/2019

## Slide 2: Muon production target system
Scraper (proton collimator)
Scrape the scattered protons by the target
Renewed in 2015.
Heat deposit：20 kW＠No1
Water cooling
Thermocouple thermometers
Rotating target
To prolong the lifetime of the target against DPA, operation started. #1: 2014-2019 #2: 2019-
Heat deposit：4 kW
Graphite disk: rad. cooling Bearing, jacket: water cooling
Monitors: motor torque, vacuum level, Q-mass for outgas. IR-Radiation thermometer is going to be operated.
Vacuum vessel
In service since 2008
Hea deposit：10 kW
Water (>90%) and air cooling
Thermocouple thermoeters
p
m
m

## Slide 3: Muon target assembly
3.5 m
3.5 m
2.4 m
plug shield
target rod
The top flange, 2.4 m from BL, is accessible for hands-on work. (Maintenance floor) The dose rate is not so low, but one can stay for several hours.
Maintenance floor in M2 tunnel

## Slide 4: Muon production target
To disperse the radiation damage, 1% shrinkage / 1DPA (~0.5 year irradiation under 1 MW), and to prolong its life time, rotating target has been applied since 2014.
Common property
  2-cm thick IG-430U from Toyo Tanso Inc., consuming 5% of protons
Different property
  direct water cooling (fixed) vs. radiation cooling (rotating)
  Max temp.: 1500 deg C (fixed) vs. 700 deg C (rotating)
Fixed target: ’08-’14
Num. of P.: 2500 MWhr.
Rotating target #1: ’14-’19
Num. of P.: 5800 MWhr.

## Slide 5: Target Replacement
We performed the target replacement work only twice.
In 2014: Fixed target  Rotating target
  The duration time of the fixed target reached at the estimated life time corresponding to 1 DPA, 0.5 years  1MW.
  We didn’t mind about tritium contamination seriously. We just placed a local IC monitor, but didn’t use any special equipment.
In 2019: Rotating target #1  Rotating target #2
  An valuable joint was found to be used in a rotating rod in 2018.
  In ’18, we decided to use a bad one until the next maintenance period of ’19 to construct a fixing/replacing scenario and to prepare special equipment like a local exhaust ventilator, an airline mask, an anorak suit etc.
  According to experiences in these years, the air in the working area was expected to be highly contaminated by tritium above the permissible level in Japanese low of 0.8 Bq/cc.

## Slide 6: Fixed target  Rotating target
Maintenance work in ’14

## Slide 7: Issues
In principle, the scenario of target replacement was constructed at the construction phase of MLF, and this work was a part of that.
  All the maintenance/replacement scenario was commissioned before the first beam in 2008.
Suppressing/controlling external exposure during transporting the highly activated target assy. from the beamline to the storage pit.
  By using of cask, the main work could be performed remotely.
  Hands-on works were minimized.
In ’14, the main concerns were
  the scenario could really be performed against highly activated devices
  the first rotating target could work properly.

## Slide 8: Conceptual design of cask in ’04
Target assy.
Cask
Target
Plug shield
Residual activity in the 1st segment of the target plug
x
Reality of the maintenance scenario is examined.
Dose on the cask surface
10-year irradiation and 1-day cooling

## Slide 9: Dose during target maintenance
In principle, target maintenance can be controlled remotely.
  Accumulated dose is negligible.
In trouble, one can work
  for several hours at FL+6m
  for a few minutes at FL+4m
A ring structure (a cover) is necessary in the service area.
service area
FL+4m
FL+6m
x
Ring structure
10cm gap

## Slide 10
Maintenance of muon target
Ceiling shields transportation
Temporary storage space
Removal of ceiling shield
Temporary storage in storage pod
Drying of cooling water in Drying room
Target exchange in Hot cell
Chamber
Storage
pit
Drying room
Hot cell
Off-beam commissioning

## Slide 11: Muon Cask
Up-down speed; 0.5m/min
Total weight; 23t
Handled apparatuses ；maximum weight; 6t
target, monitor, scraper, pillowseal
Shielding thickness; 150mm
Reproducibility of location; ±5mm
Remote controlled gripper
Off-beam commissioning

## Slide 12
Target
  Beam line, storage pit –completed
  Drying room, hot cell – this March
Monitor (M22, M23)
  Beam line –completed
  Storage pod, hot cell –this March
Pillowseal, scraper (#1, #2)
  Beam line -completed
Commissioning by Muon cask
Off-beam commissioning

## Slide 13
Target exchange in hot cell
Manipulators
Plug stand
Exchange device
Cutting device
Plug stand
Exchange device
Master slave manipulator
Cutting device
Power manipulator
Off-beam commissioning

## Slide 14: Replacement work in ’14
Vinyl sheets for a counter measure against radio isotopes on the target assy. Meaningless? “Better than do not”?
Preparation of the storage pit
The cask on the pit

## Slide 15: Rotating target #1  #2
Maintenance work in ’19

## Slide 16: Issues
In principle, the replacement work itself is almost the same as that in ’14 (fixed target  rotating target).
However, in these years, it was reported the tritium contamination in the target chamber and the vacuum components becomes serious; ~ a few kB/cm2.
Treatment of tritium wasn’t considered in the original target exchange scenario.
We refurbish the scenario, adding the use of green-house that is a small chamber covered with air-tight vinyl sheets, local exhaust ventilator, an airline mask and an anorak suit to control the internal exposure of the working staff enough low.

## Slide 17: Trouble on the shaft-joint in ’18
M1トンネル
joints
P8GVで封止
P7ポンプ位置
A Shaft-joint is used to absorb misalignment btw the shafts.
The lower side joint was located inside of the shield, and we don’t have enough time to prepare the replacement work, we decided to keep using it, after confirming the target rotation at least.
Due to cyclic stress and processing mistake, the upside shaft was broken after 5 years use.
A through-ditch by
processing mistake
Rotating machine
During the annual rotating-machine maintenance in ’18, we found the damaged shaft-joint, and then one of two was replaced with an appropriate one. But, the other one was left as it was.

## Slide 18: Target assy. removal procedure
遮蔽付きキャスク
Disconnect cables, pipes and misc.
Vacuum purge
Assemble the green-house
Remove up-down motion system Mount cask attachment
Mount cask rack and ext. shield
Mount cask
遮蔽
ミュオン標的
上下駆動機構
矩形フランジ
GH
M2 tunnel
Material for investigative commission

## Slide 19: Target assy. removal procedure
遮蔽付きキャスク
Disconnect cables, pipes and misc.
Vacuum purge
Assemble the green-house
Remove up-down motion system Mount cask attachment
Mount cask rack and ext. shield
Mount cask
Winch up the target assy.
Transport the cask on the pit.
Store the used target in the pit.
遮蔽
ミュオン標的
GH
M2 tunnel
Material for investigative commission

## Slide 20: Replacement work in ’19
GH
local exhaust
local exhaust
airline mask
anorak suit
air supply line
The work wearing airline mask and anorak suit is hot and extremely tough. 2-hour is the upper limit for one work.

## Slide 21: Tritium release to the beamline in ’18
Dry N2 only
Dry N2 + air(10%)
flange open
Manipulation like opening a flange must be done carefully because the air is drawn into the chamber rapidly.
Humidity is bad for vacuum. We need a longer time to achieve good vac. level again.
Tritium concentration in the purge gas ’18
Isotope exchange on surface: H2O  HTO
Even dry N2 contains a small fraction of water as an impurity. (~10 ppm)
H2O is useful for tritium decontamination.
Elapsed time [days]
Elapsed time [days]

## Slide 22: Tritium decontamination by air in ’19
purge
evac+purge
evac+purge
flange open
Target assy. removal
Vacuum sealing
We did tritium decontamination procedure by the humidity-controlled gas (dry N2 + air(10%)), quantitative explanation is difficult, but this was obviously efficient to reduce the tritium level.
Tritium concentration in the purge gas

## Slide 23: Summary
Originally, the used muon target assy. was designed to be replaced with new one. In the hot-cell, the used assy. would be disassembled by the manipulator, and then the rod be cut by the cutting device to reduce its volume. But the situation has been changed due to
  Treatment of the tritium becomes more serious  use of green-house, local exhaust ventilator, airline mask, anorak suit
  The life time of the target assy. is successfully prolonged to 10 years.
    Only several target assemblies will be used during the designed life time of MLF facility of 40 years.
    Original scenario was constructed to deal with a fixed target that would be spent every half a year under 1 MW beam. Two targets would be spent every year.
  Changing the human resources  simple scenario
We just started reconstructing the maintenance scenario, minimizing the consumption of time, money and human resources.

## Slide 24: Materials and Life Science Facility
RCS (3GeV333mA)
MLF
Beam transport tunnel
Exp. Hall #1
Exp. Hall #2
Muon target
Spallation neutron target
Exp. Hall #1
Exp. Hall #2
100%
>90%
graphite
mercury
m
n
Cascade target

## Slide 25
Proton beam transport tunnel
Exp. Hall #1
Exp. Hall #2
M1 & M2 tunnel
Exp. Hall #2
Exp. Hall #1
70m
30m
146m
22m
proton

## Slide 26: Muon Facility in MLF
D-line
U-line
Slow（4 MeV）-fast（50 MeV）, general-purpose beamline with 2 exp. areas.
Ultra slow beam （0.1-30 keV）, near-surface, sub-micron scale condensed matter physics, chemistry, etc.
µ+
µ±
S-line
Slow beam （4 MeV）,
dedicated to bulk µSR ultralow tem-perature/high magnetic field/ pulsed excitations.
µ+
H-line
Slow（4 MeV）-fast（50 MeV）beam, for particle physics, atomic physics (“precision frontier”)
µ±

## Slide 27: History of the target operation
Fixed target
Rotating target #1
Oct. 2014
27
Rotating Target #1
Operation time: 17000 hours Nov. 2014 ~ July. 2019 6000 MWh
Rotation : ~16 M revolutions 4.5 M revolutions /year @ 15 rpm
Stable 500 kW operation in 1 year
1 MW test beam operation
Rot. system trouble
Users’ runs were canceled for 2 weeks.

## Slide 28
Remote handling attachments for muon components
Target rod, Profile monitor rod, and CT rod are exchanged in Hot cell.
Plug stand attachment is equipped on Plug stand.
Each assemblies are put down on Plug stand attachment by Muon cask.
Rod attachment is equipped on Exchange device.
Each used rods are hooked and removed by rod attachment.
Both attachments are common for three components.
Attachment will be delivered this February.
Commissioning of target exchange will be completed this April.
Off-beam commissioning

## Slide 29: Muon target assembly
M2 tunnel
Muon target assy
Rotation speed: 15 rpm
Duration time: 5 years
Accumulated beam power: Rotating target: 5850 MWh
replaced
To be replaced

## Slide 30: Shaft-joint / Target assy. replacement
The old shaft-joint has the greatest risk for the stable operation, and thus we decided to replace it in this summer maintenance period.
The lower side joint was not designed to be replaced originally, and thus the replacement work was expected to be a fight against time under the high radiation field.
Replacement work in M2 tunnel  impossible The target assy. is moved to the pit that is originally for temporary storage of a new/old target assy. We prepared the special scaffold on the pit.
We prepared 2 options: fix the 1st target/install the 2nd target Plan A: If the dose from the assy. is enough low, joint would be replaced. Plan B: If not, new target assy. (Rot. Target #2) would be installed.
Finally, we found that the upper part of assy. was highly activated than we expected. We took plan B.