## Slide 1
Katsuya Yonehara
Gas RF detector Meeting
07/25/2019
Progress of RF beam detector R&D

## Slide 2: Beam Monitor in Neutrino Beam Line
Use multi-pixel ion chamber as beam monitor
Primary function of Hadron Monitor
  Monitor beam centroid and deterioration of target
  Evaluate systematic of beam parameter
Present Hadron Monitor shows low gain due to radiation damage
  Only use for aligning beam element (baffle, target and horns)
  Need rad-hard Hadron Monitor
NuMI Target system

## Slide 3: NuMI Hadron monitor
2’
27’
17’
~15 mrem/h after 10 days cool time

## Slide 4: Old Hadron Monitor (1x1 m2, 49 pixels)
First Hadron Monitor
Second Hadron Monitor
Backflow gas
contaminated
He gas
No outstanding damage

## Slide 5: Present hadron monitor signal
Log10[x]
Beam view
Blueish color pixel is low gain
Central pixel still useful to align
beam elements
I0 = 3.05
I = 13.79
m = 5.673
s = 0.85
I0 = 2.11
I = 2.07
m = -8.095
s = 1.190
I0 = 4.27
I = 11.77
m = 2.618
s = 0.778
Example of vertical beam scan
Found:
Baffle = -1.211 mm
Target = 2.618 - 3.7 = -1.082 mm

## Slide 6: Rad-hard RF detector (test cavity)
2.4 GHz gas-filled RF resonator
Detector and wave guide made of aluminum for rad-hard
Test cavity can be pressurized up to 2 atm
Beam ionizing gas and form plasma in the cavity
Plasma consumes RF power
Consumption is proportional to the amount of incident beam
Shunt impedance () can be measured without beam
It suggests that beam intensity can be reproduced without beam calibration
RF power consumption
at equilibrium condition
QL ~ 240

## Slide 7: Test Parameter for the First beam test
Proof Of Principle test (Spring 2019)
N2, Dry Air, He gas
Gas pressure = 1, 2 atm
V = 100, 150, 200 V in cavity gap
Beam per pulse = 1-4e12 (or 30e12) protons/pulse (6 batch)
  Bunch gap = 18 ns
  Batch gap = 20-40 ns
Pulse duration = 1.6 (or 10) micro sec (6 batch)

## Slide 8: Circuit diagram
NWA
Orange
20 dB
1 W
10 dB
5 W
White
Blue
Green
Filter
2.45 GHz
resonator
Real-time
RF power meter
N-type to SMA
SMA to N-type
Red
port1
port2
1.6 dB
DAQ
Huge insertion loss in a long ½’’ heliax cable
Used a low-level DAQ which only has a basic function

## Slide 9: Plan to upgrade DAQ system
RTP data is taken by using a vender provided DAQ system
  Easy to use
  But no tuning knob, e.g. I need to update the file name and saving location for each event
  It limits the data taking speed and someone has to be there
  Use self-trigger mode but the DAQ is often triggered by noise
Consider LabView base DAQ system
  File name can be updated automatically
  Fast data taking
  Take Q factor measurement automatically
  Remote control
  Handle external trigger signal
  Sync to beam parameter

## Slide 10: Kyle’s presentation

## Slide 11: Circuit diagram
NWA
Orange
20 dB
1 W
10 dB
5 W
White
Blue
Green
Filter
2.45 GHz
resonator
Real-time
RF power meter
N-type to SMA
SMA to N-type
Red
Unique structure of two sets of attenuator and RF amp to make a linear gain and the first RF amp has a narrow band pass to remove RF noise
port1
port2
1.6 dB

## Slide 12: Linearity of gain of RF amp
n = 2.385 GHz
2.390 GHz
2.379 GHz

## Slide 13: Q factor measurement
RF signal calibration
  Q factor measured by decay time
  Q factor measured by spectrum
  Accuracy in both measurements is within ~10 %
Forward
Reflected
Pickup

## Slide 14: Beam loaded RF signal
Beam is turned on
Five peaks during the beam on shows
the gap of MI beam batches
Beam intensity = 1.3e13
Detector filled with ambient air

## Slide 15: Linearity test
1-atm N2, V0 = 120 V
2-atm dry air, V0 = 120 V
1-atm He, V0 = 120 V
2-atm N2, V0 = 60 V

## Slide 16: Discussion of analysis
Took average of whole equilibrium voltage for first analysis
  Nice linearity vs beam intensity!
  Fraction is typically 10 %
RF detector is very sensitive to the beam intensity
  But beam intensity is changed in one batch
    Confirmed with MI group
  Beam position is also changed bunch by bunch
    SWIC data recorded
    Will calibrate RF signal with beam position and profile
RF detector is also sensitive to resonant condition
Find optimum E/p
Gas dependence
  N2 & dry air
Intensity varied
+2 MHz off
-2 MHz off

## Slide 17: Study Plasma-Beam interaction
Our plasma condition is very unique
  Cold neutral beam-induced plasma
  Dense plasma in dense neutral gas
    Differ from plasma wakefield accelerations
Need fast time RTP (?)
  Present time resolution is 20 ns
Need precise beam parameter measurement
  Time structure of beam (time vs beam intensity)
  Arrival time
  Position

## Slide 18: Next step
Make a plan to carry out beam test at MI-40
  Study influence of wave guide on the RF signal (see next slide)
    Relocate the RF detector to directly hit beam on the wave guide to study a WG noise
  Long run
    Move back the detector to hit beam to the detector center
    Endurance test
    Take more data to study beam-plasma interaction
Design High Q cavity (see next slide)
Improve DAQ
  Triggered by the MI beam
  Take Q measurement every spill
  Write a script for automatic data acquisition

## Slide 19: Design concept of new RF Detector
Design high-Q cavity
  In the cavity the gradient is 6.57 V/cm, in the Iris coupling hole is 3.2 V/cm, in the rest of the WG is 1.0 V/cm (see next slide)
  Possible power consumption on the WG is ~1/36 lower than the cavity, but the volume of WG is much bigger than the cavity
    Iris is fine since it is localized (take a part of the sense area)
  What is Q to reduce the gradient on WG to below 10 % of the cavity gradient?
N2 and dry air
< 100 V RF potential looks good
  Need to be higher than the beam noise

## Slide 20: Study influence of ionization in waveguide
Cell body
Waveguide
+
+
+
+
(+) Possible beam position to test the off-axis beam effect