## Slide 1: Strand verification testing at the NHMFL
Jun Lu and Jeremy Levitan
Magnet Science and Technology Division, NHMFL
Strand Production Readiness Review, 22 Aug 2019

## Slide 2: About our team
Team members
  Jun Lu, PhD physics,1999, Nb3Sn testing since 2006
  Jeremy Levitan, B.Sc. Mech. Eng. 2015, Nb3Sn since 2016
Relevant project experiences
  RRP Nb3Sn quality assurance tests for NHMFL series-connected hybrid magnet.
    Heat treatment, Ic, RRR, hysteresis loss, etc.
  The only ITER Nb3Sn verification lab in the US
    Heat treatment, Ic, RRR, hysteresis loss, etc.
    Tested total of over 1600 samples.
    Completed in time under budget.
Strand Production Readiness Review, 22 Aug 2019

## Slide 3: Outline of this talk
Project management
The test facility
  Heat treatment
  Critical current test
  RRR test
Test results
Summary
Strand Production Readiness Review, 22 Aug 2019

## Slide 4: Flow chart of each verification test cycle
Samples and test requests from Lance Cooley
Receiving inspection/ Labeling if necessary
Heat treatment and tests
Report Ic and RRR to Vito Lombardo (CAM) by email; copied to Cooley
Archive the raw data
Send remaining wire to LBNL archive
Strand Production Readiness Review, 22 Aug 2019
Verification samples are taken by supplier at the same time, and adjacent to, QC samples. Samples are shipped to NHMFL  Cooley

## Slide 5: Project management
Procedures are well developed.
  Benchmarking with OST during NHMFL SCH project.
  ITER world-wide benchmarking; ITER annual x-checks.
  Benchmarking within LARP.
Data reporting and security
  Data are reported in a timely manner to Vito, copied to Lance.
    Vito is CAM; Work at NHMFL is a sub-contract to FNAL
  Raw data are secured in NHMFL network and can be accessed only by the team members.
ES&H: in compliance with the safety practice of the NHMFL which implemented an integrated safety management system.
Risks:
  Magnet failure. Another magnet at the NHMFL may be used temporarily to mitigate this risk.
  Hurricane: Tallahassee has no major flooding threat, therefore delay should be no more than 2 months.
Strand Production Readiness Review, 22 Aug 2019

## Slide 6
4 furnaces: 6” diameter, 36” long 3 zone furnaces. Each can take 30 Ic and RRR samples.
Inconel retort, Ar purge and flow (100 SCCM). 4 psi positive Ar pressure.
Ic and RRR samples heat treated at the same time.
LabVIEW data monitoring. 3 thermocouples over 12” central zone.
Furnace
Temperature Trending Program
Heat treatment facility
Strand Production Readiness Review, 22 Aug 2019

## Slide 7: A typical T trend during heat treatment
Chart: Time (h)
Three K-type thermocouples reading within 5 deg.
Thermocouples are replaced annually.
Strand Production Readiness Review, 22 Aug 2019

## Slide 8
Oxford 52 mm bore 15/17 T magnet, 0.1% field homogeneity.
Ametek 0-1000 A PS, NI SCXI-1000 DAQ, LabVIEW programs.
Ic probes capable of up to 1000 A at 4.2 K.
Oxford 15/17 T
I source
Magnet power
NI SCXI
Ic testing facility
Strand Production Readiness Review, 22 Aug 2019

## Slide 9
Standard ITER Ti-6V-4Al barrel. 4 slotted roll pins to connect to each Cu terminal ring.
Pressure contact between sample and the probe,
Both 25 and 50 cm (reported) V-taps.
Chart: I (A)
Ic test probe and sample mandrel
Strand Production Readiness Review, 22 Aug 2019

## Slide 10: Ic sample prep
Heat treat on the test mandrel. No wire transfer after HT.
1 kg winding tension before heat treatment.
After HT, wire is loose. retighten carefully by pressing the wire onto the mandrel while moving along the helical groove. It’s a delicate operation and critical.
Retightening is necessary to a reliable and accurate Ic results.
Strand Production Readiness Review, 22 Aug 2019

## Slide 11
Residual Resistance Ratio (RRR)
RRR probe
Samples
Cernox
Four probe R vs. T measurement with +/- 1A current input.
Cool down to 4.2 K, take measurements while samples naturally warming up.
Eight 150 mm straight sample in contact with a horizontal G-10 plate.
A Cernox temperature sensor glued on the G-10 plate. Warm up rate < 20 mK/s.
R vs. T data are analyzed to obtain R273K/R20K or R295K/RIEC (Hi-Lumi)
Strand Production Readiness Review, 22 Aug 2019
Notes: Z:\MS&T\Materials Development and Characterization\Physical Properties\Other projects\Other data\Resistance measurements\RRR\ACE 2mm 11-14-05
Z:\MS&T\Materials Development and Characterization\Physical Properties\Other projects\Other data\Resistance measurements\ITER Copper Perer Bauer

## Slide 12
IEC RRR data analysis program
Allows efficiently analyze large number of data sets.
Strand Production Readiness Review, 22 Aug 2019

## Slide 13: Ic results
FNAL did not convey the change of HT time from 75 h to 50 h. So our Ic data were systematic difference from those of the supplier.
After 4.7% and 2.8% corrections for 15 T and 12 T data*, the results are in agreement with the supplier’s.
* J. Levitan, et al., IEEE Tran. Appl. Supercond. (2019) 6000904
Strand Production Readiness Review, 22 Aug 2019

## Slide 14: RRR results
Systematic differences are due to the difference in heat treatment and RRR definitions,
Strand Production Readiness Review, 22 Aug 2019

## Slide 15: Summary
NHMFL has ample experiences of testing large volumes of Nb3Sn strands for quality assurance.
The existing facilities and personnel are well prepared and meet the requirement of the quality plan.
Data security, personnel safety and project risks are evaluated.
We have successfully tested 107 samples for LARP. So far, our results track those of supplier’s well.
A testing work flow path has been established. Tests and reports have been timely.
Strand Production Readiness Review, 22 Aug 2019

## Slide 16: Outlook
Average of 100 tests per year is well within the capability of our facility and personnel
A batch of 20-50 samples will take about a month to complete.
Helium price has doubled since our last contract.
  Cost per test remains below that at FNAL (Opportunity RO-302-2-02-003)
Strand Production Readiness Review, 22 Aug 2019
Acknowledgement
NHMFL is supported by NSF through NSF-DMR-1157490 and 1644779, and the State of Florida.