## Slide 1: LCLS-II LLRF Design, Production & Ongoing Tests
Andy Benwell on behalf of LCLS-II LLRF Collaboration
LLRF 2019
10/3/2019

## Slide 2: Outline
LLRF overview
Recent development and testing
3.9 GHz preparation
Hardware status
Summary

## Slide 3: LCLS – II
  The Linac Coherent Light Source (LCLS-II), in Menlo Park, CA is designed produce a 4 GeV electron beam at high repetition rate (up to 1 MHz) for Soft X-ray (SXU) and Hard X-ray (HXU) Undulators
  In order to accomplish this, SLAC will install 280 1.3 GHz SRF cavities and 16 3.9 GHz SRF cavities in the first third of the 2-mile accelerator housing
  Superconducting RF is new for SLAC – we have benefitted greatly from the LLRF collaboration

## Slide 4: LCLS–II LLRF System Overview
  Hardware/Firmware
    Precision Receiver Chassis (PRC)
      Exclusively for cavity signals
      Fiber communication for high Isolation necessary for 0.01o/0.01%
    RF Station (RFS)
      Generates independent RF signal for 2 cavities
      Forward, Reverse and SSA Drive signals
      Detune calculation, PI loop control
    Resonance Control
      Piezo Control, Stepper Motor Control
      Temperature Monitoring
      Interlock processing
  Software
    Automated Scripting
      Cavity and SSA characterization
    Channel Access
  EPICS
    All control displays and waveform readouts

## Slide 5: LLRF Rack System
Down
Converter
Digitizer
Up
Converter
QF2-pre
Digital Carrier
PS Board
QF2-pre Digital Carrier
FMC Breakout
Stepper tuner
driver
Coupler/Stepper
Temperature
Piezo tuner
drivers (x4)
Interlock
digital I/O
PS Board
Resonance Control
RF Station

## Slide 6: 2 years ago
Prototype testing at FNAL
Controlled multiple cavities simultaneously in SEL and GDR mode on a cold cryomodule
  Semi-automated cavity turn-on scripts find phase in SEL and ramp amplitude, even with significant low field detuning
  Out-of-loop phase error < 2x10-3 deg rms
In the last year, efforts have been made toward
preparing the system functionality for commissioning
improving user experience
supporting cryomodule testing with custom applications
Notes: This slide demonstrates that the LLRF system can control a cryomodule with outstanding performance.
Phase noise of the LLRF system is exceptionally low.
The LLRF system can control gradient with a PI loop with high precision.
PI coefficients can be even further optimized, these were set manually to demonstrate the system performance
A step change was introduced at the same resolution of the amplitude regulation spec. The over shoot is tiny for this step change on the order of 100 us
Make a point to mention Larry’s (the architect) talk who is the expert on this system, and its tests, Larry will discuss these tests and more in detail

## Slide 7: Recent Test History - SLAC
SLAC LLRF rack system used for production hardware testing and rapid software prototyping
Full set of LLRF hardware
Test loads to simulate cryomodule
    Narrowband low power cavity
    Stepper motors, piezo tuner loads
Linux RT computer
Timing crate
Great for hardware testing and software prototyping but cannot support testing unique issues of SRF cavities

## Slide 8: Recent Test History – JLab LERF
JLab LERF facility
  RF supported exclusively by LCLS 2 design
    All production equipment must work!
  Mandate for system to be as close to production as possible
    All modes of operation (SEL, SELAP, Pulsed)
    SSA and LLRF control
    User interfaces
    Calibration routines
    Production IOC
    PV naming conventions
    Tuning scripts
    Operator and expert interfaces
  Cryomodules controlled by physicists and operators
    All difficulties and issues immediately communicated
    Tools as requested and deemed necessary for SRF cavity evaluation

## Slide 9: Control improvements
Characterization
Calibration
Steady state operation (optimization)
Interlocks
Resonance Control

## Slide 10: Characterize System and Go
4 “easy” steps
SSA calibration
Chirp detune scan
  Immediately “find” cavity and enter SEL
  Enable faster tuner response to center cavity at nominal frequency
Pulsed SEL calibration
  Measures loaded Q
  Calculates cavity scale factor, based on energy integration
Set up to target field in SELAP mode
Cavity Forward Reverse

## Slide 11: RF system calibration

## Slide 12: RF control and optimization
Expert level controller
Phase and amplitude P & I coefficients
Target stability summary
Reactive and real power status and clip limits
If a cavity frequently clips into SEL mode, envelope settings can be adjusted
  Real microphonics not suppressed here

## Slide 13: Firmware Based Interlocks
All fault information embedded into firmware
  SSA permits
  IOC watchdogs
  Stepper tuner temperatures
  Limit switch status
  Coupler temperatures
  Vacuum status
  Cryo-system summary
  Quench detection
Enables fast response and fault waveform capture
FPGA detected quench fault
  Waveform acquisition halted and displayed (with appropriate time scaling)

## Slide 14: Resonance Control
Piezo tuner feedback loop keeps cavity tuned for long periods of time
Stepper motor can be used to re-center the piezo tuner as needed
Piezo tuner DAC waveforms on acoustic time frames

## Slide 15: Detune waveforms
Resonance Control waveforms exported
Cavity detune information displayed
FWD power to maintain gradient

## Slide 16: 3.9 GHz Preparation
  A 3.9 GHz version of the LCLS 2 LLRF system is expected at Fermilab this fall
    Hardware architecture similar to 1.3 GHz chassis
    Up and down conversion separated into two chassis due to size of connectorized components and reduced total quantity
    Resonance Control is identical to 1.3 GHz
    Chassis performance must be similar to 1.3 GHz system
  This is the only LLRF system being built for 3.9GHz Harmonic Linearizer CM testing
LERF development has enabled this system to be ready for 3.9 GHz cryomodule testing

## Slide 17: Production Hardware Status
  LCLS – II is composed of 37 superconducting cryomodules
    296 SCRF cavities
  Single Source Single Cavity for precision field control with high Q
  SSSC dramatically increases RF system component count
        76 total LLRF racks
          236 RF chassis (RFS or PRC)
          74 resonance control chassis
          78 power distribution chassis
          84 LO timing distribution chassis

## Slide 18: Production Hardware Status
  All 1.3 GHz racks have been built, loaded, and installed in LCLS2 gallery
  Fiber optic cabling
  Internal RF cabling
  Power Distribution
  Internal plumbing
  Mechanical infrastructure (rails, panels, strain relief)
  Racks will undergo full LLRF checkout and calibration in field as soon as AC power is available

## Slide 19: Summary
LCLS 2 LLRF system has demonstrated its ability to meet design specifications
The control system and user interface have undergone tremendous progress towards commissioning readiness
Over 20 people (SLAC and experts from other labs) have run cryomodules at LERF
  Their pokes, prods, and comments and been incredibly valuable
3.9 GHz system has been built, and partially tested
  It is the next system to be commissioned as part of cryomodule testing at FNAL
Gun and Buncher LLRF have already transitioned to operations successfully
Heading toward commissioning at SLAC, hardware installation is progressing well
There is still plenty of work to do!

## Slide 20: Lots of ongoing development
Kintex Ethernet over fiber (high speed data transfer)
LCLS 2 timing system integration
Beam Synchronous Acquisition
Machine Protection System protocol integration and testing
Phase averaging Reference Line
Improved diagnostics on existing waveforms
Buffer freeze during non-trip events (e.g. RF clipping)
Improve/automate feedback gain
Improve clip limit interface
And if the project becomes interested
Active Resonance Control Implementation
  Previously demonstrated at Fermilab, but no automation

## Slide 21: Questions?
LBNL
K. Campbell
Q. Du
M. Davidsaver*
L. Doolittle
J. Greer
G. Huang
C. Serrano
V. Vytla
Y. Xu
FNAL
B. Chase
E. Cullerton
J. Einstein
Dan Klepec
JLAB
R. Bachimanchi
C. Hovater
O. Kumar
D. Seidman
SLAC
A. Benwell
M. Boyes
G. Brown*
J. Diaz
B. Hong*
S. Hoobler*
A. Johnson
J. Jones*
R. Kelly
A. McCollough
P. Nguyen
M. Petree
A. Ratti*
D. Steele
J. Sikora*
*not in photo
Thank you!!!