## Slide 1
Kathrine Cipriano
October 16th, 2019
Iceberg Updates

## Slide 2: Additions to Iceberg System
New LN2 flow control valve
  Smaller flow coefficient
  Reduced pressure fluctuations to ~0.1 psig (was previously +1 psig)
  PID loop tuning could still improve, reducing pressure fluctuations even further
New backpressure regulator at LN2 condenser outlet
  Increased LN2 condenser pressure from atmospheric to 25 psig
  This increases the temperature of the LN2 and prevents argon ice build-up inside the condenser
  Maintained ~54in LAr in the cryostat with condenser LN2 level at ~3.5in
Internal filter installed
  Uses same set up as Luke internal filter (actuated vapor outlet valve and liquid inlet valve)
  Cycles are based on liquid level, not time

## Slide 3: Plot of Cryostat Steady State from 10/01 – 10/03
K. Cipriano | Iceberg Updates

## Slide 4: Test of Condenser LN2 Level and Cryostat Pressure
K. Cipriano | Iceberg Updates
LAr Level
Cryostat Pressure
Cond GAr Pressure
I.F. Pressure
I.F. Level
LN2 Level

## Slide 5: Condenser Trend
K. Cipriano | Iceberg Updates
Condenser starts working, then condenser argon gas pressure starts to decrease, then condenser stops working

## Slide 6: Condenser Return Tee
K. Cipriano | Iceberg Updates
Current set up
Proposed set up
Sintered metal filter: Media grade 5, 2ft long, 0.062 wall thickness, 1.5in OD

## Slide 7: Calculations
Cryostat heat load and resulting boil off rate
  1000 W – 50.51 lb/hr
Volumetric flow rate of LAr from condenser back to cryostat
  0.074 gpm
Required cross-sectional area for LAr to flow from condenser to cryostat
  0.0119 in2
Cross-sectional area of narrowest part of condenser tube
  0.196 in2
Argon gas pressure drop from cryostat to condenser
  5.5 x 10-8 psi
Liquid pressure drop across sintered metal filter (at condenser outlet)
  0.0021 psid
Height of liquid head required to push LAr through sintered metal tube
  1.083 mm
  This indicates that the filter is being blocked by water ice
  We plan to add molecular sieve media inside the sintered metal to absorb more water
K. Cipriano | Iceberg Updates

## Slide 8: Internal Filter
K. Cipriano | Iceberg Updates