## Slide 1: CRYO: A Waveform Digitizer/Serializer for TPC Experiments
Design Changes for Next Version
Feb 5th, 2020
A. Pena Perez, A. Gupta, D. Doering, P. Tsang, M. Convery, A. Dragone*
(*dragone@slac.stanford.edu)
DUNE TPC Electronics Review
CERN - Geneva, Switzerland, 5th – 7th February, 2020

## Slide 2: About the new CRYO version
A new version of CRYO has been developed and submitted during the fall last year (2019)
This version is aimed to meet the noise specifications of the nEXO experiment and addresses the issues observed during the characterization of the first iteration
As described in the following sections, the new modifications and improvements include:
  Resized front-end preamplifier to meet noise specifications at the target Cdet (40pF)
  Improved noise contribution of baseline holder
  Modified digital back-end to eliminate “ghost effect” observed on the first iteration
  Optimized capacitor array of the ADC to improve ENOB and linearity performance
  Improved analog monitor module with input/output rail-to-rail operation to fix saturation at cold temperature
  Mode for low power operation on the ADC section to achieve ~20% overall power reduction. This feature can be enabled or disabled using SACI interface
  Improved synchronization scheme to latch CRYO with both the external master clock and the external enable signal to stream data out (Start-of-Readout)
The new chip is expected to be back mid-February

## Slide 3: OUTLINE
Comparison of CRYO Specifications
Modifications and Improvements
  Front-End Channel
  Ghost Effect Mitigation
  ADC Optimization
  Improved Analog Monitor
  Power Breakdown
  Synchronization Scheme

## Slide 4: Comparison of CRYO Specifications
TPC Experiment | DUNE |  |  |  | nEXO |  |  | 
CMOS process | 130nm |  |  |  | 130nm |  |  | 
Supply voltage | 2.5V (2V, 1V internal) |  |  |  | 2.5V (2V, 1V internal) |  |  | 
Input capacitance | ~ 220pF |  |  |  | ~ 40pF |  |  | 
Anti-aliasing filter | 5th order Bessel architecture |  |  |  | 5th order Bessel architecture |  |  | 
Peaking times | 0.6us, 1.2us, 2.4us, 3.6us |  |  |  | 0.6us, 1.2us, 2.4us, 3.6us |  |  | 
Gain settings | 6.0X (28.6mV/fC) | 3.0X (14.3mV/fC) | 1.5X (7.1mV/fC) | 1.0X (4.8mV/fC) | 6.0X (57.2mV/fC) | 3.0X (28.6mV/fC) | 1.5X (14.3mV/fC) | 1.0X (9.6mV/fC)
Max. input charge | 50fC | 100fC | 200fC | 300fC | 25fC | 50fC | 100fC | 150fC
Noise | ~500e- @ 3.0X and 1.2us |  |  |  | ~200e- @ 3.0X and 1.2us |  |  | 
Resolution | 12 bit |  |  |  | 12 bit |  |  | 
INL and DNL | ±1LSB |  |  |  | ±1LSB |  |  | 
Sampling frequency | 2MSPS |  |  |  | 2MSPS |  |  | 
Temperature | 87K (-186ºC) |  |  |  | 160K (-113ºC) |  |  | 

## Slide 5: Simulation methodology
CRYO has been designed under different simulation conditions to verify functionally and performance on all critical blocks implementing the ASIC
The evaluation takes into account PVT variations and advance analysis such as noise in AC and transient; Monte Carlo mismatch, etc; and they have been applied to both schematic level and extracted view (PEX) of layout
Digital flow implements co-simulation integrating ASIC firmware, FPGA firmware and software
  The simulation environment includes Vivado, VCS and python based tools
In the following a subset of simulation results considered during the design of CRYO are reported
Simulation conditions for the various sections | Schematic | PEX view
Supply voltage variation | ✔ | ✔
Temperature: Room, nEXO and DUNE | ✔ | ✔
Corners: SS, FF, FS, SF | ✔ | ✔
Clock speed | ✔ | ✔
Monte Carlo mismatch | ✔ | ✔
AC Noise Analysis | ✔ | ✔
DC analysis | ✔ | ✔
Transient | ✔ | ✔
Trannoise | ✔ | ✔
Top level simulations | ✔ | 

## Slide 6: OUTLINE
Comparison of CRYO Specifications
Modifications and Improvements
  Front-End Channel
  Ghost Effect Mitigation
  ADC Optimization
  Improved Analog Monitor
  Power Breakdown
  Synchronization Scheme

## Slide 7: Front-End Channel
Preamplifier
Preamplifier
DUNE Variant (Current Version)
nEXO Variant (New Version)
  Modifications
  Although the preamplifier of the front-end has been resized to meet noise specifications of the nEXO experiment simulations reported here were done at -186ºC
  Also, additional improvement in terms of noise was done in the baseline holder
  The layout of the nEXO channel follows the same footprint as the DUNE variant

## Slide 8: Front-End Channel – Collection Mode
  Simulation Conditions
  Baseline of the front-end channel is set to collection mode
  Plots show the ENC (left) and transient response (right) across different peaking times
  The analog LDO supplies the channel and temperature is -186ºC
  For this case the detector capacitance is 40pF with input charge of 25fC and gain of 6.0X

## Slide 9: Front-End Channel – Induction Mode
  Simulation Conditions
  Baseline of the front-end channel is set to middle range (induction mode)
  Plots show the ENC (left) and transient response (right) across different peaking times
  The analog LDO supplies the channel and temperature is -186ºC
  The detector capacitance is 40pF with dual polarity input charge of ±12.5fC and gain of 6.0X

## Slide 10: Front-End Channel – nEXO vs DUNE with Cdet = 220pF
  Simulation Conditions
  The ENC comparison between the DUNE preamp an the nEXO preamp is shown in the left plot
  Notice that for the given input charge (100fC), the gain of the nEXO variant is 1.5X whereas for DUNE is 3.0X
  Right plot shows the ENC of the DUNE front-end channel when it benefits from the revised baseline holder. This yields slightly noise improvement compared to the nEXO variant
  In both plots the detector capacitance is 220pF and temperature is -186ºC. Channels are supplied by LDO
1.5X (100fC)
3.0X (100fC)

## Slide 11: Front-End Channel – ENC Vs Cdet
  Simulation Conditions
  The plot shows the ENC vs Cdet of the nEXO front-end channel across different peaking times
  Detector capacitance ranges from 0pF to 250pF and input charge is 25fC
  The channel is set in collection mode with gain setting of 6.0X
  The supply is provided by the LDO

## Slide 12: Front-End Channel – Saturation Response
Characteristic shape of “Ledge Effect” was observed in LArASIC when front-end channel is saturated with a large amount of charge
DUNE management asked us to perform this simulation on all versions
Transient simulations were done in the CRYO front-end channel using similar saturation conditions in collection and induction mode and across corners to evaluate this behavior
Image taken from:
Hucheng Chen, et. al.,
Study of the “Ledge” Effect in ProtoDUNE Readout

## Slide 13: Front-End Channel – Saturation Response
Transient simulation conditions
Applied large input charge (Qin = 250fC) with pulse duration of 50us
Corners (SS,TT, FF) at different peaking times and gain set to 6.0X
Temperature is -186⁰C
Monte Carlo mismatch at -40C using model card 1
Results show that the “Ledge affect” is not observed in both modes collection and induction
Corners (SS, TT, FF)
4 Tps | -186⁰C
Corners (SS, TT, FF)
4 Tps | -186⁰C
Monte Carlo Mismatch | Corners
4 Tps | -40⁰C
Monte Carlo Mismatch | Corners
4 Tps | -40⁰C
1 Our current cryogenic models do not support Monte Carlo analysis

## Slide 14: Digital Back-End – “Ghost Effect” Solution
  “Ghost effect” observed during the evaluation of the first version of CRYO was fully understood
  This issue was reproduced in simulations using extracted view (PEX) of the layout of the digital back-end
    PEX view involved the digital MUX, encoder, serializer and LVDS TX modules
  The poor driving capability of set of latches implementing the digital MUX was the main reason of this issue
  The second prototype fixes this problem by properly buffering the output of latches
  Special care was taking into account during the layout implementation of this section

## Slide 15: Digital Back-End – Test Bench
  Simulations across PVT, corners and clock speed were performed to ensure that this problem was not present in the second version of CRYO
  Different test signals in 12 bit format were applied at the inputs of digital MUX. Some examples include constant values and dynamic signals such as ramps and sine waves using the ADC
  A dedicated Python script deserialized and decodified the data output stream of the LVDS TX to be then compared against the input signal
  Next figures shows the example of two different cases:
    Constant values in all inputs (i.e. IN0 = 1, IN 1= 2, etc)
    Zero value at inputs from IN0 to IN6 and half sinewave on IN7 using the ADC
Examples of Input Signals
12b
12b
12b
12b
Encoder
(12b/14b)
Multiplexer
12b
14b
Serializer
12b
12b
12b
12b
Stream
Deserializer
Decoder
Python
LVDS TX
ADC
IN1
1
IN2
2
IN3
3
IN4
4
IN5
5
IN6
6
IN7
7
IN8
8
EQUAL

## Slide 16: Digital Back-End – Test Bench
Examples of Input Signals
12b
12b
12b
12b
Encoder
(12b/14b)
Multiplexer
12b
14b
Serializer
12b
12b
12b
12b
Stream
Deserializer
Decoder
Python
LVDS TX
ADC
IN1
1
IN2
2
IN3
3
IN4
4
IN5
5
IN6
6
IN7
7
IN8
8
EQUAL
Corrupted Data
Correct Data
(After Fix)
Simulation results at -186⁰C
Notice that last latches of DMUX failed as they need to drive longer connections due to layout placement

## Slide 17: Digital Back-End – Test Bench
Examples of Input Signals
12b
12b
12b
12b
Encoder
(12b/14b)
Multiplexer
12b
14b
Serializer
12b
12b
12b
12b
Stream
Deserializer
Decoder
Python
LVDS TX
ADC
IN1
0
IN2
0
IN3
0
IN4
0
IN5
0
IN6
0
IN7
0
IN8
-
EQUAL
Corrupted Data
Correct Data
(After Fix)
Simulation results at -186⁰C

## Slide 18: Digital Back-End – Test Patterns
  The back-end module of CRYO has built-in pattern generators specially intended for testing and debugging
  Different modes, controlled by SACI, were implemented to bypass cells and generate unique data patterns
  Although few of these patterns did not work properly on the first iteration, they do not have any impact on the signal chain and functionally of CRYO
  The new version fixed the bugs found in the first iteration
Mode | Pattern | Description
0 | N/A – ADC data | ADC data  DMUX  Encoder  Serializer  LVDS TX (Full chain of digital back-end)
1 | Patt1: 0xA’s ↔ 0x5’s | Patt1  Encoder  Serializer  LVDS TX
2 | Patt2: 000’s ↔ 111’s | Patt2  Encoder  Serializer  LVDS TX
3 | Patt3: Ramp | Patt3  Encoder  Serializer  LVDS TX
4 | N/A – ADC data | ADC data  DMUX  Serializer  LVDS TX
5 | Patt1: 0xA’s ↔ 0x5’s | Serializer/LVDS TX
6 | Patt2: 000’s ↔ 111’s | Serializer/LVDS TX
7 | Patt3: Ramp | Serializer/LVDS TX
Test modes of digital back-end

## Slide 19: ADC Optimization
  Optimization on the capacitor array (I)
  The optimization was aimed to improve both resolution (ENOB) and linearity performance of the ADC
  These optimization was performed on the capacitor array of the ADC, which is based on a binary weighted split capacitor DAC implemented by MSB and LSB section
  Due to routing of interconnections, critical parasitic capacitors that impact performance were carefully identified from the extracted view (PEX) of the layout
  Given the results we pushed the minimum parasitic cap considered for extraction down to < 0.001fF

## Slide 20: ADC Optimization
  Optimization on the capacitor array (II)
  These critical parasitics appeared in parallel to the main capacitors of the array; from top plates to ground in the MSB and LSB sections and in cross-coupling around the bridge capacitor
  The error introduced by those parasitics was substantially minimized by properly trimming the routing of interconnections
  After optimizations, the parasitic caps in parallel to the main capacitor array followed a binary weighted arrangement with mismatch error <0.01%; the crossed coupling capacitor was removed and parasitics with respect to ground were reduced by more than 50%
OPT

## Slide 21: ADC Optimization – ENOB
Temperature = -186⁰C
FFT = 1024
Fin = 242.187kHz
Fs = 8MHz
Temperatures | Transient | Transient + Noise Noise BW: 10 kHz – 10 GHz
Room | 11.66 bit | 11.38 bit
nEXO (-113ºC) | 11.67 bit | 11.42 bit
DUNE (-186ºC) | 11.71 bit | 11.43 bit
Simulations were done using PEX views of the ADC
Without any optimization the simulated ENOB is in good agreement with measurement results (~10 bits)
After optimization the ADC achieved an ENOB = ~11.40 bits (transient + noise) across temperatures
ENOB Optimized PEX

## Slide 22: ADC Optimization – Linearity
Linearity Test of ADC
Sinewave method was used in simulations to evaluate static performance (INL/DNL)
It uses a slow sinewave input signal (i.e. ~ 30kHz) with long number of periods to exercise all codes (i.e. 127 with FFT = 32768)
Results show the performance of the optimized PEX view of the ADC
DNL and INL are within ±1LSB
Transition around middle range in INL plot is reduced after optimization
Transition

## Slide 23: Improved Analog Monitor Module
  Saturation issue observed during the evaluation of the analog monitor at cold temperature was fully understood
  The issue was reproduced in simulations using our cryogenic models
  This module is implemented by an analog multiplexer followed by an intermediate single-ended buffer and a differential output stage
  The limited input range of the single-ended buffer was the cause of the saturation
  The second version of CRYO addresses this problem with an improved rail-to-rail input stage on the buffer
  Simulations at -186⁰C were performed to ensure that this problem was not present in the new revision
  Also, new analog signals are available for debugging and testing
Input | MUX | Description
5 | Ana_Mon_Cal | External signal used for calibration of the analog monitor
6 | G_bgr_1V | Reference voltage of LVDS modules
12 | LDO1_Bk1_Mon | Output of analog LDO1 - Supply of ADC section (Bottom BANK)
13 | LDO2_Bk1_Mon | Output of digital LDO2 - Supply of SAR logic and 2V level shifters (Bottom BANK)
New monitor input signals

## Slide 24: Improved Analog Monitor Module
Simulation results at -186⁰C
Measurement Results
Input
Output
Analog monitor output after improved input range of buffer stage
Response at different Pulser levels
Linearity test of front-end
Input
Output
Simulations

## Slide 25: Power Breakdown
Second version of CRYO implements a power scaling technique
ADC section benefits from this feature and it can be enabled or disabled through SACI interface
Simulations show an overall power reduction of ~24%, which yields ~19.8mW/channel
This feature is implemented by adjusting dynamically the quiescent current of the ADC signal driver
  During sampling the driver delivers the necessary current to pre-charge the capacitor array of the ADC
  During conversion the current reduction is enabled since the driver is not required by the ADC
CRYO ASIC v0.1
CRYO ASIC v0.2
Chart: 
| FE Section | ADC Section | Vref Gen | Back-End Section | LDOs | Bias Circuits
CRYO v1 Power | 0.312 | 0.58 | 0.504 | 0.151 | 0.1169 | 0.02
Total Power | 1.27 W ¦ 2.5V
Power/Channel | 19.8 mW
Total Power | 1.68 W ¦ 2.5V
Power/Channel | 26 mW
Chart: 
| FE Section | ADC Section | Vref Gen | Back-End Section | LDOs | Bias Circuits
 | 0.186 | 0.404 | 0.392 | 0.151 | 0.1169 | 0.02
Power Scaling ON
nEXO Preamp

## Slide 26: Synchronization Scheme – External Enable Signals
Upon DUNE management request, new version of CRYO implements new external enable signals to sync with:
  Internal clock generation
  Streaming data mode
Sampling Clock Enable (SampClkEn) allows the user to control the sampling time of the ASIC based on the master clock
Start-of-Readout (SRO) allows the user to control when to streaming out valid data
WIB firmware controls the ASIC synchronization.
Master Clk
SampClkEn
SRO

## Slide 27: Synchronization Scheme – Clocking & Control at WIB
DUNE
CLOCK
56MHz
Timing logic
“uniform phase aligned counter & commands (new commands: SR0En, SHClkEn, samplingCycleSync)”
SampCLKEn
SR056MHz
Trigger logic
FIFO
Data de-serializer
CRYO
62.5MHz
clock
domain
56MHz
clock
domain
DUNE
timing
Phase
monitor
Sampling Cycle Timing
CRYO working with DUNE clock
Crossing the clock domains
  WIB generates 56MHz clock
  Monitors and align clock phases
Sampling is controllable
  SampClkEn controls the sampling time
  SRO controls the transmission
Running CRYO at DUNE clock
  Sampling of 2.2MSPS per channel

## Slide 28: Synchronization Scheme – Clock Gen. and Monitor
DUNE
CLOCK
MMCM
Phase
DATA
ISERDESE3
DDR
D
CLK
CLKDIV
1000/18~56MHZ
1000/4=250MHz
1000/16=62.5MHz
62.5*16=
1000MHz
CRYO with DUNE clock
Clock is generated internally
  Using integer option 55.5MHz clock
DUNE clock is the reference time
  CRYO clock is “data” for the phase monitor
Every 9 DUNE cycles the phase repeats
USE phase to issue fast commands to CRYO
in phase
in phase

## Slide 29: Synchronization Scheme – sampClkEn Controls Sampling
CRYO clock can be generated from the DUNE clock deterministically at the WIB level
Phase relations are predictable and can be used to control CRYO through sampClkEn defined by the user
CRYO internal clocks respond to the sampClkEn signal enabling the user to define the sampling instant

## Slide 30: Acknowledgments
SLAC ASIC Team
Aseem Gupta
Camillo Tamma
Lorenzo Rota
Bojan Markovic
Savino Petrignani
Pietro Caragiulo
Umanath Kumar
Xiaobin Xu
Faisal Abu-Nimeh
Hussein Ali
Aldo Pena Perez
Angelo Dragone
SLAC DAQ Team
Ryan Herbst
Dionisio Doering
Ben Reese
Larry Ruckman
Maciej Kwiatkowski
Mark Freytag
Thung Pan
Lupe Salgado
SLAC Sensor Team
Chris Kenney
Chu-En Chang
SLAC nEXO Team
Peter Rowson
Bob Conley
Sander Braun
Brian Mong
SLAC DUNE Team
Hirohisa Tanaka
Mark Convery
Patrick Tsang
CRYO Collaborators Team
Kurtis Nishimura
Rick Van Berg
Gary Varner
Carl Bromberg
Dean Shooltz