## Slide 1: 200K, 100K Z-scans in AUP-MQXFAP_0003
Measurement dates: Nov. 8 and 12, 2019
Superconducting Magnet Division, Brookhaven National Laboratory
Reported by P. Wanderer

## Slide 2: Measurement details (1)
Two readings at each Z-position: +I and –I,
  Run 3: I(200K)=18A; Run 4: I(100K)=30A
Background subtracted by fitting a straight line to harmonics at +I and –I A.
Data centered at each Z-position to make the dipole term zero.
Data “pre-rotated” by 18 deg. to account for observed misalignment of magnet top with the rotating coil drive top.
Data taken with both 108.74 mm long and 217.872 mm long coils
Z-scan step size = 108.74 mm for full Z-range available
  Can capture the lead end fully
  Cannot capture the non-lead end region at all
No 27.185 mm steps in the lead end region.
Plots show only straight-section data.

## Slide 3: Measurement details (2)
Axial positions:
  The axial position of the transporter is recorded in inches, with z = 0 the home position at the nonlead end.
  The CL of the 110 mm coil is 220.663 mm above the CL of the 220 mm coil.
  To align the z axis 0 in the plots with the CL of the magnet, the transporter z was converted to mm and
    For the 220 mm coil, 2177.923 mm were subtracted
    For the 110 mm coil, 1957.320 mm were subtracted

## Slide 4: Measuring coil offsets

## Slide 5: Transfer function and field angle variation

## Slide 6: Sextupole terms

## Slide 7: Octupole terms

## Slide 8: Decapole terms

## Slide 9: 12-pole terms

## Slide 10: 14-Pole terms

## Slide 11: 16-pole terms

## Slide 12: 18-pole terms

## Slide 13: 20-pole terms

## Slide 14: Compare measuring coil offsets, warm and cool. Run 1 (warm measurements), courtesy A. Jain

## Slide 15: Xo offset, run 1 (warm) & Runs 2, 3 (cooldown)
Warm – run 1
200K and 100K – runs 3 and 4
200K / 18A
100K / 30A
300K / 15A

## Slide 16: Yo offset, run 1 (warm) & runs 2, 3 (cooldown)
Warm – run 1
200K and 100K – runs 3 and 4
100K / 30A
200K / 18A
300K/15A