## Slide 1: Installation of Vacuum Components for the APS-Upgrade
Tim Clute
Interim APSU Technician Group Section Lead
Advanced Photon Source
Argonne National Laboratory, Lemont, IL
OLAV-IV: 6th Workshop on the Operation of Large Vacuum Systems
April 17, 2024
Notes: Univ of Chicago Review, Aug. 30, 2010
1

## Slide 2: Agenda
Scope of installation
  Harvesting of components from APS
  Installation Plan
Vacuum Installation
Vacuum Certification
Conclusion

## Slide 3: Scope of Installation
APS-U is composed of 40 Storage Ring Sectors that each contain:
  5 modules
  1 straight section
  2 x-ray extraction lines (only 54 of them are used)
5 of the straight sections are unique
  3 contain RF cavities
  1 contains the Bunch Lengthening System
  1 contains the septum magnet for injection

## Slide 4: Scope of the Installation
In total this represents 192 unique vacuum sections up to the user beamlines
Each vacuum section underwent the following process.
  Cable pull
  Module/Table installation
  Module/Table alignment
  Vacuum connections and certification
  Cable and water line landing
  Test and Checkout

## Slide 5: Harvesting of Components
Approximately 40% of the Front End systems were harvested during the removal process for use in the APS-U.
These components were carefully tagged prior to the removal process and moved to our offsite building for refurbishment.
  New identifying QR codes were given to every component
  List of tagged components were distributed to all stake holders
  Removal team and receiving team compared lists daily to confirm all components accounted for

## Slide 6: Harvesting of Components
Harvesting process:
  Entire machine was vented to boil off N2
  All gate valves closed
  Bellows between assemblies were removed and wrapped in UHV foil, placed in special built crates
  Table assemblies blanked and filled with high purity N2
  Transported by truck for refurbishment in clean room
  Unneeded components went with rest of ring to excess

## Slide 7: Logistics Plan
APS-U scheduled order of operation for each major component
  Daily meetings for work to take place the next day
Multiple contractors to perform bulk of installation
Vacuum work scheduled for when area was clear of possible contamination
Daily reports of which sector, module, assembly was installed and ready for work

## Slide 8: Logistics Plan
Installation Started at Sector 24 and worked counterclockwise to Sector 1 and back around to 25.
  Specific team concentrated on injection area
This meant that technical teams had to be staggered so as not to overlap.
  Fortunately, many jobs did not need to be done sequentially so teams could jump around if necessary.
Full installation took approximately 9 months
  Full checkout of the system took another 6 to 8 weeks
Tunnel Video

## Slide 9: Vacuum Installation
Vacuum installation done according to APS vacuum standards
  Component to join sections prepared
    Work surfaces covered with UHV foil
    Gloves worn
    Shipping flanges removed
    Joint cleaned with alcohol
  Static component flanges removed
    Wiped with alcohol
  Joining component installed
  Joints torqued

## Slide 10: Vacuum Installation
Type of joints used at the APS
  Several types of joints are used at the APS depending on requirements of space and material
  This includes:
    Aluminum Conflat Flanges
    Stainless Conflat Flanges
    Stainless QCF Flanges
  Many of these joints required custom RF gaskets
  Specific instruction sheets handed out for each type of joint

## Slide 11: Vacuum Installation
Custom leak test carts built for the APS in order to handle pumping and fit in required spaces
Each vacuum section was independently pumped down to ~1e-5 torr before helium spray leak check performed
  Standard for leak check was no response with a helium background of 5e-10 torr-L/sec or better
At this point the vacuum section is considered installed but not certified

## Slide 12: Vacuum Certification
Each vacuum section of the accelerator storage ring was vacuum baked
  This included the 5 SR modules and any straight sections
  Front Ends, Beamlines and Booster to Storage ring transport line were not baked unless minimum pressure requirement was not reached.
    Target of 1e-9 torr vacuum
Once the bake out was complete an RGA scan was performed and compared to the APS standard

## Slide 13: Vacuum Certification
Bake out was performed with electrical heaters
  Chamber installed heaters
  Heat tapes wrapped around unheated components
  Heat plates on large flat chambers
  Circuits were wired at build
Heaters were divided up into circuits run in either series or parallel to help balance the load and keep heating consistent.
  Circuits needed to be checked prior to use

## Slide 14: Vacuum Certification
Bake out procedure
  System pumped down to 5e-6 torr
  Heaters ramped at 20 C/hr to desired temperature
    NEG coated to 100 C / Uncoated chambers to 130 C
  Once temperature reached, dwelled for 16 hours
  Crossover process
    NEG coated to 130 C / Uncoated chambers to 100 C
    NEG strips/cartridges conditioned
  Dwelled for 16 hours
  Cooled down while NEG systems activated
  Isolated and Ion Pumps turned on

## Slide 15: Vacuum Certification

## Slide 16: Vacuum Certification
Once the bake out was completed RGA and final leak check performed
Data stored in e-traveler system
Monitoring of ion pumps and vacuum gauge readings
Current vacuum levels ~5e-10 torr with clean RGAs throughout

## Slide 17: Conclusion
Challenges
  Learning curve when it came to gasket installation
  Adjustment of bake out heater circuits
  Adjustment of NEG condition/activation procedures
  Identification of leaks
  Contamination
Lessons Learned
  Practice makes better
  Flexibility
  Communication
  Team did an amazing job completing this project

## Slide 18: Questions?