## Slide 1: FRIB Fragment Separator Vacuum Operational Experience
Brandon Ewert Vacuum Operation Group Leader ewert@frib.msu.edu
April 2024

## Slide 2: Outline
Overview of Facility
ARIS Vacuum Systems
Commissioning
Operation
Upgrades
Current Projects

## Slide 3: Overview of Facility
Facility for Rare Isotope Beams
  National Superconducting Cyclotron Laboratory merged in 2020
  A1900 Reconfigured 2020-2021
  LINAC BDS completed in 2021
DOE User Facility
  Operating user experiments since 2022 at 4000+ hrs/yr
Capabilities:
  2 Ion sources
  200MeV/u from O-16 to U-238
    Planned upgrade to 400MeV/u
  10kW CW
    20kW planned for late 2024, ramp up to 400kW over 5 years
  Beam stopping and Reacceleration up to 6MeV/u
    Planned upgrade up to 12MeV/u

## Slide 4: Vacuum Group Structure
Linac Vacuum & SRF Processing – Lead by K. Elliot
  Operate and maintain particle-free UHV in LINAC
  Support Front End vacuum system operation & maintenance
  Process cold masses for cyromodules
  Maintain insulating vacuum on cryostats
Vacuum Operation Group – Lead by B. Ewert
  Operate and maintain Fragment Separator & Experimental vacuum
  Operate and maintain detector gas handling systems
  Design and build new warm beamlines
  Component servicing and processing
Chart: 
| Operation | Maintenance | Design | Procurement | Processing | Assembly
Vacuum Operation Group Effort Breakdown | 20 | 25 | 25 | 5 | 10 | 10
Operation
(Trying to reduce!!)
Maintenance
Design
Procurement
Processing

## Slide 5: Overview of ARIS Vacuum System
Target Hall – 6 700L/s TMP stations
P < 2E-5 Torr, V ~ 15200L
Vertical Preseparator – 3 700L/sTMP stations, 3 segments
P < 2E-6 Torr, V ~ 4000L
Reconfigured A1900 – 5 550L/s TMP stations, 5 segments
P < 2E-6 Torr, V ~ 2500L
Advanced Rare Isotope Separator (ARIS) – 110m beam path
BDS (LINAC) – 1 550L/s TMP, 4 70L/s IP
P from 1E-7 to 1E-10 Torr, V ~ 300L
Transfer Lines – 5 300L/s TMP stations, 5 segments
P < 2E-7 Torr, V ~ 1000L

## Slide 6: Target Hall Vacuum System
Target Vessel houses:
  Rotating carbon target
    ~1% of beam interacts and fragments
  Dual warm-iron quad-triplets
    Focusing before and after separation
  Superconducting dipoles 1, 2
    1st order separation from primary beam
  Stationary beam dump
    Higher power requires rotating design
  Wedge & viewer system
    Reduces momentum spread from target interaction
    Allows feedback for beam tuning
  Reentrant steel shielding
    1.5m thick, Ni-plated

## Slide 7: Target Hall Vacuum System Specifications
6 700L/s TMP Stations
2 650m^3/hr RB/Screw Pumps
2 custom all-metal gate valves
  Allow isolation of Target, Wedge chambers from Beam Dump
1E-6 Torr base pressure (2E-4 Torr “Beam On”)
  Atm -> 2E-4:
    Entire vessel – 2hrs
    Target/Wedge vessel – 5 / 7min
Dual o-ring seals on chamber lids with interstitial pumping
Fully remotely operable

## Slide 8: Target & Wedge Isolation Valves
Commissioned late 2023
Developed with Mewasa
All stainless steel
Sealed via inflation of bellows to 20psia
3 concentric pillows inflated to 45psig
Interstitial areas pumped
Through-leak rate beats design of 1E-4 TorrL/s

## Slide 9: Vertical Preseparator
2 Dipoles, 4 Quad Triplets
1 Viewer station
3 700L/s TMP Stations
2 ISO400 Gate Valves
5E-8 Torr base pressure (2E-5 Torr “Beam On”)
Remotely operable, except venting
  Beamline access rarely needed

## Slide 10: Reconfigured A1900
4 Dipoles, 8 Quad Triplets
5 Diagnostic boxes
  Viewers
  PPAC Tracking Detectors (limit venting rates)
  Wedges
  Timing detectors
5 550L/s TMP Stations
5 ISO250 Gate Valves
5E-7 Torr base pressure (2E-5 Torr “Beam On”)
  2E-6 Torr typical with detectors filled
Remotely operable, partially automated

## Slide 11: Gas Handling Systems
4 in Reconfigured A1900
  More planned for upgrades and future beamlines
Supplies 5-15 Torr at 0-10sccm
  Process gas iC4H10 or CF4
  Downstream PID pressure control
  Stability mainly temperature dependent
Detector windows are 30um Aluminized Mylar
  Can usually withstand ~20 Torr dP
New systems in development with similar architecture
Special considerations for flammable gases
  ATEX rated scroll pumps
  Interlocks
  Exhaust dilution with N2

## Slide 12: Assembly and Commissioning
Issues
  Isolation valve “slinky”
    Spring solution
  Leaks in BNC, SHV feedthroughs
    Replacement with o-ring sealed versions
  Cold Cathode Gauge spikes causing nuisance
    Use of averaging in areas with redundant gauges
  Helico-flex seals with chain clamps leak after initial torque
    Re-torquing after 24hrs and leak checking
Acceptance Testing
Leak Checking
Cleaning
Pre-assembly
Assembly
Pumpdown/Leak Check
Controls Integration
Largest time-savers! Service valves on forelines proved invaluable

## Slide 13: System Operation
Pumping/venting
  Access to beamline requested by Maintenance Group
  Vacuum Operation Group (VOG) performs vent/pump
  Working to simplify, automate
System monitoring
  EPICS based controls with CS-Studio UI
  FRIB Operators contact VOG after first level of troubleshooting
  Minor/major alarms
    Minor alarms detect abnormal conditions
      50mT on foreline gauges signal tip seal wear
      50W on TMP may indicate bearing wear
    Major alarms usually ~50% of interlock threshold
      Allow decision and intervention prior to failure
  Voice alarms
    Sound when human intervention required
  “UOF” valued at $35,000/hr

## Slide 14: Operational Issues
PPACs leaking – windows, fittings, feedthroughs
  Assembly-level validation exists, vetting method for in-situ sniff testing
Long-stroke edge welded bellows with internal guides
  Internal guides dislodged and tore convolutions
  New design from Mewasa procured
Target Water Leaks 1E0 TorrL/s
  Pinhole leak on Cu e-beam weld
    Cut out, welded by hand
  Pinhole leak in Cu to SS braze
    Rebrazed
Beam Dump Water Leak – 1E-2 TorrL/s
  Leak on spring-energized seal after dump head swap
  Pressure decay test used – not sensitive enough
  Decision made to continue operation – leak is stable
  Current base pressure ~1.5E-5 Torr

## Slide 15: System Upgrades
High-vacuum quick disconnects
  Currently being vetted via limited beamline use
Automated pumpdown/venting
  “One-click” pumpdown validated and in use on A1900 diagnostic boxes
Limit switches for chamber doors
  Extra layer of safety
In-vacuum bypass valve for GHSs
  Validated use of modified poppet valves in-vacuum for low conductance bypass

## Slide 16: Other Projects
K500 SEE Facility
  Refurbishment of 1978 K500 Cyclotron and RT ion source
  Standalone beamline & user vault
High Rigidity Spectrometer (HRS)
  Expansion of ARIS, similar scope to VPS/A1900
  Vacuum calculations based heavily on empirical data from A1900
S1/S2 Beamlines
ELOSS Gas Handling System
  ~200sL gas detector at 1200Torr Xenon
  Xenon recapture & recompression
K500