## Slide 1
Curt Baffes (Fermilab)
OLAV VI
16 April 2024
Field Installation of Cryomodule Vacuum Systems at Fermilab
FERMILAB-SLIDES-24-0087-AD-PIP2

## Slide 2: Cryomodules Built, Tested, and/or Operated at Fermilab
1.3GHz cryomodules
  “ILC style” cryomodules incorporated in FAST/IOTA accelerator
  Fermi-design cryomodules built for LCLS-II project at SLAC
Fermilab-Style Cryomodules
  325MHz Single Spoke Resonator Cavity Cryomodules for Fermilab’s PIP-II Linac
  650HMz Elliptical Cavity Cryomodules for PIP-II
Others
  ATLAS-style 162.5MHz Half Wave Resonator provided by ANL
  1.3GHz single-cell cryomodules in collaboration with DESY
PIP-II HB650 Cryomodule
See Tug Arkan’s Talk

## Slide 3: Cryomodule Characteristic Vacuum
PIP-II HB650 Cryomodule Shield Cooldown
Beamline
E-9 range warm
Beamline
E-10 range cold
Cryostat
E-5 range warm
Cryostat
E-7 range cold

## Slide 4: Field Installation Scope
Installation in Cryomodule Qualification Test Stands
  CMTS1 Test stand for LCLS-II 1.3GHz Cryomodules
  PIP-II Integrated Test Stand for PIP-II 325 and 650MHz Cryomodules
Installation in Accelerators
  FAST/IOTA
  PIP-II Injection Test Accelerator
  PIP-II Linac (installations in 2025+)
PIP-II SSR1
Cryomodule
Cryostat Vacuum Connection:
Single-CM pumping stations
Beamline Vacuum Connection(s):
Low-Particulate UHV
Softwall Cleanroom/Cleanbox

## Slide 5: FAST/IOTA
Fermilab’s Accelerator R&D Facility
  SRF accelerator chain supporting Integrable Optics Test Accelerator (IOTA) ring
  300MeV e- Linac
  Two single-cavity 1.3GHz Cryomodules
  One eight-cavity 1.3GHz ILC-style Cryomodule
Conservative Vacuum Strategy
  Cryomodules are stand-alone
  Brute force ion pumping for beamline
  Long low-particulate section lengths
  Ion pumped coupler vacuum
  Redundant turbopump stations for cryostat
  Extensive diagnostics
ILC-Style Cryomodule
Talk To:
Lucy Nobrega

## Slide 6: Cryomodule Test Stand 1 (CMTS1)
Fermilab is building and testing cryomodules for LCLS-II at SLAC
  At CMTS1, they are cooled down and powered
  Cryomodules installed/tested on 1 month cycle
Vacuum design informed by lessons-learned from FAST/IOTA
Some simplifications relative to FAST/IOTA
  Beamline vacuum connection at 1 side only
  Simplified cryostat vacuum logic
  Very clean interfaces to enable quick turnaround
CMTS1
See Brian Hartsell’s Talk

## Slide 7: PIP-II Linac
See Raul Campos’sTalk
25 Cryomodules
~200m
PIP-II is an essential upgrade to Fermilab accelerator complex to enable the world’s most intense beam of neutrinos to LBNF/DUNE, and a broad physics research program for decades to come
Construction underway, completion in 2029
PIP-II Capabilities
  800MeV/1GeV H- Linac
  25 Cryomodules of 5 styles
  All cryomodules are stand-alone (rather than strings)

## Slide 8: PIP-II Linac Cryomodules
First cryomodules tested with beam, demonstrated PIP-II beam parameters
Prototype cryomodule –
first phase of cold RF testing completed
PIP2IT

## Slide 9: PIP-II Injection Test Accelerator
A full-up build of the upstream portion of PIP-II from the Warm Front End through the second cryomodule
Allowed investigation of vacuum architecture for PIP-II
  HWR Cryomodules unique vacuum systems
  Beamline interlock and fast-closing valve implementation
  More aggressive low-particulate lengths
  Combination NEG/Ion Pumps (though not yet near the cryomodules in this system)
HWR 162.5 MHz Cryomodule
See Alex Chen’s Talk
SSR1 325MHz Cryomodule

## Slide 10: PIP-II Integrated Test Stand
The same physical space was repurposed as a cryomodule test stand
Testing for prototype and production PIP-II Cryomodules
  One test stand for 325MHz cryomodules, One test stand for 650MHz cryomodules
  Testing ongoing now, continues until 2029
Test stands have independent vacuum systems
HB650 650MHz Cryomodule
SSR1 325MHz Cryomodule

## Slide 11: Cryostat Vacuum Field Installation
Conventional High Vacuum Systems (~1E-5 mbar while warm)
  Dry screw rougher for new “wet” installations
  Turbopump stations for each cryomodule to achieve ultimate warm vacuum
    Typically left on during cryomodule operation
    PIP-II Integrated Test Sand – station interface is field connection
    CMTS1 – station interface is permanent and cryostat bellows is field connection
Portable turbo station for
650MHz cryomodule test stand
Permanently installed rougher at CMTS1
To cryostat

## Slide 12: Cryostat Vacuum Field Installation
Leak detection
  Standard He Mass Spectrometer leak detection
  High background / O-ring permeation
Troubleshooting via conventional bag of tricks
Cryo-to-vacuum leak check via pressurized cryogenic circuits
  Rough localization via time of flight techniques
Direct cryostat connection example
Talk To:
Ron Kellett
N2 bagged flange

## Slide 13: Beamline Vacuum Field Installation: Clean Volumes
Temporary clean volumes of various sizes are set up to connect at cryomodule beamline gate valves
Training and experience critical to setup of volume, QC particle counter, and performance of work
Particle-safe vacuum pumping cart
Talk To:
Dan Lambert
Roll-around softwall cleanroom

## Slide 14: Beamline Vacuum Field Installation: Clean Volumes
We have used
Permanent overhead filter units
Roll-around soft-wall cleanrooms
Glovebox setups
Overhead filter carriage at FAST/IOTA
Overhead filter carriage at CMTS1
More flexible
More ergonomic

## Slide 15: Glovebox Approaches: PIP-II HWR-SSR1 Connection
Connection Mockup
Filter setup
Installation

## Slide 16: Glovebox approaches: PIP-II Integrated Test Stand
SSR1 Connection to Test Stand Beamline Vacuum Station
Talk To:
Elias Lopez
Kyle Kendziora
HB650 Connection to Test Beamline Vacuum Station

## Slide 17: Clean Volumes for PIP-II
HB650 Connection to Test Beamline Vacuum Station
For the PIP-II Linac, a mix of these approaches will be necessary
We are studying RAON’s T-shaped Clean Booth Concept
J. Kwon, Institute for Rare Isotope Science, TTC 2023

## Slide 18: Beamline Vacuum Field Work
HB650 Connection to Test Beamline Vacuum Station
We try to minimize scope of field work

## Slide 19: Cryomodule Connection In Test Stands
CMTS1 Beamline Vacuum
Talk To:
Brian Hartsell
We avoid the field connection if possible
To Cryomodule
To Beamline Pumping Station
Field Connection Spool

## Slide 20: Cryomodule Connection In Test Stands
Talk To:
Brian Hartsell
Window manifold connected in cavity string assembly cleanroom
Pumping a cryomodule from one end has been sufficient
To Faraday cup
To cryomodule
Thin AL window
Thin SS window
Dark current
CMTS1 Beamline Vacuum

## Slide 21: Cryomodule Connection In Test Stands
Talk To:
Genfa Wu
Jacopo Bernardini
Pump/Gauge manifold installed in cavity string assembly cleanroom
Rides along for
Assembly
Transport to test
RF test
Transport to PIP-II
Installation of cryomdule
Use of small NEG/IP pumps in low-particulate service has opened the door to reasonably-sized room-temperature pumping manifolds
Thin AL window
Thin SS window
PIP-II HB650 Cryomodule

## Slide 22: Adjacent Warm Beamlines
Talk To:
Alex Chen
Low-particulate handling of adjacent warm beamlines has been trending less conservative
  Many 10s of meters in FAST/IOTA
  Just a few meters in PIP-II Injection Test Accelerator
See Alex Chen’s talk
Thin AL window
Thin SS window
PIP-II Injector Test Accelerator

## Slide 23: Particle Safe Vacuum Cart
Mass Flow Control for slow pumping/backfilling in high pressure regime where particulate transport is more likely
Standard design deployed across Fermilab
Cart design includes
  Mass flow controllers
  Scroll, Turbo and Ion pumping
  Gauges and Diagnostics
Talk To:
Lucy Nobrega

## Slide 24: A skilled workforce is the key to success
Pay no attention to the man behind the curtain!
Do pay attention to the man or woman behind the curtain!
Skilled and dedicated technicians are essential to this work
Developing and continuously improving these skills is a “must’
Thanks to the OLAV organizers for this opportunity to cross-pollinate

## Slide 25
This work has been authored by Fermi Research Alliance, LLC under Contract No. DE-AC02-07CH11359 with the U.S. Department of Energy, Office of Science, Office of High Energy Physics