## Slide 1: LCLS-II Vacuum System Requirements and Performance
Giulia Lanza and Robert Coy
OLAV - April 2024

## Slide 2: LCLS-II Vacuum Systems
LCSL-II vacuum system consists of 3 independent systems:
Beam line system
RF Coupler vacuum system
Thermal isolation system within Cryomodule
1
2
3
Notes: There are multiple stages: the injector gun, L0 CM, laser heater, L1 CM, first bunch compressor, L2, second bunch compressor, L3 which accelerate the beam up to the final energy of 4GeV. Then 2km of transport line. The beam get separated in three lines: one is the dump, one is the Soft X-Ray line and one is the Hard X-Ray line.

## Slide 3: Beam Line Vacuum System
>3.5km from Gun/Injector to the Experimental Stations.
LCLS-II
  Requirement is Ultrahigh Vacuum (UHV) cleanliness along full length (though pressures may not be operationally achievable)
  Adjacent particle free cryo and ambient sections.
Beam line parts came from SLAC, Argonne, LBNL, FNAL, JLab

## Slide 4: SLAC UHV Acceptance Test Criteria
FP-202-631-14-REV8
Leak Tight: Helium leakage rate of < 2x10-10 std. cc/sec helium.
RGA Acceptance: Mass spectrometer scan performed at 150°C on each lot of components in the same test chamber
Conditions/Criteria | Limits
 | SLAC AD UHV
Temperature of entire vacuum system and RGA head ionizer | 150°C
Ratio of partial pressures of water vapor, ( 18 AMU ) to hydrogen (2 AMU) | 
Partial pressure from sum of all peaks >44 AMU | P < 1 x 10-11 Torr
Maximum single-peak partial pressure for >44 AMU | P < 5 x 10- 12 Torr
Conditions/Criteria | Limits
 | SLAC AD UHV
Temperature of entire vacuum system and RGA head ionizer | 150°C
Ratio of partial pressures of water vapor, ( 18 AMU ) to hydrogen (2 AMU) | 
Partial pressure from sum of all peaks >44 AMU | P < 1 x 10-11 Torr
Maximum single-peak partial pressure for >44 AMU | P < 5 x 10- 12 Torr

## Slide 5: SLAC’s Process for UHV Assembly and Verification
If executed properly, this sequence has many SLAC years’ proof of success to achieve < 1x10-8 Torr levels.
For LCLS-II there wasn’t any beamline in-situ bakeout after installation
(Except the VHF Gun)
Fabricate & Clean Parts: Per UHV spec FP-14
Assemble: UHV best practices that keep UHV Parts clean!
Bake:
250C
48 hrs
Verify Outgassing
Spot checks & vendor qualifications
Verify Hydrocarbons
RGA @150C
RFI
PF cleaning

## Slide 6: Argonne National Laboratory collaboration for Particle Free cleaning
High pressure rinsing
mechanical particles displacement
No hydrocarbon contamination
Bake:
250C
48 hrs
PF cleaning
Needs a periodical check of the quality of the DI water

## Slide 7: LCLS-II: Particle Free Pump Carts
LCLS-II installation Particle Free zones:
  Venting
  Pumping
  Leak detection
For pumping and venting the Mass Flow Controller is set up at 3 lN/min (40 Torr l/s)
Soft start mass flow controller
Automatic valves and interlocks to minimize human errors

## Slide 8: Actual pressure and pump down time
  On 06/22/2023
  09:00AM atmospheric pressure, start slow pump down
  12:40PM switch to fast pump down
  01:30PM start ion pumps
  09:00PM pressure 4.3x10-9 Torr
  on 08/18/2022 pressure 4.5x10-10 Torr up to now
Gauge off and on
Venting and activity

## Slide 9: Actual pressure and pump down time

## Slide 10: Beam Line Vacuum System
Cold
Room temperature
Differential pumping system

## Slide 11: LCLS-II: Differential Pumping System
To comply with the vacuum requirements described in the LCLS-II Vacuum Engineering Specifications (LCLSII-1.1-ES-0231-R0)
pressure should drop from 1x10-8Torr to 1x10-10Torr passing from the room temperature region to the cold region.
Engineering Note: LCLSII-1.1-EN-0658: A differential pumping system for LCLS-II beam pipe cold-warm transitions

## Slide 12: LCLS-II Field Installation: Portable Overhead Clean Room
Portable cleanrooms are used to connect the rafts.
Estimated time for two adjacent connection is 2 days.
3 days for additional pump down, leak check and venting (slow flow to avoid particle displacement).

## Slide 13: LCLS-II Field Installation: Portable BLA Clean Room
BLA
Portable Clean Room
Cryomodule
BLA
Interconnect bellow
Particle Free Cart

## Slide 14: Beam Line Vacuum System
Particle free areas
A region of warm beam line in the neighborhood of Cryomodules designed, fabricated, and installed following best practices for minimizing particle contamination in order to reduce the risk of degrading SRF cavity performance.
Recommended Particle Free Length is 20 m.
Base upon Fast Valve Response Time and mean speed for Nitrogen at 35 C.
(Defined in Engineering Note: LCLSII-4.1-EN-0337)

## Slide 15: Example of leak in BC2
Moving parts – weak point for possible leak
Gate valves close
Interlock
1.
2.
3.
Fast shutter closes

## Slide 16: Cryomodule Outgassing
CM outgassing performance metrics updated for HE based on LCLS-II delivery and installation data
Outgassing reduced to 4.4x10-9 Torr-L/sec per CM
Determined LCLS-II modeling (prior to fabrication) was too conservative by about 50X
Outgassing result similar to modeling Niobium at zero and Copper-coated Stainless per SLAC spec
Cryomodule Vacuum – L4B String (Warm)

## Slide 17: LCLS-II Cryomodule Installation Performance
Gradient performance is in line with CM acceptance test measurements at FNAL and JLab
No observable change in Field Emission onsets or magnitude from installation
Attribute result to production quality and best-in-class installation techniques, including Particle-Free processes
Change in FE Onset

## Slide 18: LCLS-II – Soft and Hard X-Ray ”First Light” in 2023
93 kHz e- beam rate
First SXR Lasing - 450 eV
First HXR Lasing - 1050 eV
CD-4 Project Milestone completed October 18, 2023
All LCLS-II Threshold KPPs have been Verified
Project had more than a decade of work, including 4.5 M hours from 5 DOE National Laboratories and many collaborators
Work continues at SLAC to support user science with LCLS-II

## Slide 19
Questions?