## Slide 1
Joe Ozelis
05/14/20
HWR Status and Cryogenic Overview
A Partnership of:
US/DOE
India/DAE
Italy/INFN
UK/UKRI-STFC
France/CEA, CNRS/IN2P3
Poland/WUST

## Slide 2: Introduction – HWR CM
Cold Mass & Lid Going into the Box
The HWR CM employs “half-wave resonator” cavities, mounted on a Ti strong back, with a He manifold connecting the He ckts above them. Beam focusing elements (solenoids) and BPMs are placed after each cavity. This “cold mass” is suspended from the cryostat “lid” and inserted into the vacuum vessel “box”. This is the so-called “top down” CM design.

## Slide 3: Cryogenic Valves & Bayonets

## Slide 4: Cryogenic Valves & Bayonets
There are 7 bayonet penetrations (FNAL design) w/manual valves:
  5K input, 5K output
  50K input, 50K output, 50K “jumper” (2)
  Sub-Atm ckt
There are 6 WEKA valves/controllers:
  J-T
  J-T Bypass
  5K cooling circuits (2)
  50K cooling circuits (2)

## Slide 5: Cryogenic Circuits
50K thermal shield – upper and lower
50K coupler cooling intercepts
5K strongback cooling circuit
5K coupler cooling circuit
4K/2K cavity/solenoid supply
  2K JT H/X

## Slide 6: HWR CM Heat Load (as documented in ED0008200)
  Cavity heat loads based on nominal operating gradients and measured Q0 at that gradient.

## Slide 7: Cryogenic Instrumentation - Thermometry
Cavities : 2 ea., top (He V flange), bottom (cooldown fitting) – 16 total, Cernox
Couplers : 2 ea., 5K (flange), 50K (flange) – 16 total (8 PT, 8 Cernox)
Solenoids : 1 ea., top (He V flange) – 8 total, Cernox
Tuners : 1 ea., top plate – 8 total, PT
Thermal Shield : 5 ea., top, 4 sides – 5 total, PT
Tuner HX : 1 ea., side – 1 total, PT
2K He Manifold : 2 ea., pipe, 50K intercept – 2 total, 1Cernox, 1PT
J-T Output Valve : 1 ea., output – 1 total, Cernox
LHe Input & Output Valves: 2 ea, 5K, 50K intercepts – total 4, 2 Cernox, 2 PT
TI Strongback : 1 ea., hanger mount – 1 total, Cernox
TI Hanger : 1 ea., 50K intercept – 1 total, PT
63 Installed RTDs
Connectors : Solid Sealing FA16523 – 8 total

## Slide 8: Cryogenic Instrumentation - Heaters
Cavities : 1 ea., bottom – 8 total
Solenoids : 1 ea., bottom – 8 total
Helium Manifold : various locations – 4 total
Strongback : various locations – 4 total
2K HX Output Line : on output side – 7 total
Couplers : on coupler #1 – 1 total
32 installed heaters Caddock M850-50.0-1% : 50Ω , 50W
Connectors : Solid Sealing FA16523 – 2 total

## Slide 9: Cryogenic Instrumentation - He Level Probes
He level probes installed in 2K manifold, protrude into solenoid “chimney”.
2 probes – AMI Model 3D-28-20
  28 cm active length
  30 cm overall length
  Interfaces via Solid Sealing FA16523
  connector

## Slide 10: HWR Status
Delivered to PIP2IT in late October 2019 after attempted thermal shield repair (arrived from ANL in August 2019)
Installation followed:
  Install couplers
  Install vacuum systems
  Complete tuner gas manifold/line fabrication/installation
  Cabling
  Electrical checkout
  HPRF distribution
  Fabricate and install u-tubes
  Alignment (external) and alignment (cold mass)measurements
  Leak checking and pressure testing
    Leak in lower 50K shield, thought to be repaired, was still present
      Moved ahead with cooldown using only top shield

## Slide 11: HWR Status – Cooldown
HWR was “buttoned up” and ready for cooldown, pending ORC approval, by late January.
Began 50K top shield cooldown 2/24/2020
  Followed by 50K coupler/intercepts circuit
  Additional 2K heat load due to loss of bottom shield ~ 14W
Cooldown of 5K circuit began 2/27/2020
  Cavities, solenoids, coupler and strongback intercepts
  No cooldown (or warm-up) rate restrictions (fast or slow)

## Slide 12: HWR Status – Cooldown
A few hiccups along the way
  Balancing flow to 5K strongback circuit and 5K cold mass circuit
  Some interference with LCLS-II CM
  Strongback flow stagnated until proper valve opened
  Uncertainty about target liquid level
  PID loop tuning less than optimal initially
But now have established a parameter set that provides stable 2K operation and adequate liquid level.
  Setpoint of 90-91% on 2nd LL probe
  1st probe is lower, so equivalent setpoint of ~100% doesn’t provide control “headroom”
  Difference of about 3% (after taking account height difference) between probes indicates slight tilt of 2K header – about 1.2” between 1st/8th cavities
  Difference in fill position between solenoids (at ~0” wrt header) and cavities (~0.75” wrt header) caused some confusion in interpretation of data
  Additional 2K heat load due to absence of lower/side thermal shield estimated at 14W (8W calculated)

## Slide 13: HWR Status – Cooldown
What it looks like when it goes well…

## Slide 14: HWR Recent Status – Test Activities
Preparing for RF testing
Logistical issues delayed beginning of 2K RF testing until week of 3/23/20
Proceeded with (completed) 2K testing of magnet packages while waiting on “logistics”.
Then COVID19 lab shutdown suspended operations…
  HWR CM is sitting at 4K