## Slide 1: SMuRF: SLAC Microresonator RF Readout
Josef Frisch for the SMuRF group,
CPAD Instrumentation Frontier Workshop 2017
Technical Innovation Directorate
Advanced Instrumentation for Research

## Slide 2: Cold System
Resonators tune with TES current
Flux Ramp eliminates 1/f noise in resonators, SQUIDS and warm electronics
MKID readout systems are similar, but no flux ramp so 1/f noise is important

## Slide 3: SMuRF system for TES / SQUID resonators
Generate a set of tones to excite cryogenic micro-resonators
Detect phase shifts due to resonator frequency changes
Feedback to track resonator frequencies
Drive flux ramp and detect resonator frequency vs flux ramp
Technical challenges
  Noise
  Linearity
  Wide bandwidth
  Computation power

## Slide 4: Noise and Linearity
The requirement for low noise in the readout electronics is clear, but for a multi-line system linearity is also critical.
Original tones
Out of band
Out of band
In band!
Signal band
~N3 3rd order terms!

## Slide 5: Linearity:
If the system bandwidth includes an octave or more bandwidth, 2nd order non-linear products are important (f1-f2) is in-band.
  This can still work but requires careful design
Otherwise the first important nonlinear terms are 3rd order (f1+f2-f3) which are typically smaller.
Signal levels need to trade off noise vs nonlinearity throughout the design
  SMuRF uses a low gain HEMT and 50K amplifier
SMuRF design avoids octave bandwidths.
With randomly spaced lines, the intermodulation products look like broadband noise
With evenly spaced lines the intermodulation products are “hidden” underneath the main lines, invisible but still there
Linearity needs to be tested with random line spacing

## Slide 6: Receiver Design
Variety of trade-offs between analog and digital I/Q receivers
Transmitter design concepts very similar to receiver design
Analog I/Q scheme used for many systems
Digital I/Q scheme used in SMuRF

## Slide 7: Line Tracking
Drive RF tones track resonances
Resonance dips typically 20dB which reduces the power in the receiver system for the same S/N
Same ADC/DAC hardware and bandwidths
Low sensitivity to amplitude variation: works near a signal null
Always operates were is maximum.
Allows use of resonators where the frequency shift > linewidth.
By including reference tones away from the resonances, also insensitive to phase shifts in cables / electronics
Tracking 12 cryogenic resonators with flux ramp using earlier generation hardware
Complex amplitude vs resonator frequency (calculation)

## Slide 8: SMuRF Block Diagram

## Slide 9: SLAC Hardware Configuration (ATCA common platform)
Carrier card: FPGA, memory, backplane connections
AMC cards: ADCs, DACs, high performance front end electronics
RTM: General purpose IO, extra networks, miscellaneous
1: Carrier Card (Xilinx KU060, 2.7K slices)
2: Crate (ATCA, 1-14 slot available)
3: RTM
Each carrier supports 2 AMC application cards

## Slide 10: RF Cards
Daughter Card
Multiplexer filter
1:4 multiplexing filters
ADCs each 2X 2.5Gs/s 14 bit
DACs each 2X
2.5Gs/s 16 bit
Mixers
AMC Card
LO generators 4X
2X (AMC + daughter card) in each carrier

## Slide 11: Low Frequency Systems
18 bit DACs (X3) TES bias
RTM board (ATCA crate)
Communication to FPGA
Cryostat board
TES Bias resistors
De-latch relays
Mounted to cryostat
Flux ramp DAC

## Slide 12: System Configuration

## Slide 13: Configuration
Single ATCA includes RF, TES bias, flux ramp and amplifiers for a 4GHz block of signals
  Baseline is 4-8GHz
  Modification for 2X 2GHz blocks to 8X 500MHz blocks is straightforward
  Frequency changes (below 10GHz) are straightforward
  Synchronization data stream provided on ATCA backplane. Multiple formats supported. (includes frequency reference)
System density
  Space is approximately 4GHz / rack “U”.
  Power <250W / 4GHz system (doesn’t depend on number of resonators)
Environment:
  ATCA rated for telecom use, up to ~50C (sea level)
  Investigating high altitude operation, expect operation to ~17000’ is OK.
  Temperature stability requirements under investigation but flux-ramp algorithm is insensitive to drifts.
  Resistor bias cards may have high temperature sensitivity – but low power dissipation makes temperature control straightforward.

## Slide 14: Algorithms / Firmware (Should be its own talk!)
Firmware functions
  Process ADC data into individual channels
  Calculate line resonance frequencies
  Digitally synthesize tones to track the frequencies
  Drive flux ramp and measure line frequency vs. flux ramp phase
  Decode into per-pixel, per-flux ramp data.
  Operation to ~10KHz flux ramp
High bandwidth modification for fast sensors with low occupancy
  Feed-forward to predict line motion
  Allow flux ramp to >1MHz
  Energy resolving X-ray detectors
  Cryogenic particle detectors
  Works in simulation, testing soon.
Line tracking and flux ramp decoding of uMUX resonator
Glitch in data was determined to be a real glitch in the flux ramp drive signal generator!

## Slide 15: Status
Prototype hardware tested for single 500MHz block
Designed for 107dBc/Hz spec for 4000 detectors in 4GHz
  ADC, DAC, upmix, downmix, LO generation tested.
  Most lines >110dBc/Hz, Worst case 101 dBc/Hz, expect to fix with tuning.
TES bias and flux ramp circuits tested. Boards in fab
Line tracking firmware working using simple DDC/DDS algorithm
  Computationally efficient algorithm works in simulation and fits comfortably in FPGA
Expect full hardware / firmware in prototype stage by end of 2017
Generating 2000 randomly space lines in 4-6GHz
One full AMC card operating
Not that block gain adjustment will level output power
60 cryogenic resonators in 500MHz read out with prototype electronics

## Slide 16: Future
ADCs, DACs, high dynamic range electronics are all advancing
  6.4Gs/s, 12 bit ADC with 10GHz input bandwidth available
  9Gs/s 14 bit DAC available
In not too long it should be possible to directly digitize and synthesize the 4-8GHz RF with a single DAC / ADC pair. (not yet!)
Xilinx has announced RFSoC chips with integrated FPGA and RF A-D and D-A. (not ready for our application yet)
Expect technology to improve performance, decrease size and cost.
  We are fortunate to be using the same sort of technology that industry is driving
When to update designs? Its inevitable that by the time a project is ready for production there will be better cheaper parts available.
The best advice on upgrades comes from the late Arthur C. Clarke’s short story “Superiority”