## Slide 1
LSST sensors P. O’Connor
Brookhaven National Laboratory, Upton, NY Subsystem physicist, LSST science rafts Cosmic Visions Dark Energy
LBNL 20171115

## Slide 2: Science drivers
Filter | 5s limiting mag (Texp=30s, AM=1, Seeing = 0.67”)
u | 23.9
g | 25.0
r | 24.7
i | 24.0
z | 23.3
y | 22.1
Single image depth
Image size
Seeing | Delivered image FWHM | Instrumental contribution
0.44” | 0.56” | 0.31”
0.60” | 0.69” | 0.34”
0.80” | 0.87” | 0.34”
Spectral
Spatial | FPA active area | ≥ 90%
Temporal | Shutter open fraction | ≥ 80%
Throughput losses

## Slide 3: Sensor requirements
Parameter | Goal (Thresh.) | unit
Total active area | 3096 (2654) | cm2
QE u | 41 | %
QE g | 78 | %
QE r | 88 | %
QE i | 81 | %
QE z | 75 | %
QE y | 21 | %
Diffusion | 5.0 | mm rms
Read time | 2 (3) | sec
Read noise | 9 (13) | e- rms
Video crosstalk | 0.2 | %
Gain stability | 0.1 | %
Pixel size | 10 | mm
Nonplanarity | 18 | mm p-v
Fill factor | 92 | %

## Slide 4: Detector design
High resistivity p-type silicon CCD, 100mm thick, fully depleted, back illuminated
10mm pixels, 4K x 4K full-frame format
16-fold parallel output
4-side buttable package, ±9mm coplanarity across full set
Specified mechanical mounting and alignment features, and electrical interface
Operating temperature -100C

## Slide 5: R&D prototype phase, 2005 - 2012
Prototype contracts with 2 CCD suppliers
Establish EO and metrology test lab
Demonstrate existence proof of sensor performance
Confirm diffusion model
Measure electronic crosstalk between channels (intra-CCD)
Investigate ‘tall pixel’ effects and other nonidealities
Demonstrate 9Hz ROI readout (guide mode)

## Slide 6: 2 detector variants
Uniform illumination
Test target image
CCD in test fixture
T-E2v CCD250
T-DALSA/ITL STA3800C
(overscan displayed)
(overscan suppressed)
Package

## Slide 7: Sensor prototype performance
QE
Dark current
Charge transfer inefficiency
Height of image surface
Read noise
PSF
Usable pixel fraction > 99.99%
Radeka, V., et al. "LSST sensor requirements and characterization of the prototype LSST CCDs." Journal of Instrumentation 4.03 (2009): P03002..

## Slide 8: Crosstalk – intra-CCD
O'Connor, P. "Crosstalk in multi-output CCDs for LSST." Journal of Instrumentation 10.05 (2015): C05010.
Staggered aggressor pattern
Averaged and stretched to show victims
Crosstalk matrix

## Slide 9: Speed tests
Crosstalk vs. pixel rate
.08% worst xtalk @ 550kpix/s
Read noise vs. pixel rate
<6e- noise@ 550kpix/s

## Slide 10: “Tall pixel” effects
A. Nomerotski
flatfields
point sources: optical
X-ray
flux
position
shape

## Slide 11: Vertical slice tests, 2012 - 2016
Complete development of raft readout electronics
Establish test lab with elx. + multiple sensors in Dewar
Re-verify EO performance in integrated configuration
Demonstrate virtual oscilloscope + diagnostics

## Slide 12: Multi-CCD readout, near-final electronics
Imaging – 3 CCDs, 48-channel readout board in Dewar
(lack of) inter-CCD crosstalk

## Slide 13: Focal plane organization
3100 cm2 active area
189 science CCDs, 4K x 4K, 16-output
21 science rafts, 9 CCDs on common mount
Corner rafts for guiding and wavefront sensing

## Slide 14: Raft Tower Module (RTM)
CCDs
electronics
Photons in – photons out
Complete 144-Mpixel imager
Provide control signals for CCD operation
Low noise analog signal processing
Digitizing, multiplexing of pixel data
Diagnostics
Support and maintain alignment of sensors
Thermal management of sensors and electronics
12.7 x 12.7 x 42cm3; 12kg

## Slide 15: RTM prior to REC installation

## Slide 16: Tower in cryostat
Friday, Jan 27 2017
ETU#1 in TS-8 at BNLj
Image from first engineering RTM
Equivalent to 70 HDTV frames

## Slide 17: Production phase 2016 – early 2019

## Slide 18: RTM production status Nov. 2017
ETU1
ETU2
RTM8
RTM1
RTM2
RTM3
RTM4
RTM5
RTM6
RTM7
RTM9
RTM10
RTM11
RTM22
…
T-e2v sensors
ITL sensors
future production
assembled and tested
awaiting test
Assembled this year: 11 rafts
99 CCDs
1.584Gpix

## Slide 19: QE: values, inter- and intra-CCD variation
%
%
%
%
%
%

## Slide 20: EO performance statistics: 3 rafts (432 channels)
gain e-/ADU | 
spec | --
<x> | 0.70
1s | 0.02
95% | 0.08
read noise e- | 
spec | 9
<x> | 4.66
1s | 0.16
95% | 0.68
full well ke- | 
spec | <175
<x> | 136
1s | 5.5
95% | 20.3
PSF s | 
spec | 5
<x> | 4.12
1s | .062
95% | .160
nonlin % | 
spec | 3
<x> | 0.15
1s | 0.03
95% | 0.11
HCTI ppm | 
spec | 5
<x> | 0.72
1s | 0.76
95% | 2.69

## Slide 21: Parameter distributions, 5 rafts:
gain
noise
full well
dark current
(95th percentile)

## Slide 22: Absolute height of image surface – 6 rafts
Room temperature measurements
10mm
target

## Slide 23: Image surface coplanarity

## Slide 24: Performance goals
  2 second readout time
  9e- read noise
  QE > 75% (450 < l < 870nm)
  QE > 20% from 320 < l < 1050nm
  > 99.9995 % CTE serial and parallel
  99.0% operable pixels
  91.4% area fill factor
  < 0.24’’ FWHM diffusion PSF
  0.3% electronic crosstalk
  < 58W power per RTM

## Slide 25: Performance goals achieved
  2 second readout time 1.975s
  9e- read noise 4.7 ± 0.17 e-
  QE > 75% (450 < l < 870nm) 85%
  QE > 20% from 320 < l < 1050nm 23%
  > 99.9995 % CTE serial and parallel 99.9999%
  99.0% operable pixels 99.97%
  91.4% area fill factor 92.97%
  < 0.24’’ FWHM diffusion PSF 0.18”
  0.3% electronic crosstalk 0.03%
  < 58W power per RTM 39W

## Slide 26: Final thoughts
Combination of characteristics that optimize LSST sensors for a wide, fast, deep imaging survey don’t map readily into other DE programs:
  Noise is good but not state-of-the-art
  NIR sensitivity is good but not state-of-the-art
  Fast readout less important for spectroscopy
  More susceptible to radiation damage than p-channel CCDs for space applications
Potentially useful for transient studies (GW follow-up?)
Compact, low power mosaic building-block construction techniques adaptable to alternative sensor types
Other suggestions?

## Slide 27
THANK YOU

## Slide 28: REFERENCES
SDW2013 “System Tests of LSST Integrated Readout Chain” Florence, POC
SPIE2016 “Integrated System Tests of the LSST Raft Tower Modules”, Edinburgh, POC
”Electro-optical testing of fully depleted CCD image sensors for the Large Synoptic Survey Telescope camera”, Vol. 9154 (2014), pp. 915418-915418-17, doi:10.1117/12.2056733 by Peter E. Doherty, Pierre Antilogus, Pierre Astier, et al.
LSST sensor requirements and characterization of the prototype LSST CCDsV Radeka, J Frank, JC Geary, DK Gilmore, I Kotov, P O'Connor, P Takacs, Journal of Instrumentation 4 (03), P03002
Development of the LSST raft tower modules, Society of Photo-Optical Instrumentation Engineers (SPIE) Conference Series, Vol. 8453 (July 2012), doi:10.1117/12.926645 by P. O'Connor, I. Kotov, P. Z. Takacs, et al.
Technology of the LSST focal plane, Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, Vol. 582, No. 3. (01 December 2007), pp. 902-909, doi:10.1016/j.nima.2007.07.120, by P. Oconnor, J. Geary, K. Gilmore, J. Oliver, V. Radeka, P. Takacs
Lu, W., et al. "CCD emulator design for LSST camera." SPIE Astronomical Telescopes+ Instrumentation. International Society for Optics and Photonics, 2016.
O'Connor, P., J. Fried, and I. Kotov. "An advanced CCD emulator with 32MB image memory." SPIE Astronomical Telescopes+ Instrumentation. International Society for Optics and Photonics, 2012
P. O'Connor, I. Kotov, P. Z. Takacs, J. S. Frank, S. Plate, R. Van Berg, M. Newcomer, P. Antilogus, Hervé Lebbolo, V. Tocut, C. Juramy, P. Doherty, N. Felt, "Development of the LSST raft tower modules", Proc. SPIE 8453, High Energy, Optical, and Infrared Detectors for Astronomy V, 84530L (25 September 2012); doi: 10.1117/12.926645; http://dx.doi.org/10.1117/12.926645

## Slide 29: Design thickness study
Choice of operating temperature impacts thickness decision.
  Warmer operation improves long-wavelength QE
  Limit set by dark current and/or bright defect rate; 173K is baseline
Internal Quantum Efficiency at 1000nm
0.1
0.2
0.4
0.3
diffusion PSF s, mm
model including velocity saturation, 5kV/cm
5
4
3
6
0.20
0.15
0.25
0.30
PSF FWHM, arcsec

## Slide 30: Raft electronics
Functions:
Video CDS processing
18-bit digitization and serialization
HV clock and bias generation
CCD temperature monitoring and control
Onboard diagnostic features including waveform recorder
Implementation:
Three 4” x 16”, 14-layer boards
Video processing in CMOS ASIC (x18)
Xilinx XC7K160T FPGA (x3)
AD7982 18-bit 1MSa/s SAR ADC (x144)
+40/±15/+8/+7/+5V supplies, onboard regulators

## Slide 31: Virtual sampling scope mode of video digitizer
Arrange a repetitive CCD output waveform (i.e. dark, flat)
Acquire one sample per pixel period, incrementing the ADC sampling clock 10ns per pixel
Combined with “transparent mode” feature of DSI ASIC, creates virtual 100MHz, 18b oscilloscope for probing video channels
(a) ASPIC DSI mode
(b) ASPIC Transparent mode – amplified CCD output

## Slide 32: 144-channel waveforms using virtual ‘scope

## Slide 33: Noise correlations
0.1
0.2
0
Single CCD
Full raft CCD