## Slide 1: High Speed PADs at Cornell
Practical experience with high-speed
x-ray pixel array detectors and x-ray
sensing materials

## Slide 2: Who we are
Permanent members
Some past members
Kate Shanks
Sol Gruner
Hugh Philipp
Mark Tate
Prafull Purohit
Lucas Koerner
(U. St. Thomas)
Dan Schuette
(MIT-LL)
Darol Chamberlain
Joel Weiss
(Apple)
Julian Becker
(DESY)
Marty Novak
Marianne Hromalik
(SUNY Oswego)
Present Student members
Divya Gadkari (graduate), Jonathan Okasinski (undergrad)

## Slide 3: Outline Slide
Landscape of detector development.
Cornell detectors and practical use
Way forward

## Slide 4: “Landscape” of detector capabilities…
1s
100 ms
10 ms
1 ms
100 µs
10 µs
1 µs
100 ns
1 Hz
10 Hz
100 Hz
1 kHz
10 kHz
100 kHz
1 MHz
10 MHz
Phosphor-coupled CCDs
GE flat-panel detector
(30 Hz)
SACLA MP-CCD
(60 Hz)
CS-PAD
(120 Hz)
Pilatus 100k
(300 Hz)
MMPAD (1.1 kHz)
Medipix3
(12-bit @ 2 kHz
1-bit @ 24 kHz)
Keck PAD
(10 MHz)
AGIPD
(4.5 MHz)
pnCCD
(200 Hz)
Pilatus 6M X Series (100 Hz)
In-pixel frame storage
Eiger
(12-bit @ 8 kHz,
4-bit @ 22 kHz)
*Frame rates listed are maximum stated frame rates as of June 2015
**Maximum frame rate is often defined as 1/(read time)

## Slide 5
“Landscape” of detector capabilities…
Continuous framing more difficult
Sensor material
Limitations
100 MHz
1 GHz
10 GHz
10 ns
1 ns
100 ps
Original
MaRIE Imager Specs
Type I,Type II
500 MHz, 3 GHz
Continuous framing
APS 1296 bunch mode
(350 MHz)
Strong function of x-ray energy
5

## Slide 6
“Landscape” of detector capabilities…
Sensor material
Limitations
100 MHz
1 GHz
10 GHz
10 ns
1 ns
100 ps
Original
MaRIE Imager Specs
Type I,Type II
500 MHz, 3 GHz
Continuous framing
APS 1296 bunch mode
(350 MHz)
Strong function of x-ray energy
6
Increased frame rates
Triggering & synchronization
Data transfer
Data Storage

## Slide 7: Cornell detector development efforts
Materials
  Si, CdTe, Ge, GaAs
Continuous framing
  Developing 10 kHz+ continuous framerates.
High Dynamic Range
  MMPAD* and related detector development
Burst mode imaging
  KeckPAD*
Scientific collaboration - to find new applications for detectors.
  Tools create opportunities!
* Being commercialized by Sydor Instruments

## Slide 8: PAD structure
Not to scale
ASIC
sensor
HV
sensor
thickness
Cornell typical sensor:
Si (500 μm)
CdTe (750 μm)
ASIC: custom CMOS, most recent TSMC 0.18um

## Slide 9: Materials

## Slide 10: Materials
Material | ρ [g/cm3] | Z | Egap [eV] | E [eV] Pair creation | µe [cm2/Vs] | µh [cm2/Vs] | e | h
Silicon | 2.33 | 14 | 1.12 | 3.6 | 1500 | 400 | >1ms | >1ms
Ge | 5.33 | 32 | 0.67 | 2.9 | 3900 | 1900 | >1ms | >1ms
GaAs | 5.32 | 31,33 | 1.42 | 4.3 | 4000 | 300 | ~10ns | ~ 1ns
CdTe | 5.85 | 48,52 | 1.44 | 4.4 | 1100 | 90 | ~ 3µs | ~ 2µs
Si: low ρ,Z => low QE at high E
Ge: low Egap => needs cryo cooling
GaAs: low 
=> energy dependence
CdTe: slow; material polarization
=> speed, complexity
Biggest drawbacks:

## Slide 11: Detector Speed
Oversimplified
Mobility and peak carrier velocity in context of x-ray absorption length.
What thickness of materials are available?
153 ns
2.9 ns
2.0 ns
333ps
Thicker sensors

## Slide 12: Outline Slide
General capabilities of PADs and sensor materials
Cornell detectors and practical use
Way forward

## Slide 13: Keck PAD
In-pixel frame storage provides down to ~100 ns spacing between stored frames (~10 MHz)
Storage caps
300, 466, 500, 700 fF
Read time | 0.86 ms/frame
Storage caps/pixel | 8
Read noise | 1 photon @ 8 keV (CF = 300 fF) 4 photons @ 8 keV (CF = 1966 fF)
Well capacity | 1112 photons @ 8 keV (CF = 300 fF) 7288 photons @ 8 keV (CF = 1966 fF)
H. Philipp, et. al., (2015), J. Synchrotron Radiation. 23: 395-403.
Integration caps -
Sensors: Si and CdTe diodes.
Pixel size: 150 µm x 150 µm
Each tile: 128 x 129 pixels.
Above: Simplified Pixel Schematics

## Slide 14: Metals under stress @ CHESS G3
Collaborators: Hufnagel group @ Johns Hopkins & Army Research Office
Goal: study deformation of metal compounds under high strain rates
Energy: 10 keV
Kolsky bar apparatus
Optical video, 1 million FPS
Silicon diode sensor
C. J. Hustedt et al., J. of Dynamical Behavior of Materials. 44 (2018) 222-230.
Lambert et al., Review of Scientific Instruments 85, 093901 (2014)

## Slide 15: More on CdTe diode
CdTe : Acrorad Co., Ltd., Japan.
Bumping and module mating to ASICs:
Oy Ajat Ltd., Finland.
Diode structure: In/Pt Schottky type with a pixelated, Pt contact. Hole collecting.
Pixel size: 150 um x 150 um.

## Slide 16: CdTe Keck-PAD data from APS (beamline: DCS, 35IDE)
Bunch period
153 ns
750 μm thick sensor
e- : initial feature
holes: lagging feature
Both charge carriers relevant.
  Think GaAs:
    μn ~ 8500 cm2 V-1s-1
    μp ~ 400 cm2 V-1s-1
Figure: Julian Becker, et. al. (2017) JINST 12: P06022
~40 ns integration window

## Slide 17: Mixed-Mode PAD (MM-PAD)
Mixed analog and digital readout provides extended dynamic range
Read time (frame rate) | 0.86 ms (1.1 kHz)
Read noise (RMS) | 0.16 photons @ 8 keV
Well capacity | 4.7 x 107 photons/pix/frame @ 8 keV
Sustained count rate | >108 ph/pix/s
Instantaneous count rate | > 1012 ph/pix/s (e.g. ~200 ph/pix/pulse)

## Slide 18
MM-PAD: application (Si sensor)
Coherent diffractive imaging
Capture scattering pattern from Au test object @ 100 ms per frame
106 photons per pixel full scale
Dynamic range and tolerance of high flux are key
(PETRA III: Mancuso group from European XFEL –
see Giewekemeyer et al., Journal of Synchrotron Radiation (2014) 21, 1167-1174)
per pixel per 100 ms frame
No beamstop

## Slide 19
MM-PAD: application (CdTe sensor)
80 keV
APS Beamline: 1-ID-E
Armand Beaudoin, Kamalika Chatterjee
Ti alloys

## Slide 20
MM-PAD: application (CdTe sensor)
80 keV
Sum of all times
APS Beamline: 1-ID-E
Armand Beaudoin, Kamalika Chatterjee

## Slide 21
MM-PAD: application (CdTe sensor)
APS Beamline: 1-ID-E
Armand Beaudoin, Kamalika Chatterjee

## Slide 22: CdTe Problems
Significant polarization 1010 photons/pixel or long period of time.
Polarization: dose dependent accumulation of space charge.
Reset in few minutes with small forward bias on diode.
Becker, et. al., (2016) JINST 11: P12013
doi:10.1088/1748-0221/11/12/P12013

## Slide 23: GaAs
Testing Chromium compensated GaAs in ohmic configuration.
Carrier lifetimes low.
Charge collection strongly dependent on temperature.
Not optimized for integrating detectors.
Reference: J. Becker, et al., Characterization of chromium compensated GaAs as an X-ray sensor material for charge-integrating pixel array detectors. JINST 13 (2018) P01007.

## Slide 24
2017 submission further improvements to MM-PAD pixel
24
Dual track/hold circuit Differential output
Expose-during-readout operation
Dual 16-bit counters Expose-during-readout operation
Faster readout:
ASIC capable of < 100 ns/pixel
128 x 128 pixels in 16 channels: ~10 kHz
8 channels: ~5 kHz
4 channels: ~1 kHz
CFLG
CFHG
Notes: 24

## Slide 25
MM-PAD 2.0 - High Current Testing
25
Max oscillator
frequency
Inferred Input vs Actual Input
Design goal
No sensor, external current source
MMPAD 1.0 Saturation
Notes: 25

## Slide 26: Outline Slide
General capabilities of PADs and sensor materials
Cornell detectors and practical use
Way forward

## Slide 27: How to speed up ASIC for continuous readout?
A lot of room for improvement on the readout/ASIC side before detector limitations.
MM-PAD 2.0 running at 10 kHz:
  (10k x 1282 pixels x 32 bits/pixel) = 5.24 Gbits/s per ASIC.
Some options:
  Brute force, parallelize readout.
  Sparse readout
  Lossless data compression (easiest, don’t transfer lots of zeros)
  Use on chip optical modulators (optical Mach-Zehnder modulator 10’s of Gbps) – problem is identifying an accessible technology.

## Slide 28: Conclusions
New detector materials or geometries are essential for meeting high-speed, high energy imaging. Absorption length and mobility must be considered.
The Cornell detector group has demonstrated both high speed and high energy imaging with new materials coupled to existing PADs.
Development of high-speed, high-dynamic range imaging with continuous readout is continuing.
We are always looking for collaborators to help us put detectors in the hands of users.

## Slide 29: Who we are
Permanent members
Some past members
Kate Shanks
Sol Gruner
Hugh Philipp
Mark Tate
Prafull Purohit
Lucas Koerner
(U. St. Thomas)
Dan Schuette
(MIT-LL)
Darol Chamberlain
Joel Weiss
(Apple)
Julian Becker
(DESY)
Marty Novak
Marianne Hromalik
(SUNY Oswego)
Present Student members
Divya Gadkari (graduate), Jonathan Okasinski (undergrad)

## Slide 30: Additional Slides after here

## Slide 31
Sensor material
Limitations
100 MHz
1 GHz
10 GHz
10 ns
1 ns
100 ps
Continuous framing

## Slide 32: CdTe Problems
Polarization significant after X photons/pixel or long period of time.
Polarization: dose dependent accumulation of space charge.
Can be reset in a few minutes with small positive bias on diode.
Region A was undosed and shows increased variance but the same average signal.
Region B is in between the dosed and
undosed region and shows the lateral displacement artifact.
Region C was dosed and shows
a reduction in signal, but less variance than in region A.
Before reset
After reset
Becker, et. al., (2016) JINST 11: P12013
doi:10.1088/1748-0221/11/12/P12013

## Slide 33
MM-PAD: applications (Si sensor)
Example #2: coherent diffractive imaging
Capture scattering pattern from Au test object @ 100 ms per frame
106 photons per pixel full scale
Dynamic range and tolerance of high flux are key
(PETRA III: Mancuso group from European XFEL –
see Giewekemeyer et al., Journal of Synchrotron Radiation (2014) 21, 1167-1174)
5
4
3
2
1
0
Number of photons
Wide dynamic range gives extraordinary experimental flexibility
Example #1: materials science
Capture diffraction from Ni3Al multilayer foil undergoing thermite-like reaction @ 2 ms per frame
5 photons per pixel full scale
Single-photon sensitivity is key
(CHESS: Hufnagel group at Johns Hopkins)

## Slide 34
MM-PAD: application (CdTe sensor)
Tracking Bragg peaks of polycrystalline sample in a volume of reciprocal space.
Observe grain evolution of Ti sample under stress (in load stage).