## Slide 1: Studies at JLAB of Photodetectors for EIC Experiments
Carl Zorn*
Jefferson Laboratory
Newport News, VA
*on behalf of EIC/DIRC and EIC/RICH collaborations

## Slide 2
Radiation Detector & Imaging Group
Support design and construction of new detector systems
Technical consultants for the lab scientists and users
Development and use of imaging and non-imaging detector systems
Expertise in nuclear particle detection
Dr. Drew Weisenberger- Group Leader/detector scientist
Brian Kross- mechanical engineer
Jack McKisson- electrical engineer
John McKisson- software engineer
Dr. Seung Joon Lee- detector scientist
Dr. Wenze Xi- detector scientist
Dr. Carl Zorn- detector scientist
The Group makes use of expertise in the JLab
  Fast Electronics Group
  Data Acquisition Group

## Slide 3: Jefferson Lab – QCD Physics
Built originally as 4 GeV high intensity electron accelerator in early 90’s. Eventually reached 6 GeV.
Now beginning new 12 GeV physics program to continue QCD studies.
Baltimore
D.C.
Richmond
JLAB
Hall D
A B C
Halls

## Slide 4: 21st Century Science Questions
What is the role of gluonic excitations in the spectroscopy of light mesons?
Where is the missing spin in the nucleon?
What is the role of orbital angular momentum?
Can we reveal a novel landscape of nucleon substructure through measurements of new multidimensional distribution functions?
What is the relation of short-range nuclear structure and parton dynamics?
Can we discover evidence for physics
beyond the standard model
of particle physics?
Excited Glue

## Slide 5: Jefferson Lab 12 GeV Upgrade
CHL-2
JLab’s 12 GeV Upgrade was ranked the highest priority in the 2007 Nuclear Science Advisory Committee (NSAC) Long Range Plan
  Enhanced capabilities in existing Halls A/B/C
  Max energy: 11 GeV
New Hall-D
Max energy: 12 GeV
Two 1.1 GeV LINACs, maintain capability to deliver lower pass beam energies: 2.2 – 8.8 GeV

## Slide 6: Hall A
Hall A
nucleon form factors,
& future new experiments like Moller & SOLID

## Slide 7: Hall B (CLAS12)
CLAS12 (Hall B)
understanding nucleon structure via generalized parton distributions

## Slide 8: Hall C
SHMS (Hall C)
precision determination of valence quark properties in nucleons and nuclei

## Slide 9: CLAS12 RICH
H8500 series
need400
H12700
better single pe resolution

## Slide 10: New Hall D – Photon Beam
BCAL
BCAL
CDC
FDC
TOF
FCAL
start counter
Excited Glue

## Slide 11: Hall D – Oct,2014 – now taking beam
GLUEX (Hall D)
exploring origin of confinement by studying hybrid mesons

## Slide 12: Hall D – First Use of SiPMs
Barrel Calorimeter
3840 SiPM arrays
4x4 array of 3x3 mm2 50 µm microcell (Hamamatsu)

## Slide 13: FDIRC PID Add-on to GlueX
4 of the orginal bars
given to GlueX

## Slide 14: The prototype READOUT BOX

## Slide 15: Options for Readout
 | MaPMT | SiPM | MCP-PMT
Cost | ~ $1M/m2 | ~ $1M/m2 | Low
Position resolution | ~ mm | ~ mm | ~ mm
Time resolution | fair | good | very good
Dark rate | low | high | low ?
Single photon detection | fair | good | fair ?
Rate capability | good | good | fair ?
Radiation hardness | good | poor | good ?
Field sensitivity | poor | very good | fair ?

## Slide 16: The Future for QCD Studies - EIC

## Slide 17: RICH Detector for the EIC’s Forward Region Particle Identification (RD2013-4)
F. Barbosa, W. Brooks, M. Contalbrigo, A. Datta, M. Demarteau,
J.M. Durham, D. Fields, X. He, H.van Hecke(co-PI), J. Huang,
M. Liu, J. McKisson, Y. Qiang(co-PI), P. Rossi, M. Sarsour,
R. Wagner, W. Xi, L. Xue, Z. Zhao, C. Zorn
an EIC R&D Proposal – July 21, 2014

## Slide 18: EIC PID Requirements
Very rich physics program:
  Nucleon tomography and spin structure
  Quark hadronization
  Spectroscopy
  Many more …
Dedicated EIC machine and spectrometer
  Hermetic detector system
  Large momentum range
  Multi-particle detection in final states
Cherenkov Counter
Magnet
RICH Detector
EM Calorimeter
EM Calorimeter
Barrel EM Calorimeter
Tracking
9 m
TOF/DIRC
Ion
beam
electron
beam
Notes: We all know that EIC is a very promising project which covers a lot of important physics topics.
Therefore, dedicated …
Forward region with respect to the ion beam

## Slide 19: Dual Radiator Concepts
MEIC design geometry constraint: ~150 cm length
Focusing RICH
  Aerogel RICH: 3 – 10 GeV
  C4F10: 10 – 20 GeV
  Focused by a Fresnel lens
    absorbs scattered UV lights
  Readout resolution: < 2 mm
Proximity Focusing
  Aerogel RICH: 3 – 10 GeV
  CF4 threshold counter: 8 – 17 GeV
    signals from readout window mixed with photons from CF4
  Readout resolution: < 4 mm
Notes: SUVT acrylic, > 280 nm
Two concentric rings will be formed on the readout plane

## Slide 20: Options for Readout
 | MaPMT | SiPM | MCP-PMT | GEM
Cost | ~ $1M/m2 | ~ $1M/m2 | Low | Low
Position resolution | ~ mm | ~ mm | ~ mm | ~100 µm
Time resolution | fair | good | very good | poor
Dark rate | low | high | low ? | fair
Single photon detection | fair | good | fair ? | poor
Rate capability | good | good | fair ? | good
Radiation hardness | good | poor | good ? | good
Field sensitivity | poor | very good | fair ? | good

## Slide 21: MCP-based LAPPD
ALD MCP
1.2 mm
20 μm pore
large area photocathode
vacuum transfer system
transmission line readout
15 GS/s waveform
sampling DAQ
20×20 cm2
6×6 cm2

## Slide 22: First 6x6 cm2 samples

## Slide 23: Test Set for mini-LAPPD sample 28
Hamamatsu picosecond laser
405 nm
60 ps pulse width
diffuser
Trigger PMT R4998
Collimator
2.5 mm ø
X/Y stage

## Slide 24: Initial DAQ – 1 GHz BW, 5 Gs/s scope
A few quick results
Readout strip width
4.7 mm physical
Scan FWHM = 4.5 mm
Strip Signal transmission speed
178 µm/ps (ANL)
180 µm/ps (JLAB)
Some Key Investigations to Come
Irradiation – Neutrons and Gammas
High B-field test
Ponder a pixel-based readout
Also plan to implement full standard ADC/TDC DAQ

## Slide 25: Pulse comparison
5 ns
5 ns
ANL sample
3.5GHz BW 40 Gs/s scope
JLAB sample
1 GHz BW
5 Gs/s scope
Rise time 0.75 ns
Scope limit 0.35 ns

## Slide 26: High B field Test Setup
Up to 5 Tesla
5 inch diameter bore
30 inch length
Sample Tube – fits bore
18 inch length
PMT can moved axially
or rotated in a plane
Light source is blue LED (470 nm) fed through quartz fiber optic
All components non-magnetic

## Slide 27: High B field Test Setup
Up to 5 Tesla
5 inch diameter bore
30 inch length
Sample Tube – fits bore
18 inch length
PMT can moved axially
or rotated in a plane
Light source is blue LED (470 nm) fed through quartz fiber optic
All components non-magnetic

## Slide 28: Use the GlueX magnet
Used for LASS at SLAC, for MEGA at Los Alamos, refurbished for Hall-D
Bore inner diameter 1.85 m, length 4 m, BMAX 2.2 T @ 1500 A

## Slide 29: Add Title

## Slide 30
Purpose: Evaluation of PMT gain in B fields
Commissioning: July/August 2014
Data taking: November 2014
People: JLab: P. Nadel-Turonski, C. Zorn; USC: Y. Ilieva, T. Cao, E. Bringley; ODU: K. Park, G. Kalicy, L. Allison
Notes: Purpose: to evaluate the gain of small photon sensors in magnetic fields
The underlined names are of Tongtong and Greg who are featured working this November taking data in the facility.

## Slide 31: Some More Snaps
L->R: Yordanka Ilieva (USC)
Pawel Nadel-Turonski (JLAB)
some guy, Tongtong Cao (ODU)
Kijun Park (ODU), Eric Bringley (USC)
The limiting factor – cooling down the magnet

## Slide 32: Initial Samples Tested
Photonis PP0365G
6 µm pore
18 mm window
Double MCP
Photek 240 & 210
10 & 40 mm windows
double MCP
S20 photocathode

## Slide 33
Major Components

## Slide 34
Sensor Orientation Capabilities
X
Z (along B-field)
Y (Y’)
θ
X’
φ
Z’ (along sensor’s axis)
φ: rotation about Z’
θ: rotation about Y(Y’)
Notes: The range and step-size of phi rotation depends on the sensor’s shape.
The step-size of the theta rotation is 5deg and the range depends on the sensor’s size.
Theta is the angle between Z and Z’.
The way we collected data was: Fix theta, fix phi, then change the field from 0 T until signal disappears.
We also inserted the dark box from the other side of the magnet to simulate rotation of B field to 180 deg.

## Slide 35: Welcome to 3D printing
We use this regulaly now to make custom plastic parts “on the fly”
Several pieces were made to improve the performance of the “Dark Tube”

## Slide 36
Dark Tube
Trigger
Pulser
LED
470 nm
MCP-PMT
Diffuser
PreAmp
Discriminator
Trigger Board
fADC250
Sync
MVME
Fan In/Fan Out
NIM to ECL
Notes: I think the most notable feature here is the sampling flash ADC. The Quantum efficiency of the sensors we tested is about 10 - 15% for this wavelength.

## Slide 37
Charge Generator PS7120
Q(t)
fADC250
ADCch
ADCch
Q (pC)
Notes: In case there is time for details. But we do not make use of the calibration anywhere later in the talk, so just to demonstrate that we did it.
The fADC calibration is a two-step process: first we calibrate the negative DC output, Vm, of the charge generator: Q vs Vm. Then we take fADC calibration data for various Vm and calibrate the fADC response, ADCch: Q vs ADCch.
Currently, the uncertainties of the points are smaller than the symbol size and only account for the uncertainty of the total Q due to the choice of range of integration of the V(t) pulse. Statistical and pedestal subtraction uncertainties are not taken into account yet.
ADCch is the average sum of all background-subtracted pulse heights: ADCch=[sum(ADcch_i-ADCch_pedestal)]/N, where i is a counter over all the samples for each pulse, ADCch_pedestal is the pedestal peak position in ADCch, and N is the number of events.

## Slide 38
Average-Pulse Height (ADC ch)
time (a.u.)
θ=0°
B = 0.5 T
B = 0.0 T
B = 2.5 T
B = 3.5 T
B = 4.0 T
Photek PMT210
signal height is averaged over all events in the run
average pedestal is subtracted from pulse height at each time
increasing the field from 0. to 0.5 T leads to an increased gain
above 0.5 T the signal amplitude continuosly deacreses
Notes: Profile of the average signal (average over all events and all amplitudes in one run). The X axis shows time in units of ADC samples. The fADC digitizes the pulse height each 2 ns (our signals are very fast, however the ADC stretches the signal without, hopefully, changing its total area).
The figures are showing raw data taken in November 2014 with Photek PMT210 and just demonstrate the effect of the field and sensor rotation on the PMT output.
The pedestal is determined for each B (data run) separately and then subtracted.
The signals shown contain single- and multi-photoelectron hits.

## Slide 39
θ=0°
HV=-4.76 kV
HV=-4.44 kV
Average-Pulse Area (a.u.)
B (T)
Photek PMT210
max. high voltage: -4.8 kV
very preliminary 10% uncertainty shown
operating the sensor at nearly maximum high voltage extends the range of applied field to 5 T
Notes: This figure shows the area under the average pulse as a function of B-field for standard orientation of the sensor (Photek PMT210). Data were taken at an operating value of the high voltage and at nearly maximum high voltage. The latter allows to extend the range of applied field while still obtaining a relevant signal from the PMT.
To obtain the area reported above, only the positive pulse heights (in ADC ch) of the average signal (shown in the previous slide) are summed up. The time interval covering the positive part of the signal is the same for all B but I have not scaled down the points, thus, the units on the Y axis are arbitrary.
It is interesting to note that applying a nearly-maximum HV increases the + area of the average pulse almost 4 times, but this factor actually depends on the field: it is 3.6 times at 2 T and about 5 times at 4 T.

## Slide 40
events
ADCch
θ=0°
B = 0.5 T
B = 0.0 T
B = 2.5 T
B = 3.5 T
B = 4.0 T
Photek PMT210
high voltage: -4.4 kV
pedestal peak at about 1530 ch
gain can be extracted from fits to spectra to identify position of single-photoelectron peak
analysis of spectra is ongoing
Notes: Event distribution of pulse area in units of ADC ch. The spectra exhibit the same features as the average-pulse spectra on slide 5: 0.5 T seems to be the best field for largest output signal. From 0.5 T up, the signal amplitude decreases. These spectra are to be fit to extract the absolute gain of the sensor (that may turn out to not be feasible given that it is hard to “see” where the single phe peak is at higher fields, still we can report relative gain change with field and angles). The figure is shown to demonstrate how the increasing field changes the ADC spectra. The number of events in each spectrum is not necessarily the same and no quantitative conclusions are to be drawn.

## Slide 41: Summary
The facility is established and fully operational
Data collected
  Photek PMT210: B = (0, 5) T; θ = (0∘, 30∘, 180∘); φ=0∘, 90∘, 135∘
  Photek PMT240: B = (0, 5) T; θ = 0∘; φ=0∘
  Photonis PP0365G: B = (0, 3) T; θ = (0∘, 30∘, 180∘); φ=0∘, 90∘, 135∘
Gain analysis is ongoing
Future runs: Summer 2015

## Slide 42: Future Developments
Installation of a 4.7-T superconducting solenoid with 25-cm bore
Time resolution evaluation in B-field up to 5 T

## Slide 43: More MCP-PMTs to Test
XP85012 (25 µm)
XP85112 (10 µm)
Katod UFK-1 (5 µm)
Hamamatsu R10754-07-M16M

## Slide 44: In Summary
Continue high B-field tests of MCP-based photomultipliers
Test new mini-LAPPD (“LAP-P-Dino”?) in high B field
Also test its radiation tolerance
LAPPD are of interest because of several possibilities that are addressed:
[1] compact in format and adaptable to different sizes
[2] can have excellent time and position resolution
[3] should be operable in high magnetic fields
[4] should be reasonably radiation tolerant
[5] may be affordable as large area photodetectors
However, will keep an eye on other technologies – e.g. SiPMs
*in quantities of 500-1000, full 5x5 cm2 array can be had for $1.5K per unit