## Slide 1
Challenges and Opportunities in Biomedical Imaging
Creating & Growing New Knowledge in
Biomedical Imaging
And So Be Human Life Enriched
Chin-Tu Chen, PhD, Chien-Min Kao, PhD, Heejong Kim, PhD, and UC/FNAL/ANL & Other Colleagues
Workshop on MCP Based Detectors, Argonne, IL, USA, 12/3/2014
The University of Chicago
Crescat scientia; vita excolatur
“Let knowledge grow
from more to more, and
so be human life enriched”

## Slide 2
“Georges Charpak’s Company” - Low Dose X-ray
EOS Imaging
www.eos-imaging.com
Booth: South - Hall A: 3171
EOS was developed from a NobelPrize-winning technology by a teamof engineers, orthopedic surgeons and radiologists as a complete ultra low dose orthopedic imaging solution. EOS allows full-body imaging of patients at a dose reduction up to 90% compared to CR systems. It enables global assessment of balance and posture as well as a 3D modeling image in a weight-bearing position, and provides automatically over 100 clinical parameters to the orthopedic surgeon for pre- and post-operative surgical planning. EOS is a total solution inorthopedic imaging available in theU.S. since receiving FDA 510(k)clearance.
Notes: 2

## Slide 3
G Charpak’s Company - Low Dose X-ray
Gas Chambers
Biospace Radiologie
Majewski
TIPP2014
Notes: 3

## Slide 4
10 X Lower Dose X-ray
Special Focus:
From Pediatrics
To
Geriatrics
Biospace Radiologie
(EOS)
Majewski
TIPP2014
Notes: 4

## Slide 6: Spin-off from particle physics
David Townsed was affiliated with Georges
Majewski
TIPP2014

## Slide 7: Dual-Modality PET/CT Imaging
PET/CT Dual Modality Imaging
Townsend et al,
J Nucl Med
45 (2004) 4S–14S
Majewski
TIPP2014
Future “Total-Body” PET Imaging
EXPLORER
> $10M
Notes: L’une des principales préoccupations avec les systèmes TEP/TDM est la durée totale de la séance d’imagerie pour un scan pancorporel.
Un protocole typique pour une étude oncologique comprend:
  6 positions du lit ~ 75 cm (commençant sous les yeux jusqu’au pelvis)
  Durée du scan TEP: 1-5 min par position du lit, dépendant de la taille du patient (e.g. ~1-2 min/position pour < 60 kg, ~2-3 min/position pour 60-80 kg, ~3-4 min/position pour 80-110 kg, ~4-5 min/position pour >110 kg)
  Dose: 0.1-0.2 mCi/kg (3.7-7.4 kBq/kg)
  Traceur: FDG, 1 h après injection
  Durée du scan TDM: ~70 s (dépendant des ajustements: pas de l’hélice, épaisseur de tranche, collimation)
  Durée totale de la séance d’imagerie TEP/TDM: ~16-30 min

## Slide 8
Conventional PET Detector Design

## Slide 9
8” X 8” MCP Development from LAPPD

## Slide 10
A MCP-PET Detector Design

## Slide 11
Micro-Channel Plate PMT

## Slide 12
A MCP-PET Detector Design – Sample MCP

## Slide 13
Transmission-Line Readout

## Slide 14
MCP/TL Module + RDS4 Sampling Board

## Slide 15
A MCP-PET Detector Design (Simulation)

## Slide 16
A MCP-PET Detector Design (Simulation)

## Slide 17
A MCP-PET Detector Design (Simulation)

## Slide 18
A MCP-PET Detector Design (Simulation)

## Slide 19
A MCP-PET Detector Design (Simulation)

## Slide 20
A MCP-PET Detector Design (Simulation)

## Slide 21
A MCP-PET Detector Design (Simulation)

## Slide 22
Summary of Preliminary Investigation

## Slide 23
LAPPD PET
Two 80x60 cm2 LAPPD detectors, 4x4x20 mm LSO crystals, 40 cm spacing; axial FOV =80 cm
energy res. 15%; perfect time resolution
Count statistics (simulation):
  Zubal phantom (~191 lbs); 2mCi injected; scan starts right away
  450-650 keV energy window, 3.5 ns time window, 210-second scan
  54M trues+scatters, 9.9M randoms; scatter fraction ~32%, random fraction~21.28%
84.2 cm-diameter detector ring, 4x4x20 LSO crystals, 52 rings; axial FOV = 21.6 cm (33% longer than others); 4 bed positions to cover ~80 cm
energy res. <12%; time res ~ 555ps
Count statistics (measurement reported):
  160 lb patient; 16 mCi injected; scan starts 90 min post injection (~9.1 mCi at start of scan)
  350-650 keV energy window, 4.5 ns coincidence window; 210-second scan
  44M trues+scatters, and 67M randoms; scatter fraction ~30%, random fraction ~68%
21.6 cm
42.1 cm
Siemens mCT
80 cm
40 cm
60 cm

## Slide 24: Initial Simulation Results Whole-body phantom
Have obtained a 6-minute scan (1-minute frame), 7-12 minutes are under way.

## Slide 25
2mCi with LAPPD PET vs. 9.1mCi with mCT:
  LAPPD PET has more trues+scatters (54M vs 44M, 10M or ~23% more)
  LPPPD PET has much fewer randoms (9.9M vs 67M due to use of lower dose)
  For NECR=T^2/(T+S+R) (a measure of SNR-equivalent count): LAPPD is ~21.0M and mCT is only ~8.54M, ~2.5x gain.
  LAPPD PET uses 22% dose but yields 23% more trues (and scatters) and a SNR gain of ~2.5. This does not consider the effect of TOF measurement.
  TOF: 555 ps with mCT => an additional SNR gain of 1.36, 1.67, or 2.36 due to TOF if the LAPPD PET system has a TOF resolution of 300 ps, 200 ps, or 100 ps.
  An order of magnitude reduction in dose is possible.
Comparison

## Slide 26
Model-Based Reconstruction  Algorithm Enabled Instrumentation (PhysicsComputingEngineering)
Integrate detector, system geometry and reconstruction approach to achieve high-resolution, high-sensitivity and large field-of-view imaging.
Utilize novel algorithm enabled novel system designs.
2.4 mm
2.0
1.0
1.7
0.75
1.35
~30 g mouse
(a coronal slice)
Model based
conventional
Kao et al 2007, 2009; Chou et al 2012
~300 g rat
(MIP 3d rendering)
Key New Perspectives
Modular
Re-Configurable
Application-Specific

## Slide 27
Fully Digital PET Instrumentation
g photon
scintillators
diffusor
Multiplexer
Strip line
row/col sum
Sampling
MVT
WS
Digital Signal Processing (DSP)
DAQ (digital)
Detector (analog)
CPU
control
data
GPU
Reconstruction
visualization
Photodetectors (SiPM)
Current UChicago effort
Current UChicago effort
Control/Recon/Display
Current UChicago effort

## Slide 28
Strip-Line (SL) and Waveform Sampling (WS) DAQ
SiPMs
preamp
The position of the firing SiPM (or other PD) is determined using the propagation time difference.
PET signal pulse
~ 200 ns
~ 2-5 ns
x
x
x
x
x
x
x
x
x
x
Sampling at ~1GHz and above
DRS4 chip, 0.8-5 GHz sampling, 8 channels
DRS4 evaluation board (4 ch)
Clock
IN
Out
Waveform
stored
Inverter “Domino” ring chain
0.2-2 ns
FADC
33 MHz
10-100 mW
Switched capacitor array (analog memory)
shift register
Domino Ring Sampler
Kim et al 2009/10/12/14; Xie et al 2009; Niu et al 2012

## Slide 29
6x6 SiPM array
SensL FM30035
4x4mm pixels
6x6 LYSO array 3.9x3.9x20mm LYSOs
FPGA based MVT sampling
FPGA based MVT sampling
strip line
strip line
Row sums
Column sums
delay
v1
v2
v4
v3
PET signal pulse
Reference
voltages
pulse samples
Readout
Panel PET detector (w/ HUST)
MVT sampling Concept
(3x2 such units)
Ethernet I/O
Sync clock
Notes: Top Left: A SiPM based PET detector module. Top Right: The module contains 3x2 units of 6x6 LYSO/SiPM arrays, 1:1 coupling. The SiPM is SensL FM30036, consisting of 4 mm pixels (3 mm active area). The LYSO crystal is 3.9x3.9x20 mm. The pitch between crystal is 4.2 mm. Bottom Right: Readout for each LYSO/SiPM unit is done by using row/column sums followed by strip-line readout, resulting in 4 readout channel for 6x6 SiPMs. The strip lines are read by using MVT sampling (Bottom Left), implemented entirely on FPGA. Due to limited timing resolution of the current FPGA TDC implementation, delays are introduced on strip lines between SiPMs. Preliminary testings successful; no TOF.

## Slide 30: Combine PET & CT
CT Image
PET Image
Fused Image
Post-Therapy
Anatomy from X-Ray CT, Function from PET
Current “Standard of Care”
Images courtesy of Stig Larsson, Karolinska Institute

## Slide 31
PET/MRI: PET Inserts in 4.7/9.4/14T MRI
8x2 LYSO crystals/SiPMs read by using 2 strip lines
A prototype MRI insert holding 8 modules
SiPMs
8x1 LYSO crystals
Insert in a 9.4T Bruker scanner, main DAQ electronics are outside in the control room
Identification of crystals (5.2mm pitch)
Coincidence timing resolution ~500 ps FWHM, in/out MR
Gradient Coil
RF Coil
Kim et al, 2014 (submitted)

## Slide 32
CZT/CdTe – Based SPECT Inserts for
Hybrid SPECT/MRI & SPECT/CT Imaging
The proposed energy-resolved photon-counting (ERPC) detector. (1) CZT crystals of 4.4cm  4.5 cm  2-4 mm in size, (2) ERPC ASICs, (3) Readout PCBs, (4) indium bump-bonding between CZT detector to the ASIC, (5) wire-bonds between the ASIC and the PCBs and (6) Cathode signal out.
Readout PCB
MRI Scanner Bore
Rotary motor
RF coil
CdTe Detector
Collimator
Non-magnetic SPECT system chassis
Patient bed
1.5 m long arm to allow a remotely mounted motor
Spacing = 350µm
SPECT Insert Inside 3T MRI
Cai et al, NIMA, 2014
NIH/NIBIB R01 EB011640
Tc-99m
Super-Resolution Tracking of T-Cells
Images of 6 nCi I-125
in 750 Cells
Notes: Molecular Imaging Instrumentation – SPECT Inserts for SPECT/MRI [Collaboration with University of Illinois – Urbana/Champaign]
We are employing energy-resolved photon-counting (ERPC) solid-state semiconductor detectors such as CZT (or CdTe) to build SPECT inserts that can be incorporated with the existing MRI system to offer hybrid SPECT/MRI imaging
Upper-Left: (1) integrated detector head assembly
Upper-Right: (2) detector heads integrated with readout and signal-processing electronics
Central-Right: (3) [Left]: diagram of the 10-detector ring assembly; [Center]: SPECT assembly with MRI-compatible housing structure/support produced by 3-D printing technology; [Right]: a desk-top stand-alone laboratory SPECT system
Lower-Left: (4) Diagram of a SPECT insert assembly inside a typical MRI scanner;
Lower-Right: (5) the SPECT insert assembly used inside an actual 3T MRI system (can be inserted into a typical 9.4T animal MRI system as well)
Central-Bottom: (6) Image of a test phantom with the use of Tc-99m, the spacing between neighboring hot dots is 350 microns, illustrating the ability to achieve the targeted 200-500 microns resolution

## Slide 33
In vivo delivery vehicle of neural stem cells (NSCs) with pH-sensitive mesoporous silica nanoparticles (MSNs) loaded with doxorubicin (MSN-Dox), promoting programmed self-destruction of NSC carriers after migration to glioma cells
Yu, et al, Small, 2013
NIH/NINDS R01/R21
Control
Tumor
Stem cell
Injection (left)
Stem cell
Injection (left)
24h
24h
tumor cell
Injection (right)
Demonstration of NSCs loaded with MSN-In-111 radiotracer for tracking stem cells migrating toward glioma
SPECT/CT/MRI/Optical Imaging of Stem Cells in Brain Tumor Targeting
Chen/Lo/Cheng/Roman, Mou (NTU) & Lesniak/Balyasnikova (Surgery)
In-111
24h
SPECT/CT
CT
CT
SPECT/CT
Tumor
Bioluminescence
Notes: Molecular Imaging Chemistry – Nano-Theranostics: [collaboration with UChicago/Department of Surgery, and NTU/Chemistry’s Dr. Chung-Yuan Mou, who is an expert of mesoporous silica nanoparticles (MSNs)]
We are developing multi-modality nano-theranostic platforms for effective delivery of imaging probes as well as therapeutics to brain tumor
Upper & Left – (1) Nano-platforms using neural stem cells (NSCs), conjugated with ph-sensitive mesoporous silica nanoparticles (MSNs), to (2) carry either [Left] therapeutic drug (doxorubin); or [Upper-Right] imaging tracer (in this case, In-111 for SPECT imaging)
Central Upper – (3) electron-microscopic image of the nano-platforms
Molecular Imaging Clinical Application – Targeted delivery of imaging and therapeutic agents for brain tumors
Central Center - (4) Survival curves show improved delivery of therapeutics (Dox) to tumor by MSNs and NSCs
Central Lower – (5) Optical imaging by bioluminescence to illustrate NSC moving toward brain tumor cells
Lower Right – (6) CT and SPECT/CT images demonstrate the nanoplatform of NSC+MSN+In-111 moving toward brain tumor, and the control (no tumor) image shows no such migration
Comments: this nano-platform can carry contrast agents for other modalities such as MRI or PET, and other therapeutic drugs for treating brain tumor, or for other cancers as well

## Slide 35: Upgrades and Location of Proposed Shielded Irradiation Cave
209Bi(7Li,5n)211Rn
209Bi(a,2n)211At
78Se(p,2n)77Br
Other Medical Radioisotope Production

## Slide 36: Addressing the Impracticality of 211At Short Half Life (7.2 h) & Extremely Limited Supply
Only 4 production sites in the entire USA
For SPECT Imaging, 13.2 h t½ 123I is commercially viable and supplied from limited production sites
Benchmark: 211At would be more attractive if t½≥ 123I
211Rn
211At
14.6 h
Generator

## Slide 37
Theranostic Use of MIBG in Neuroblastoma
Cohn/Vochenbaum (Pediatrics) &
Appelbaum/Pu/Armato/O’Brien-Penney/Chen (Radiology)
M123IBG whole-body Imaging  Predict and monitor the outcome of treatments
Automated detection &
scoring of lesions  Improved reliability
computer- aided image
analysis
 Very challenging
M131IBG therapy trial  High-risk NB patients
ANL/ATLAS-produced At-211  MABG therapy

## Slide 38
Ultrafast Photonic Switching – Super TOF PET
Souris/Cheng (BSD), Wiederretch (ANL) & Others
Plasmonic
Materials
Time-of-Flight
PET
Transient Absorption
Pump/Probe Spectroscopy
CNM/ANL

## Slide 39: Novel X-Ray Imaging La Riviere (BSD) & ANL Colleagues
100 μm
Elemental Mapping
Color X-Ray Histology
~350 μm
Optimizing synchrotron micro-CT for
high-throughput phenotyping
Zebrafish
56 dpf, 14 mm
Practical Resolution Limit: ~250 nm
Source: http://zfatlas.psu.edu
60 dpf, ~14 mm
Voxel Size:1.43 µm
Source: ANL/APS