## Slide 1: Laser re-melting technique at FANL
M. Ge, G. Wu, J. Ruan, J. Ozelis, T. Nicol, D. Sergatskov, D. Hicks, L. Cooley
Fermi National Accelerator Laboratory
TESLA TECHNOLOGY COLLABORATION MEETING
April 19th ~ 22th, 2010
Fermilab, IL, USA

## Slide 2: outline
Motivation
Technical approaching
Cavity test results

## Slide 3: Cavity RF performance statistic
The data was from Jlab and DESY 9-cell cavity activity summarized by R.L. Geng and C.Ginsburg

## Slide 4: Surface defects (Pits and Bumps) caused quench Identified by Thermometry and Optical Inspection
Bump found by KEK Optical Inspection with CCD camera in AES 9-cell cavity with thermometry (Jlab and FNAL) Quench at 18 MV/m
Pit found by KEK optical inspection with CCD camera in AES #1cavity Quench at ~ 18 MV/m
500 mm
500 mm
26 um
Defects in cavity AES003 correlated to heating detected by T-mapping

## Slide 5
1.3GHz 9-cell cavity TB9ACC017
EB-Welding seam
Y-Y’
X-X’
X-X’
Y-Y’
TB9ACC017 quenched at 12.3MV/m, Pit was found at Cell #4 equator 180 deg region (quench location), the pit is 150 µm deep and 200 µm wide on the top.

## Slide 6
Diameter: 400µm,Depth: 60µm
Diameter: 1300µm, Depth: 60µm
A 15µm tiny bump in the center.
36MV/m quenched at pit region
39MV/m quenched at pit region
TE1ACC003
TE1AES004
Equator welding seam
Equator
welding seam
HAZ
Experience from Fermilab 1.3GHz single-cell cavity activity
Notes: 6

## Slide 7: Laser repair tool
Cavity
Mirror
Laser beam
Goals:
Try to push cavity gradient from 20MV/m limited by pits to 40MV/m.

## Slide 8
8
Laser source
Ar gas cylinder
CCD camera
Laser beam deliver line
Laser focus platform
X-Y stage
V-block with rollers
Monitor
Gas nozzle holder
Rotation stage

## Slide 10: Profile Comparison before and after laser re-melting
After re-melting the pit profile changed from 120µm deep to 30µm flat
Manmade pit

## Slide 11
Laser processing of 1.3GHz single-cell cavity TE1ACC003
The Pit before re-melting
After re-melting
Images was taken from Kyoto Optical Inspection machine
Screenshot of the monitor before and after laser re-melting

## Slide 12
Profile comparison before and after Laser processing
The pit profile changed from 60µm deep to 30µm flat
after re-melting and 50µm light EP
400µm in diameter
60 µm in depth
700µm in diameter
30 µm in depth

## Slide 13: TE1ACC003 vertical test results before and after laser processing
After Laser processing: EP 20µm+HPR+120C baking;
Gradient achieved 39.4MV/m, quenched at molten region;
After flux trapped into cavity, cavity quenched at 32MV/m.

## Slide 14
To confirm flux trapping
Cavity was warmed up to room temperature and cooled down to 2K

## Slide 15
After another light EP (20 micron)
Cavity was EP’d another 30µm+HPR+120C baking;
Gradient achieved 40.3MV/m, quenched at molten region;
Cavity quenched at 38MV/m after first quench.

## Slide 16
Fast thermometer response on the laser spot
Pre-heating was observed

## Slide 17: Summary
Successfully re-melted pit inside 1.3GHz single-cell cavity.
Cavity gradient achieved nearly 40MV/m.
Flux trapping degraded gradient from 39MV/m to 32MV/m.
Flux trapping was improved by additional light EP. (totally about 50µm EP).

## Slide 18: The next step
Produce manmade pit in cavity and recover by laser
Extend system to 1.3GHz 9-cell cavity.
Surface analysis via Nb coupons and samples cut off from single-cell cavity.
Pit in 1.3GHz 9-cell cavity TB9ACC017
Quenched at 12.3MV/m.