## SRF Fac
SRF FACILITIES Lessons Learnt / Learning from Experience Log
Wednesday 02/12/20
Topics in yellow, Lessons identified can be added below each topic
Flythrough of current SuRF Lab Facility 10 mins per speaker [Subject area]
Summary of Lesson / Information | Organisation Identifying The Learning Point | Partners who may benefit | Process Step / Activity Affected | Action Required? | Proposed Action | Originator / Point of Contact
There is a disparity across institutions on how vacuum systems are calibrated | Fermilab | All | Cavity testing; Cavity string assembly | Yes | Focused discussion by Vacuum Scientists | Genfa Wu; Keith Middleman; Stuart Wilde; Stéphane Berry.
Add extra rows as necessary
Setting up infrastructure for Assembly of 650MHz Cryomodule (3D models) Workstations configurations/requirements
Summary of Lesson / Information | Organisation Identifying The Learning Point | Partners who may benefit | Process Step / Activity Affected | Action Required? | Proposed Action | Originator / Point of Contact
Working area around cavity string in clean room needs to be consistent with operations expected (i.e. within the constraints of the building at SFTC's laboratory, as large as possible) | Fermilab | STFC | Clean room build | Yes | Mark Pendleton & Mitchell Kane (STFC) to discuss with Tim Ring (Fermilab) | Tim Ring, Fermilab
Add extra rows as necessary
Instrumentation and Calibration standards
Thursday 03/12/20
SuRF Cryogenic systems / Cryomodule system performance acceptance planning
Summary of Lesson / Information | Organisation Identifying The Learning Point | Partners who may benefit | Process Step / Activity Affected | Action Required? | Proposed Action | Originator / Point of Contact
Add extra rows as necessary
Facility Preparedness
Summary of Lesson / Information | Organisation Identifying The Learning Point | Partners who may benefit | Process Step / Activity Affected | Action Required? | Proposed Action | Originator / Point of Contact
Add extra rows as necessary
HPR current set-up and proposal for 650MHz cavities (20 mins per institution)
Summary of Lesson / Information | Organisation Identifying The Learning Point | Partners who may benefit | Process Step / Activity Affected | Action Required? | Proposed Action | Originator / Point of Contact
Add extra rows as necessary
SuRF Lab SuraBee system overview: Data management systems - what are the requirement for outputs from the system?
Summary of Lesson / Information | Organisation Identifying The Learning Point | Partners who may benefit | Process Step / Activity Affected | Action Required? | Proposed Action | Originator / Point of Contact
Add extra rows as necessary
Data transfer mechanisms. Responsibility of designer of record.
Summary of Lesson / Information | Organisation Identifying The Learning Point | Partners who may benefit | Process Step / Activity Affected | Action Required? | Proposed Action | Originator / Point of Contact
Add extra rows as necessary
Close Out/Summary Report
Summary of Lesson / Information | Organisation Identifying The Learning Point | Partners who may benefit | Process Step / Activity Affected | Action Required? | Proposed Action | Originator / Point of Contact
Add extra rows as necessary
OPEN ISSUES LOG
Summary of Open Issue | Organisation Identifying The Issue | Partners who may benefit | When does issue need to be resolved? | Action Required? | Forum to resolve issue (technical meeting, DCB, CCB, Tech Coord meeting, Technical Workshop, other) | Originator / Point of Contact
Day 1 - Wednesday 02 December 2020
Understanding the condition of 3 cavities whilst under test @ 2K | Fermilab | STFC / CEA | Position of thermometry, configuration of magnetic field suppression equipment. Comparison of radiation measurements across different infrastructures | Yes | Technical meetings between RF, VS and Cryogenics specialists | Andrew May (STFC) / Genfa Wu (Fermilab)
Active pumping of cavities | STFC | All | What is the requirement for establishing / maintaining vacuum in cavities during testing | Yes | Requirements meeting on cavity testing | Mark Pendleton (STFC)
The STFC proposed configuration for 'dual track' cavity string support arrangements are not consistent with Fermilab's single track arrangement . Further discussions on infrastructure layout are required. | Fermilab | All | Before STFC infrastructure build plans are finalised | Yes | Meeting Fermilab, STFC, RRCAT, CEA on best layout for cavity string assembly | Genfa Wu (Fermilab)/ Mark Pendleton (STFC)
STFC to report on staff levels available for dual cold mass construction track system | Fermilab | STFC | Before STFC infrastructure build plans are finalised | Yes | Internal STFC meeting - reported to All Partners meeting | Genfa Wu
Cool down - gradient for controlled cool down of 3 cavities in STFC's cryostat (Fermilab team also investigating cool down gradients) | STFC | All | Devise best cool down gradient for testing | Yes | Potential R&D project | Andrew May
String assembly purging, pumping and back-fill requirements | Fermilab | All | Agree requirements | Yes | Technical meetings between VS specialists | Tim Ring (Fermilab) / Stuart Wilde (STFC) / Stéphane Berry (CEA)
Work flow management - how do deal with a CM that is defective on test and occupying production space needed by 2nd CM | Fermilab | STFC | Before STFC infrastructure build plans are finalised | Yes | STFC to conduct review and report back via All Partners meeting | Mark Pendleton (STFC) / Mitchell Kane (STFC) / Genfa Wu (Fermilab)
Routine meeting is required to address infrastructure issues. Are Fermilab team considering constraints imposed by partners' infra? | Fermilab | All | Before all partner labs' infrastructure build plans are finalised | Yes | Establish regular meeting | Mark Pendleton (STFC) / Mitchell Kane (STFC) / Genfa Wu (Fermilab) / Stéphane Berry (CEA) / S.K. Suhane (RRCAT)
Day 2 - Thursday 03 December 2020
HPR: Is it necessary to 'flip' the cavity? | Fermilab | All | Prior to HPR commissioning | Yes | Working party to agree requirement | Tim Ring (Fermilab)
HPR: Requirement to be formalised on flow-rate in HPR pipework that discourages / eliminates bacterial growth | Fermilab | All | Prior to HPR commissioning | Yes | Working party to agree requirement | Tim Ring (Fermilab)
HPR: Operating pressure to be agreed | Fermilab | All | Prior to HPR commissioning | Yes | Working party to agree requirement | Tim Ring (Fermilab)
HPR: Nozzle profile - especially upwards-facing nozzle, to be agreed | Fermilab | All | Prior to HPR commissioning | Yes | Working party to agree requirement | Tim Ring (Fermilab)
HPR: Wash direction. Only wash on down-stroke? | Fermilab | All | Prior to HPR commissioning | Yes | Working party to agree requirement | Tim Ring (Fermilab)
Cleanroom - consider noise associated with use of nitrogen 'air gun' | CEA | All | Prior to clean-room fit-out | Yes | Working party to agree requirement | Stéphane Berry
Standardise working distance and usage procedure for 'air guns' | Fermilab | All | Prior to clean-room fit-out | Yes | Working party to agree requirement | Tim Ring (Fermilab)
Add extra rows as necessary
#VALUE!

## Data Validation
Yes
No - info only