## Slide 1
IWAA 2012 - Fermilab
September 2012
HIE ISOLDE
Alignment and Monitoring System
Technical Design and Project Status
Jean-Christophe Gayde BE/ABP-SU
Guillaume Kautzmann BE/ABP-SU
Sebastian Waniorek BE/ABP-SU

## Slide 2
Contents
Introduction
Alignment Specifications
BCAMs
Viewports
Targets
Mechanical supporting and adjustment system
Schedule
Conclusions

## Slide 3
INTRODUCTION
Transfer line
HIE-ISOLDE SUPERCONDUCTING LINAC
~16m
High Intensity and Energy (HIE)-ISOLDE  Important upgrades of REX-ISOLDE
Goal: Increase the energy and the quality of post-accelerated Ion Beams
HIE-ISOLDE

## Slide 4
NBL
Temp. 4 K
High Vacuum
Linac and Alignment Specifications
+/-300 micr
+/-150 micr
Alignment and monitoring of the Cavities and Solenoids in the Cryomodules
Alignment w.r.to a common nominal beam line along the Linac
Permanent system
Precision demanded along radial and height axis at 1 sigma level :
  300 microns for the RF Cavities
  150 microns for the Solenoids
RF Cavities
Solenoid

## Slide 5
RF cavities and solenoid equipped with targets
CONCEPT
Creation of a closed geometrical network continuously measured
Observation and position reconstruction of Cavities and Solenoid in this Network
Alignment System
Pillars
Pillars
Cryo-module
RF Cavity
Solenoid
SYSTEM
BCAM
Metrologic Table
BCAM cameras fixed to inter-module metrological tables
Interface Atmosphere / High Vacuum  Precise viewports
BCAM observations
External Line
External Lines => Position and orientation of metrological tables and BCAMs
Internal Line
Internal Lines => Position of the targets inside the tank

## Slide 6
Alignment System
Top view
BCAM
Metrological Table
BCAM to BCAM observations
BCAM to Pillar observations
Pillar
Pillar
Side view
Ext. line
Overlapping
Overlapping zone of BCAM obs. on external lines
Double sided targets observations on internal lines
=> Redundancy

## Slide 7
Original BCAM:
Camera focal length: 72 mm
Sensor: 336 x 243 pixels 10 microns
Field of view: 40 mrad x 30 mrad
Sources: Laser Diodes 650 nm
Mounting: "Plug-in" isostatic system under the BCAM body
Double sided model  Chain of BCAMs
Resolution: 5 micro radians constructor (OSI)
Accuracy of 50 micro radians to absolute
Cable length BCAM/Driver > 60 m
Delivered calibrated (focal length, position diodes, geometric relationship with plate support)
Original BCAM  HBCAM
Camera focal length: 72 mm  50 mm
Sensor: 336 x 243 pixels 10 microns  659p × 494p, 7.4 microns
Field of view: 40 mrad x 30 mrad  ~ 100 x 70 mrad
Sources: Laser Diodes 650 nm
+ Additional synchronized illumination system
Mounting: "Plug-in" isostatic system under the BCAM body
Double sided model  Chain of BCAMs
Resolution: 5 micro radians constructor (OSI)
Accuracy of 50 micro radians to absolute
Cable length BCAM/Driver > 60 m
Delivered calibrated (focal length, position diodes, geometric relationship with plate support)
BCAMs
Developed on 1999 by Brandeis University for ATLAS Muon alignment
OSI (Open Source Instruments)
http://alignment.hep.brandeis.edu/
http://www.opensourceinstruments.com/

## Slide 8
1 micron on CCD
HBCAM - LAST NEWS FROM OSI - BRANDEIS
First tests results:
No cyclic error
Resolution spot position of 0.1 um on CCD
Spot separation over a CCD scan: rms 0.15 um
Test conditions
No cover on HBCAM Box
Ambient light on
Many thanks to Kevan Hashemi and
Jim Bensinger
HBCAM – Some news
HBCAM proto

## Slide 9
HBCAM integrated illumination system for retro-reflective targets observations
HBCAM – Some news
Illumination of Retro Reflective Targets
1st Prototype
  Validation of the components / intensity …
2nd Prototype - Brandeis
  Remotely controlled by HBCAM Driver
  No extra power supply needed
  Synchronized with the HBCAM
1st illumination prototype

## Slide 10
Viewports
Atmosphere / Vacuum interface
Parallel plates window
Viewports at CM ends (off the shelf)
Study of viewport effects on BCAM observations
Viewport 6.55 mm thick
3 opt. quality classes tested

## Slide 11
Wedge angle and wedge angle effect evaluation
Viewport Study
Wedge Angle
Window | Given wedge angle (microrad) from window’s technical data | Wedge angle observed (microrad) | Influence on target at 1m (micr) | Influence on target at 2m (micr)
A | 25 | 5 | 2.5 | 5
B | 50 | 10 | 5 | 10
C | 500 | 300 | 150 | 300
Window A,1 (micr)
Window A,2 (micr)
Window B (micr)
Window C (micr)
In red: measurements on the CCD
In blue: best fit circle
10 microrad wedge angle acceptable
Viewports better than manufacturer data
Principle: Measure a fix point through the window  Rotation of the window around the main axis  Observation of the point image coordinate change  Calculation of the wedge angle
Tests: Nicolas Gauthé

## Slide 12
Optical fiber attached to a Coordinate-measuring machine controlled with an interferometer
Window mounted on a theodolite (rotation) and a translating holder
BCAM W0226
Parallel plate effect on image at different incident angles
Incident angle change of 1gon (0.9deg)  37 microns radial object “displacement”
Match the theory by a few microns  Easy observation correction by software
Adjustment of the Window within less than 1 degree  Ease the correction
Difference Theory/observed:
Average: 0 micr
Standard deviation: 6 micr
Viewport Study
Parallel Plate Effect
50 microns
BCAM to Target distance: 1.3 m

## Slide 13
Preliminary design for a viewport alignment system
Drawings by A. Bouzoud
Adjustment of the viewports within less than one degree
  Viewport adjustment system
  Collimator (under development) or Standard Survey methods
Viewport Study
Viewport Adjustment System

## Slide 14
Y.Leclercq – CERN
Viewport Study
Vacuum Deformation
Less than 7 microns deformation at the center
Less than 0.015 degree of angular deviation
7 microns
Viewport: 6.55mm thick
Deformation measurements at Liberec University (CZ):
Results match the calculated deformations by a few microns
Same deformation on both side  Parallelism kept

## Slide 15
The Targets
Overview
Constraints  HIGH VACUUM - CRYO CONDITIONS - SIZE
Studied Target Types
  Silica Silica optical fiber end
      - feed-through needed, one-sided target
      + easy light level control, OK with cold and vacuum (tested)
  Silica Silica optical fiber ended by a ceramic ball
      - feed-through needed, connection fiber/ball
      + visible from all positions, good diffuser
  Retro-reflective targets
      - illumination needed, all targets in one shot
      + double-sided, passive target, no feed-through
      More tests in cold and vacuum ongoing
IWAA12 - Sept 2012
HIE ISOLDE Alignment System
15

## Slide 16
Double Sided Targets
Two types of “double sided” targets considered
Double sided retro-reflective target Prototype tests on-going
Test prototype for an illuminated ceramic ball synchronized to the acquisition system
3 mm diffusion ball
Fiber
Light injection
Fiber inside the support
4mm
2mm
Slot for double sided retro foil
Retro-reflective bi-directional target
Laser illuminated ceramic balls

## Slide 17
Target Movement Reconstruction
BCAM measurements on Optical Fiber End – Ref CMM measurements
Large part of the CCD surface covered
Target at 1.415 m from the BCAM lens
[BCAM to Optical Fiber End] compared to CMM  Better than 5 microns
Comparison of different types of targets  Differences at 7 microns level
5 microns in object space

## Slide 18
Target Study
Cryogenic Conditions
Targets and fiber tests in cold conditions – Liquid nitrogen at 70K
Tested: fibers, retro-reflective ball, retro-reflective targets and ceramic diffusion balls
All of them resist to 70K cold conditions
No visible crack on the fibers (Microscope)
Light transmission in the fiber still OK
Test carried out in the Cryolab with Mario Herrmann (TE/VSC)

## Slide 19
Target Study
Vacuum
Test performed by Mario Herrmann (TE/VSC)
ALL TESTED TARGET CAN BE USED
Inside the Cryomodules: Common beam and insulation vacuum
Outgassing tests of:
Ceramic balls (Al2O3 and ZrO2)
Silica/Silica optical fibers
Photogr. retro-reflective ball on anodized support
Macor plate
Retro-reflective tape: under test

## Slide 20
Integration
Target Distribution
Drawing A. Bouzoud
BCAM1 Simulation
BCAM2 Simulation
Cryomodule
Internal Line BCAM
Target
From BCAM2
From BCAM1
CCD
Target env.
Opp. BCAM
Opp. Viewport

## Slide 21
Supporting and Adjustment
Cryomodule assembly in ISO Class 5 clean room
Cavity and solenoid isostatic support: Sphere – V-shape
Precise adjustment
Solenoid adjustment allowed in operational conditions
Used as Target support
Tie-rods  Frame suspension and adjustment
Survey sockets  Assembly / CM pre-alignment

## Slide 22
Project Staging
LINAC 2014
LINAC 2016
LINAC 2017
Modular Alignment System  Adapted to staging of the project

## Slide 23
Project Planning
Minimum disturbances to the Experiments presently running

## Slide 24
Conclusions
ALIGNMENT SYSTEM MAINLY BASED ON WELL KNOWN ELEMENTS
BCAMs
  Proved and used devices
  HBCAM development  Very promising results
VIEWPORTS
  Fit well to the theory  Easy BCAM observation corrections
  Be careful in the choice of the viewport  High optical quality needed
TARGETS
  All alternatives seem to work well (Fibers – Ceramic balls – Retro targets)
  Retro targets looks promising  Passive targets
  Target support ~ Std Survey  Easy control when CM open (Clean room…)
SOFTWARE:
  Development well advanced
  Simulation of metrol. table position reconstruction  ~20 microns at 1 sigma level
  Validation tests on going  Promising results
GOAL: BE READY FOR THE VACUUM AND CRYOGENIC TESTS OF 1st CRYOMODULE

## Slide 25
Thank you

## Slide 26
Viewport Study
Vacuum Deformation
Viewport of 3.55mm thick
   25 micr (±0.4) in the center
   0.08 deg angular deviation
Results match the theory by a few microns
Same deformation on both side
   Parallelism kept
Chart: Deflection function of radius [mm]
| -31.5 | -30.24 | -28.98 | -27.72 | -26.46 | -25.2 | -23.94 | -22.68 | -21.42 | -20.16 | -18.899999999999999 | -17.64 | -16.380000000000003 | -15.120000000000001 | -13.86 | -12.600000000000001 | -11.34 | -10.079999999999998 | -8.82 | -7.5599999999999987 | -6.2999999999999972 | -5.0400000000000027 | -3.7800000000000011 | -2.5199999999999996 | -1.2600000000000016 | 0 | 1.259999999999998 | 2.5200000000000031 | 3.7800000000000011 | 5.0399999999999991 | 6.2999999999999972 | 7.5600000000000023 | 8.82 | 10.080000000000005 | 11.340000000000003 | 12.599999999999994 | 13.86 | 15.119999999999997 | 16.380000000000003 | 17.64 | 18.900000000000006 | 20.159999999999997 | 21.419999999999995 | 22.68 | 23.939999999999998 | 25.200000000000003 | 26.46 | 27.72 | 28.979999999999997 | 30.240000000000002 | 31.5
Supported Edges | 0 | -0.0016555 | -0.003315 | -0.0049695 | -0.0066103 | -0.0082294 | -0.009819 | -0.011372 | -0.01288 | -0.014338 | -0.015739 | -0.017077 | -0.018347 | -0.019544 | -0.020662 | -0.021698 | -0.022647 | -0.023505 | -0.024271 | -0.02494 | -0.02551 | -0.02598 | -0.026347 | -0.02661 | -0.026768 | -0.026821 | -0.026768 | -0.02661 | -0.026347 | -0.02598 | -0.02551 | -0.02494 | -0.024271 | -0.023505 | -0.022647 | -0.021698 | -0.020662 | -0.019544 | -0.018347 | -0.017077 | -0.015739 | -0.014338 | -0.01288 | -0.011372 | -0.009819 | -0.0082294 | -0.0066103 | -0.0049695 | -0.003315 | -0.0016555 | 0
Chart: Slope function of radius [deg]
| -31.5 | -30.24 | -28.98 | -27.72 | -26.46 | -25.2 | -23.94 | -22.68 | -21.42 | -20.16 | -18.899999999999999 | -17.64 | -16.380000000000003 | -15.120000000000001 | -13.86 | -12.600000000000001 | -11.34 | -10.079999999999998 | -8.82 | -7.5599999999999987 | -6.2999999999999972 | -5.0400000000000027 | -3.7800000000000011 | -2.5199999999999996 | -1.2600000000000016 | 0 | 1.259999999999998 | 2.5200000000000031 | 3.7800000000000011 | 5.0399999999999991 | 6.2999999999999972 | 7.5600000000000023 | 8.82 | 10.080000000000005 | 11.340000000000003 | 12.599999999999994 | 13.86 | 15.119999999999997 | 16.380000000000003 | 17.64 | 18.900000000000006 | 20.159999999999997 | 21.419999999999995 | 22.68 | 23.939999999999998 | 25.200000000000003 | 26.46 | 27.72 | 28.979999999999997 | 30.240000000000002 | 31.5
Supported Edges | -0.075091 | -0.075479 | -0.075451 | -0.075025 | -0.074219 | -0.073049 | -0.071533 | -0.069689 | -0.067533 | -0.065083 | -0.062356 | -0.05937 | -0.056141 | -0.052688 | -0.049026 | -0.045175 | -0.041151 | -0.03697 | -0.032652 | -0.028212 | -0.023669 | -0.019039 | -0.01434 | -0.0095886 | -0.004803 | 0 | 0.004803 | 0.0095886 | 0.01434 | 0.019039 | 0.023669 | 0.028212 | 0.032652 | 0.03697 | 0.041151 | 0.045175 | 0.049026 | 0.052688 | 0.056141 | 0.05937 | 0.062356 | 0.065083 | 0.067533 | 0.069689 | 0.071533 | 0.073049 | 0.074219 | 0.075025 | 0.075451 | 0.075479 | 0.075091
25 microns
0.08 deg
Simulation
Experimental results by Liberec University Many thanks to Miroslav SULC
Viewport: 3.55mm thick
25 microns

## Slide 27
Effect of a lateral translation of the B viewport in front of the BCAM
Measurement of a fix point through a window translated in front of the BCAM
No visible influence / In the instrumental precision (about ±7micr at this distance)
Viewport Study
Window Uniformity
11.5 mm
Window Center
65 mm
Window
Influence on observed X object (mm)
Chart: 
| 18 | 16 | 14 | 12 | 10 | 8 | 6 | 4 | 2 | 0 | -2 | -4 | -6 | -8 | -10 | -12 | -14 | -16 | -18
X Mvt (mm) | -0.0035168 | -0.0072217 | -0.0065259 | -0.0048521 | -0.0062814 | -0.0091776 | -0.010475 | -0.010663 | -0.0089895 | -0.0028586 | -0.0047769 | -0.0055479 | -0.0058676 | -0.001072 | -0.0034604 | -0.0036861 | -0.0048897 | -0.012111 | -0.014011
0
Set-up less stable
Set-up less stable
10 microns
in Object space
Window translation (mm)
BCAM1
: Windows
: Optical axis
Window translation
X

## Slide 28
Validation of calculated target distribution - Apr 2012 - Setup for a 8 targets design
BCAM1
BCAM2
In White: the expected position according to simulations.
Integration
Target Distribution
No standard length and position for target support
Decrease of BCAM focal length to 50 mm  approx. 30% more field of view
To be refreshed with the new supporting design (more plates)
Done with prototype W0226
Field of view 70x50 mrad

## Slide 29
Large Scale Test Bench
2,5m
2m
2,5m
Complete set-up in SMI2
Almost 1:1
Cavity support mockup
Adjustable viewport table
Adjustable BCAM table
: BCAM
: Viewport
: Targets
Tests with standard BCAMs lent by OSI
Under construction