## Slide 1: 302.4.01 Magnets Vertical Test at BNL Weekly Report
Joseph F Muratore – BNL
18-Jan-2021

## Slide 2: Return to Work Planning– Building 902 High Bay
Magnet Division Electrical Group |  |  |  |  | 
Date | Days | Probable Staff | Job | # of SMD staff | Staff %
Ph 1 | D+0 | William McKeon,Frank Teich, John Cintorino | Tower water, compressors for pump and purge. | 6 | 15
 |  | Ray Ceruti, Glen Jochen, Toby Levin | Coil fabrication preparation. |  | 
 |  |  | DI water, QA cable, start cryo plant. |  | 
Ph 2 | D+15 | Above+ Piyush Joshi, Honghai Song | Validate and debug field measurement | 21 | 50
 |  | Sebastian Dimaiuta, Daniel Sullivan | LabView program. |  | 
 |  | Joseph Muratore, Patrick Doutney, Richard Felter | AUP magnet preparation. |  | 
 |  | William Themann, Dominick Milidantri | QA cable, start cryo plant, start coil fabrication. |  | 
 |  | Matthew Milidantri,Drew Carleton, Michael Hartsough | Alpha G. |  | 
 |  | Thomas Van Winckel, Jesse Schmalzle, Andy Marone |  |  | 
Ph 3 | D+30 | Above + Peter Galioto, Suresh Deonarine, Jorge Galarraga | Start AUP magnet testing. | 30 | 75
 |  | Shresht Joshi, Anis Ben Yahia, Henry Hocker, Paul Kovach | Prepare 10kA and 8kA PS for |  | 
 |  | Michael Anerella, Chris Runyan | sweet spot magnet testing. |  | 
Ph 4 | D+45 | Above + John Cozzolino, Steve Plate, Ramesh Gupta | Engineering and scientific work. | 40 | 100
 |  | Brett Parker, Kyle Capobianco, Kathleen Amm |  |  | 
 |  | Peter Wanderer, William Sampson, Marteenio Sams, |  |  | 
 |  | New Admin |  |  | 
P. Joshi and K Amm.
Started 1-June
Started 18-June
Started 6-July

## Slide 3: CRYOGENICS UPDATES
Summary of work being done on the CTI4000:
  New Phase 1 heat exchanger (LN2) and all associated plumbing is in place. One weld is presently leaking.
  Two new heat exchanger control/protection valves have been installed and wired, and the cryogenics system LabVIEW programming has been updated to include control of the new valves. So, this is DONE.
  The new in-line purifier heater has been installed and wiring is now IN PROGRESS.
  All scheduled maintenance is DONE.
  Timeline for repair of the leaking weld is uncertain at this time.
  2 out of 3 BNL welders are out; welder on medical is back but on restricted duty (bench welding only for a few days at least).
  Leak checking will then continue to make sure no other leaks.
  CTI4000 outer shell will be put back in place when all the above is finished.
So, it is now unlikely that scrubbing of the re-assembled CTI4000 will start late this week; few days further delay is possible.
Second cryogenics technician is back today, pending BNL clinic approval this morning.

## Slide 4: QUENCH ANTENNA UPDATE
  Quench antenna to be shipped to Fermilab for amplifier replacement on noisy channels.
  Packed in warm bore tube crate and ready to ship.
    No need to use coil crate.
    Fits without bending cables on lead end.
  Shipped Fri 12-Dec afternoon.
    Dedicated, air ride truck, as for coils.
    Should arrive today Mon 14-Dec.
  Arrived Fermilab on Mon 14-Dec
  Quench antenna received Mon 11-Jan at BNL.
  Test of quench antenna will be done before use in MQXFA05.

## Slide 5: MQXFA05 UPDATE
  Magnet bore video done.
  New centering fixture installed and checked for concentricity within 0.1 mm.
  Spacers to be added to WBT. Manufacturing request has been sent to BNL Central Shop.
  There might some delay since some shop personnel are out on medical/contact quarantine.
  Main and MM Test Plans are finished.
  Magnet is now in vertical test stand (Dewar 2).
  WBT with return end centering fixture installed.
  WBT spacers/bumpers not yet ready.
  Room temperature z-scan at ±15 A. DONE.
  Preliminary analysis: Centering offsets about half that of MQXFA04.
  Gravity sensors being tested and calibrated. Then gravity z-scan.
  Z-scan will be repeated after WBT (w/spacers) is installed.
  Cooldown of refrigerator and magnet 300-100K cooldown to start next week at earliest.

## Slide 6: MQXFA06 UPDATE
  Magnet MQXFA06 received at BNL on Thurs 7-Jan.
  Crate has been opened and visual inspection started. So far, no issues observed.
  Initial incoming tasks still to be done:
    Accelerometers to be removed and data downloaded.
    Initial electrical checkout including 3680V hipots on horizontal stand.
    Initial (horizontal) strain gauge readings.
    Video inspection of magnet bore.
  Tasks completion listed above will depend on status of MQXFA05.

## Slide 7: 2nd TEST FIXTURE UPDATE
  Work on second top plate/test fixture in progress.
    Mains and CLIQ leads are installed.
    Lambda plate and hanging hardware are installed.
    Still to be done:
      Connector trees and wiring.
        Wiring for connector trees is in progress with tech dedicated to this work. Need 2 months work.
        New electrical and mechanical techs are being hired and will be helping with remaining work.

## Slide 9: CRYOGENICS UPDATES
  Work on CTI4000 in progress:
    New Phase 1 heat exchanger (LN2) is in place. Welding to pipes is finished.
    One expansion engine rebuilt.
    Other expansion engine had regular maintenance done.
    Still need:
      Pressure test. WELD LEAKING. Welder out on medical. Maybe back tomorrow. Weld needs to be redone. Then continued pressure and leak testing.
      New valves installed to protect heat exchanger. DONE.
      Programming for control of new valves. DONE..
      In-line purifier heater has a short that requires repair. INSTALLED – needs to be wired.
  Cryogenic tech out on medical. WILL BE BACK TOMORROW.
  CTI4000 cooldown: latest estimate – week of 18-January 2021.

## Slide 10: MQXFA04 UPDATE
  Final electrical measurements are done – all tests OK. Electrical check sheets (Excel) are finalized.
  Test summary for CERN done.
  Magnet at Fermilab.

## Slide 11: MQXFA03 UPDATE
  Magnet at Fermilab.
  Missing jam nuts from return end found.
  Were removed to install micarta centering fixture for warm bore tube.
  Nuts placed in hardware cabinet and not re-installed before shipping.
  “Rust” on threads probably micarta dust.