## Slide 1
Lorentz Force Detuning Compensation for Pulsed Linac:
First Results and Short Term Plan
Yuriy Pischalnikov Warren Schappert
Project X Collaboration Meeting, Oct.25-27, FNAL

## Slide 2
Adaptive Least Square LFD Algorithm
has been developed at Fermilab as a part of SRF Resonance Control R&D program (Developed by Warren Schappert.)
Details of Adaptive LS LFD Algorithm at :
“W. Schappert, Y.Pischalnikov, “Adaptive Lorentz Force Detuning Compensation”. Fermilab Preprint –TM-2476-TD.
The response of the cavity frequency to the piezo impulse (TF) can be easily measured when cavity operated in CW-mode.
Since it is often not convenient to connect a pulsed cavity to CW source we developed alternative technique to measure this response (TF) when cavity operated in RF-pulse mode.
Piezo/cavity excited be sequence of small (several volts) narrow (1-2ms) pulses at various delay.
The forward, probe and reflected RF waveform recorded at each delay and used to calculate detuning.
[Response Matrix]
“1”
“10”
“34”

## Slide 3
Time during RF pulse
Piezo Pulse # (or delay)
DETUNING
0
100Hz
-100Hz
Invert the response matrix and determine combination of pulses needed to cancel out the LFD using LS
Any part of RF pulse could selected for Compensation:
“Fill+FlatTop” only “FlatTop”
Response Matrix
Adaptive LS LFD Algorithm
“1”
Optimal Compensation Pulse
for 1.3GHz 9cell cavity equipped
with KEK tuner
As operating conditions vary, the RF waveforms can be used to measure any residual detuning. The response matrix can then be used to calculate the incremental waveform required to cancel that residual detuning.
“10”
“34”

## Slide 4: HTS at MDB FNAL
Piezo OFF
Piezo ON
LFD during 1,3ms RF-pulse (Fill+FlatTop) was ~2300Hz
LS LFD compensation -- to less than 20Hz during 1,3ms pulse
Eacc=35MV/m
Open loop operation
Piezo OFF
Piezo ON

## Slide 5: Eacc=15MV/m PIEZO OFF RF pulse 4ms-Fill & 4 ms Flat-top Blade Tuner at HTS Open loop
160 DEG
4ms

## Slide 6: LFD Compensation Eacc=22.8MV/m RF pulse 4ms-Fill & 4 ms Flat-top Blade Tuner at HTS Open loop
FlatTop=4ms
Fill=4ms
8ms-RFpulse
10ms
Start of RF pulse
Piezo Stimulus Pulse
Apiezo p-to-p=100V

## Slide 7
HTS test with long (8ms) pulse --- Blade Tuner
  2 piezo
  DF=3000Hz (at 200V on piezo)
      For 23MV/m ~50% of piezo stroke required….
Plan is to use CM1 cavities #5 & #6 test with SACLAY I/DESY tuner - 1 piezo
  DF=700Hz (at 200V on piezo)

## Slide 8
Transient spectral analysis
Response of cavity /tuner system on short piezo stimulus pulse

## Slide 10: Summary
First “proof-of-principle” LFDC test with 8ms RF pulse (at 22MV/m) was successful.
FNAL’s adaptive LS LFD compensation algorithm is “key” for success…
  At S1-Global LS LFD algorithm successfully used for compensation of the different type of tuner with different transfer function
CM#1 (cavity #5&#6) is not the best starting point for LFDC test with “long pulse”
  DF=700Hz (instead of 3000Hz for Blade Tuner)
  Saclay I/DESY tuner transfer function – “good mechanical oscillator”