Purpose of This Procedure

The purpose is to provide radionuclide ground water monitoring to comply with the Illinois Environmental Protection Agency (IEPA) corrective action requirements for the Meson and Neutrino Soil Activation Areas. The objective is to monitor for radionuclides in ground water in the Silurian dolomite Class I aquifer.

Materials/Pre-requisites Needed

WH7E

Meson and Neutrino Soil Activation Areas Field Logbook

Flow-Thru Cell Logsheets - Form EP 001

Chain-of-Custody (COC) - Form EP 002

Sample Identification (SID) Labels

Calculator

A6 Key

Site 39

Calibrated pH/Temperature Meter

Conductivity Meter

Electronic Water Level Meter

Multi-probe Monitoring Chamber (MMC)

Grundfos Power Converter Box

Grundfos Sampling Tube with Tubing Adapter and Flow Diverter

Electrical Generator with GFCI Cord

Disposable Nitrile Gloves

5-Gallon Rinse Bucket with Distilled Water

5-Gallon Carboy

125-ml Poly Sample Bottles

Stopwatch

References

1)	Geotech Multiprobe Monitoring Chamber, Geotech Environmental Equipment, Inc.

2)	Making the Most of Field-Measurable Ground Water Quality Parameters, Garner, Stuart, Hydrolab Corporation, Pub No. 3043-688.

3)	Practical Guide for Ground Water Sampling, M. Barcelona et al, Illinois State Water Survey, Contract Report 374, 1985.

4)	Redi Flo 2 Installation and Operating Instructions, Grundfos Pumps Corporation, 11/94.

DEFINITIONS

Applicable to this procedure

RESPONSIBILITIES

## Beginning in 1972, high-energy particle beam interactions on fixed targets and dumps in the Meson and Neutrino areas resulted in radioactivation of surrounding soils. The targets and dumps were designed with impervious liners or membranes filled with sand and underlying gravel. Drains were installed inside the liner with underdrains outside to collect activated water that might leach from the surrounding soils, sand, and gravel. Experiments with potential to affect these areas were discontinued in 1982.

## The Resource Conservation and Recovery Act (RCRA) Part B permit issued to Fermi National Accelerator Laboratory (FNAL), designated the Meson and Neutrino Soil Activation Areas as Solid Waste Management Unit 15 (SWMU 15). See Appendix A for the SWMU 15 boundaries and monitoring well locations and Appendix B for digital images of the monitoring wells. As a SWMU, the IEPA specified the accelerator-produced radionuclides (H-3, Be-7, Na-22, Mn-54, Co-60) to be analyzed, along with routine indicator parameters (pH, temperature, conductivity), and quarterly sample frequency during the months of: January/February, April/May, July/August, and October/November.

## In 2006, as part of the RCRA Part B permit renewal process, the SWMU designation was removed from the Meson and Neutrino Soil Activation Areas site; however, quarterly monitoring and annual reporting are to continue as a courtesy to the IEPA, per agreement with the U. S. Department of Energy (DOE).

## A Multi-probe Monitoring Chamber (MMC) is used in-line with a Grundfos Redi-Flo 2 pump to enable real-time collection of pH, temperature, and conductivity data throughout the purge event. Evaluation of these three parameters allows the determination of their stabilization, implying that water representative of the screened stratigraphic unit is being withdrawn, at which time a sample is collected.

## All samples are submitted to the Radionuclide Analysis Facility (RAF) at Site 39 for analysis.

## DETAILED PROCEDURE

1)	Calibrate the pH/temperature meter following the steps outlined in the manufacturer’s manual, fill the generator with gasoline, and load all listed equipment into the field vehicle.

2)	At the well site, unlock the casing using the A6 key.

3)	Record the field sampling conditions in the logbook.

4)	Wearing a pair of disposable Nitrile gloves use the electronic water level meter to measure the depth-to-water. Reference each measurement from the top of the outer protective casing.

5)	Decontaminate the water level meter tape by running it through the five-gallon bucket of distilled water.

6)	Calculate the bore volume (gallons), using Equation 2 (given in the Special Instructions section of this document), and record the depth-to-water and the bore volume on EPG Form 001 (Flow Thru Cell) and in the field logbook.

7)	Assemble the MMC and generator apparatus (Figure 1). Loosen the appropriate fittings on the MMC and insert the pH/temperature and conductivity sensors. The sensors should be adjusted so that the tips are hanging just above the bottom of the MMC to ensure accurate readings. Carefully tighten the fittings. Do not over-tighten or the sensors and the MMC may be damaged. Screw the threaded end of the tubing adapter to the sampling port on the well and clip the quick-connect together linking the well to the MMC. Make sure the valves on the flow diverter are set to send water to the MMC. Fit one end of the motor lead plug to the power converter box and the other end to the well. Plug the power converter box into the GFCI outlet and plug the GFCI cord into the generator.

8)	On the generator, set the Engine switch to the “on” position and the A/C circuit breaker to the “off” position then start the generator allowing it to warm if necessary. Adjust the choke as needed.

9)	Turn the pH/temperature and conductivity meter(s) on.

10)	Set the A/C circuit breaker on the generator to the “on” position.

11)	Check the frequency display on the converter, it should read “0”. Start the pump by turning the Start/Stop switch into the “Start” position. Increase the flow of water by turning the speed dial until water is steadily flowing out of the MMC discharge. Do not exceed 8-psi inside the MMC. There is no gauge that shows the pressure inside the chamber, but a frequency setting of about 200 Hz should ensure the 8-psi limit is not exceeded (S-1290 requires a setting of about 270 Hz).

12)	Record pH, temperature, and conductivity readings every minute on an EPG Form 001 (Flow Thru Cell).

13)	Measure the time taken to fill the five-gallon carboy, calculate, and record the flow rate and the time to purge two bore volumes.

14)	Figure 1: Equipment Setup

15)	When at least two bore volumes have purged and the three indicator parameters stabilize, as described in the Special Instructions, it is assumed water representative of the screened geological unit is being discharged and a sample is collected by switching the valves on the flow diverter to first, discharge directly from the diverter and second, to prevent backflow from the MMC. If the parameters do not stabilize continue purging until the well purges dry. If a well purges dry prior to stabilization, collect a sample as soon as it has recharged sufficiently.

16)	Turn the Start/Stop switch on the converter box to the “Stop” position and, to prevent damage to the converter box, set the generator’s A/C circuit breaker to the “off” position BEFORE shutting down the generator. Dismantle the sampling apparatus.

17)	Label the bottle with the proper Sample Identification (SID) and analysis type.

18)	Fill out the appropriate entries in the field logbook and the COC.

19)	Take the samples to Site 39 to ensure that all labels are accurate and attached securely to the bottles. Also ensure that the COC is complete and accurate.

20)	Complete a Radiochemical Analysis Facility (RAF) Work Request Form and relinquish custody of the samples to RAF personnel. After they sign the COC, make a copy for the EP Group file on WH7E.

21)	Decontaminate the MMC and related tubing with Liqui-Nox™, tap water, and a final rinse of distilled water. Hand-dry the MMC to prevent mineral deposits.

22)	Return the COC copies to the EP Group file on WH7E for storage until analytical results are received.

23)	At WH7E enter all depth-to-water and flow-thru-cell data in their appropriate locations on the Fermilab Environmental Database (FED).

Appendix A

SWMU 15 Boundaries and Monitoring Well Locations Map

## Appendix B

Meson and Neutrino Soil Activation Areas Monitoring Well Digital Images

- MS-1 & NS1-4

- NS2-4, NS2-5, & S-1290

MS-1 - Located east of the MS-1 service building. View is to the northwest.

NS1-4 - Located east of Road C West and south of KTEV (in background). View is to the north.

NS2-4 and NS2-5 - Located east of Road C West and Proton Ditch near KTEV (not shown). View is to the south.

S-1290 – Located east of Road C south of Batavia Road. View is to the east.